## IDC Nomenclature
# | Reference ID | Version | Identifier | Partition Code | Key | Optionality | Reference ID | Co-Constraints | Display Name | Definition | Systematic Name | Class Type | Metric Type | Data Format | Unit of Measure | Discriminator | Implementation Notes | Note | Example-Value | Enumerator | Report Section Name | Report Sub Section Name | Report Field
1 | MDC_IDC | [7 empty] | Implantable Cardiac Device System |  |  | MDS | GROUP
2 | MDC_IDC_DEV |  | 0 |  |  |  | 1..1 |  | Implantable Cardiac Device | The group of cardiac device attributes. |  | VMD | GROUP
3 | MDC_IDC_DEV_TYPE | 1.0.0 | 1 | 720897 |  | R |  |  | Implantable Cardiac Device Type | The type of cardiac device. | type ¦ cardiac device ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | CRT_D | _MDC_IDC_ENUM_PG_TYPE | Demographics | Device Demographics | Device Type:
4 | MDC_IDC_DEV_MODEL | 1.0.0 | 2 | 720898 | Y | R |  |  | Implantable Cardiac Device Model | The model identifier of a cardiac device. | model ¦ cardiac device ¦ idc ¦ mdc | Metric | String | 50 |  |  |  |  | "Superpacer" |  | Demographics | Device Demographics | Device Model:
5 | MDC_IDC_DEV_SERIAL | 1.0.0 | 3 | 720899 | Y | R |  |  | Implantable Cardiac Device Serial Number | The serial number of a cardiac device. | serial number ¦ cardiac device ¦ idc ¦ mdc | Metric | String | 40 |  |  |  |  | "PM88881234" |  | Demographics | Device Demographics | Device Serial Number:
6 | MDC_IDC_DEV_MFG | 1.0.0 | 4 | 720900 | Y | R |  |  | Implantable Cardiac Device Manufacturer | The manufacturer of the cardiac device. | manufacturer ¦ cardiac device ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | BIO | _MDC_IDC_ENUM_MFG | Demographics | Device Demographics | Device Manufacturer:
7 | MDC_IDC_DEV_IMPLANT_DT | 1.0.0 | 5 | 720901 |  |  |  |  | Implantable Cardiac Device Implant Date | The implant date of the cardiac device. | implant date ¦ cardiac device ¦ idc ¦ mdc | Metric | DateTime | N/A |  |  |  |  | 20090511170125 |  | Demographics | Device Demographics | Device Implant Date:
8 | MDC_IDC_DEV_IMPLANTER | 1.0.0 | 6 | 720902 |  |  |  |  | Implantable Cardiac Device Implanter | The name of the physician that implanted the cardiac device. | implanter ¦ cardiac device ¦ idc ¦ mdc | Metric | String | 50 |  |  |  |  | "Dr. A. Miller" |  | Demographics | Device Demographics | Device Implanter:
9 | MDC_IDC_DEV_IMPLANTER_CONTACT_INFO | 1.0.0 | 7 | 720903 |  |  |  |  | Implantable Cardiac Device Implanter Contact Information | The contact information of the physician that implanted the cardiac device. | implanter contact information ¦ cardiac device ¦ idc ¦ mdc | Metric | String | 100 |  |  |  |  | "Phone: 43214321, e-mail: miller@cityclinic.org" |  | Demographics | Device Demographics | Device Implanter Contact:
10 | MDC_IDC_DEV_IMPLANTING_FACILITY | 1.0.0 | 8 | 720904 |  |  |  |  | Implantable Cardiac Device Implanting Facility | The facility (clinic / hospital) where the cardiac device was implanted. | implanting facility ¦ cardiac device ¦ idc ¦ mdc | Metric | String | 100 |  |  |  |  | "London City clinic" |  | Demographics | Device Demographics | Device Implanter Facility:
11 | MDC_IDC_LEAD |  | 64 |  |  |  | 0..N |  | Implantable Lead Attributes | The group of lead attributes. |  | VMD | GROUP
12 | MDC_IDC_LEAD_MODEL | 1.0.0 | 65 | 720961 | Y | RE |  |  | Implantable Lead Model | The model of the lead. | model ¦ lead ¦ idc ¦ mdc | Metric | String | 50 |  |  |  |  | "OptiSense XXL" |  | Demographics | Lead Demographics | Lead Model:
13 | MDC_IDC_LEAD_SERIAL | 1.0.0 | 66 | 720962 | Y | RE |  |  | Implantable Lead Serial Number | The serial number of the lead. | serial number ¦ lead ¦ idc ¦ mdc | Metric | String | 40 |  |  |  |  | "5020304789" |  | Demographics | Lead Demographics | Lead Serial Number:
14 | MDC_IDC_LEAD_MFG | 1.0.0 | 67 | 720963 | Y | RE |  |  | Implantable Lead Manufacturer | The manufacturer of the lead. | manufacturer ¦ lead ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | BSX | _MDC_IDC_ENUM_MFG | Demographics | Lead Demographics | Lead Manufacturer:
15 | MDC_IDC_LEAD_IMPLANT_DT | 1.0.0 | 68 | 720964 |  |  |  |  | Implantable Lead Implant Date | The implant date of the lead. | implant date ¦ lead ¦ idc ¦ mdc | Metric | DateTime | N/A |  |  |  |  | 20090121 |  | Demographics | Lead Demographics | Lead Implant Date:
16 | MDC_IDC_LEAD_POLARITY_TYPE | 1.0.0 | 69 | 720965 |  |  |  |  | Implantable Lead Polarity Type | The number of electrodes on the lead. | polarity type ¦ lead ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | BI | _MDC_IDC_ENUM_LEAD_POLARITY_TYPE | Demographics | Lead Demographics | Lead Polarity Type:
17 | MDC_IDC_LEAD_LOCATION | 1.0.0 | 70 | 720966 |  |  |  |  | Implantable Lead Location | The fixation location of the lead, usually indicating a chamber of the heart. | location ¦ lead ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | RV | _MDC_IDC_ENUM_LEAD_LOCATION_CHAMBER | Demographics | Lead Demographics | Lead Location Chamber:
18 | MDC_IDC_LEAD_LOCATION_DETAIL_1 | 1.0.0 | 71 | 720967 |  |  |  |  | Implantable Lead Location Detail 1 | The first word of a combination of words describing further details of the location. | detail 1 ¦ location ¦ lead ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | Apex | _MDC_IDC_ENUM_LEAD_LOCATION_DETAIL | Demographics | Lead Demographics | Lead Location Detail:
19 | MDC_IDC_LEAD_LOCATION_DETAIL_2 | 1.0.0 | 72 | 720968 |  |  |  |  | Implantable Lead Location Detail 2 | The second word of a combination of words describing further details of the location. | detail 2 ¦ location ¦ lead ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | not used in this example | _MDC_IDC_ENUM_LEAD_LOCATION_DETAIL | Demographics | Lead Demographics | Lead Location Detail:
20 | MDC_IDC_LEAD_LOCATION_DETAIL_3 | 1.0.0 | 73 | 720969 |  |  |  |  | Implantable Lead Location Detail 3 | The third word of a combination of words describing further details of the location. | detail 3 ¦ location ¦ lead ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | not used in this example | _MDC_IDC_ENUM_LEAD_LOCATION_DETAIL | Demographics | Lead Demographics | Lead Location Detail:
21 | MDC_IDC_LEAD_CONNECTION_STATUS | 1.0.0 | 74 | 720970 |  |  |  |  | Implantable Lead Connection Status | The physical connection status of the lead, either connected or abandoned. | connection status ¦ lead ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | Connected | _MDC_IDC_ENUM_LEAD_STATUS | Demographics | Lead Demographics | Lead Connection Status:
22 | MDC_IDC_LEAD_SPECIAL_FUNCTION | 1.0.0 | 75 | 720971 |  |  |  |  | Implantable Lead Special Function | A description of any special attribute or function of the lead. | special function ¦ lead ¦ idc ¦ mdc | Metric | String | 20 |  |  |  |  | "incl. p-sensor" |  | Demographics | Lead Demographics | Lead Special Function:
23 | MDC_IDC_SESS |  | 128 |  |  |  | 1..1 |  | Interrogation Session | The group of session attributes. |  | Channel | GROUP
24 | MDC_IDC_SESS_DTM | 1.0.0 | 129 | 721025 |  | R |  |  | Date Time Interrogation Session | The date and time of the in-clinic or remote interrogation session. | date time ¦ session ¦ idc ¦ mdc | Metric | DateTime | N/A |  |  |  |  | 20090525095530 |  | Demographics | Session | Interrogation Date:
25 | MDC_IDC_SESS_TYPE | 1.0.0 | 130 | 721026 |  | R |  |  | Type Interrogation Session | The type of device interaction that generated the current data set. | type ¦ session ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | RemoteDeviceInitiated | _MDC_IDC_ENUM_SESS_TYPE | Demographics | Session | Interrogation Type:
26 | MDC_IDC_SESS_REPROGRAMMED | 1.0.0 | 131 | 721027 |  |  |  |  | Re-programmed During Session | The indication of whether the device was re-programmed during the session. | reprogrammed ¦ session ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | NO | _MDC_IDC_ENUM_SESS_REPROGRAMMED | Demographics | Session | Reprogrammed:
27 | MDC_IDC_SESS_DTM_PREVIOUS | 1.0.0 | 132 | 721028 |  |  |  |  | Date Time Previous Interrogation Session | The date and time of a previous in-clinic or remote interrogation. | previous date time ¦ session ¦ idc ¦ mdc | Metric | DateTime | N/A |  |  |  |  | 20090515095530 |  | Demographics | Session | Previous Interrogation Date:
28 | MDC_IDC_SESS_TYPE_PREVIOUS | 1.0.0 | 133 | 721029 |  |  |  |  | Type Previous Interrogation Session | The type of device interaction of the previous interrogation. | previous type ¦ session ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | InClinic | _MDC_IDC_ENUM_SESS_TYPE | Demographics | Session | Previous Interrogation Type:
29 | MDC_IDC_SESS_REPROGRAMMED_PREVIOUS | 1.0.0 | 134 | 721030 |  |  |  |  | Re-programmed During Previous Session | The indication of whether the device was re-programmed along with the previous interrogation. | reprogrammed previous ¦ session ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | YES | _MDC_IDC_ENUM_SESS_REPROGRAMMED | Demographics | Session | Previous Reprogrammed:
30 | MDC_IDC_SESS_CLINICIAN_NAME | 1.0.0 | 135 | 721031 |  |  |  |  | Clinician Name | The name of the clinician that is responsible for the examination. | clinician name ¦ session ¦ idc ¦ mdc | Metric | String | 50 |  |  |  |  | "Mrs. Monitor" |  | Demographics | Session | Clinician Name:
31 | MDC_IDC_SESS_CLINICIAN_CONTACT_INFORMATION | 1.0.0 | 136 | 721032 |  |  |  |  | Clinician Contact Information | The contact information for the responsible clinician. | clinician contact information ¦ session ¦ idc ¦ mdc | Metric | String | 100 |  |  |  |  | "Phone: 123456789" |  | Demographics | Session | Clinician Contact:
32 | MDC_IDC_SESS_CLINIC_NAME | 1.0.0 | 137 | 721033 |  |  |  |  | Clinic Name | The name of the clinic where the examination took place. | clinic name ¦ session ¦ idc ¦ mdc | Metric | String | 50 |  |  |  |  | "Super FU clinic" |  | Demographics | Session | Clinic Name:
33 | MDC_IDC_MSMT |  | 192 |  |  |  | 0..1 |  | Measurements | The group of measurement attributes. |  | Channel | GROUP
34 | MDC_IDC_MSMT_DTM_[STRTEND] | 1.0.0 | 256 | 721152 |  |  |  |  | Measurement Date Time | The date and time of a discrete or a group of measurements. | date time_[STRTEND] ¦ measurement ¦ idc ¦ mdc | Metric | DateTime | N/A |  | STRTEND |  |  | 20090525095010
35 | MDC_IDC_MSMT_BATTERY | [5 empty] | 0..N |  | Battery Measurements | The group of battery measurement attributes. |  | Channel | GROUP
36 | MDC_IDC_MSMT_BATTERY_DTM | 1.0.0 | 320 | 721216 |  |  |  |  | Battery Date Time of Measurements | The date and time of the battery measurements. | date time ¦ battery ¦ measurement ¦ idc ¦ mdc | Metric | DateTime | N/A |  |  |  |  | 20090525095010 |  | Status / Measurements | Battery | Battery Measurement Date:
37 | MDC_IDC_MSMT_BATTERY_STATUS | 1.0.0 | 384 | 721280 |  |  |  |  | Battery Status | The different stages of battery depletion. | status ¦ battery ¦ measurement ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | BOS | _MDC_IDC_ENUM_BATTERY_STATUS | Status / Measurements | Battery | Battery Status:
38 | MDC_IDC_MSMT_BATTERY_VOLTAGE | 1.0.0 | 448 | 721344 |  |  |  |  | Battery Voltage | The measured battery voltage. | voltage ¦ battery ¦ measurement ¦ idc ¦ mdc | Metric | Numeric | F6.2 | V |  |  |  | 6.2 |  | Status / Measurements | Battery | Battery Voltage:
39 | MDC_IDC_MSMT_BATTERY_IMPEDANCE | 1.0.0 | 512 | 721408 |  |  |  |  | Battery Impedance | The measured battery impedance. | impedance ¦ battery ¦ measurement ¦ idc ¦ mdc | Metric | Numeric | F6.0 | Ohm |  |  |  | 2450 |  | Status / Measurements | Battery | Battery Impedance:
40 | MDC_IDC_MSMT_BATTERY_REMAINING_LONGEVITY | 1.0.0 | 576 | 721472 |  |  |  |  | Battery Remaining Longevity | The estimated amount of battery capacity remaining in months. | remaining longevity ¦ battery ¦ measurement ¦ idc ¦ mdc | Metric | Numeric | F6.0 | mo |  |  |  | 38 |  | Status / Measurements | Battery | Battery Remaining Pct:
41 | MDC_IDC_MSMT_BATTERY_REMAINING_PERCENTAGE | 1.0.0 | 640 | 721536 |  |  |  |  | Battery Remaining Percentage | The estimated amount of battery capacity remaining in percent. | remaining percentage ¦ battery ¦ measurement ¦ idc ¦ mdc | Metric | Numeric | F3.1 | % | [5 empty] | Status / Measurements | Battery | Battery Remaining Time:
42 | MDC_IDC_MSMT_BATTERY_RRT_TRIGGER | 1.0.0 | 704 | 721600 |  |  |  |  | Battery RRT Trigger | The criteria for determining battery end of service condition. | recommended replacement time trigger ¦ battery ¦ measurement ¦ idc ¦ mdc | Metric | String | 50 |  |  |  |  | "Battery capacity < limit for 3 months" |  | Status / Measurements | Battery | RRT (ERI) Trigger:
43 | MDC_IDC_MSMT_CAP | [5 empty] | 0..N |  | Capacitor Measurements | The group of capacitor measurement attributes. |  | Channel | GROUP
44 | MDC_IDC_MSMT_CAP_CHARGE_DTM | 1.0.0 | 768 | 721664 |  |  |  |  | Capacitor Last Charge Date Time | The date and time of the capacitor charge. | last charge date time ¦ capacitor ¦ measurement ¦ idc ¦ mdc | Metric | DateTime | N/A |  |  |  |  | 20090501095010 |  | Status / Measurements | Capacitor | Charge Date:
45 | MDC_IDC_MSMT_CAP_CHARGE_TIME | 1.0.0 | 832 | 721728 |  |  |  |  | Capacitor Charge Time | The duration in seconds of the capacitor charge. | charge time ¦ capacitor ¦ measurement ¦ idc ¦ mdc | Metric | Numeric | F6.1 | s |  |  |  | 9.4 |  | Status / Measurements | Capacitor | Charge Time:
46 | MDC_IDC_MSMT_CAP_CHARGE_ENERGY | 1.0.0 | 896 | 721792 |  |  |  |  | Capacitor Charge Energy | The amount of energy stored after the capacitor charge. | charge energy ¦ capacitor ¦ measurement ¦ idc ¦ mdc | Metric | Numeric | F4.0 | J |  |  |  | 30 |  | Status / Measurements | Capacitor | Charge Energy:
47 | MDC_IDC_MSMT_CAP_CHARGE_TYPE | 1.0.0 | 960 | 721856 |  |  |  |  | Capacitor Charge Type | The type of the capacitor charge. | charge type ¦ capacitor ¦ measurement ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | Reformation | _MDC_IDC_ENUM_CHARGE_TYPE | Status / Measurements | Capacitor | Charge Type:
48 | MDC_IDC_MSMT_LEADCHNL_[CHAMBER] | [5 empty] | 0..1 |  | Lead Channel Measurements | The group of measurement attributes per a given chamber. |  | Channel | GROUP |  |  | CHAMBER
49 | MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_DTM_[STRTEND] | 1.0.0 | 1024 | 721920 |  |  |  |  | Lead Channel Measurements Date and Time | The date and time of the lead channel measurements. | date time_[STRTEND] ¦ lead channel_[CHAMBER] ¦ measurement ¦ idc ¦ mdc | Metric | DateTime | N/A |  | STRTEND |  |  | 20090501095010 |  | Status / Measurements | Lead Channel Measurements / Status | Lead Channel Measurement Date:
50 | MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_LEAD_CHANNEL_STATUS | 1.0.0 | 1088 | 721984 |  |  |  |  | Lead Channel Status | The indication whether to check the lead. | lead channel status ¦ lead channel_[CHAMBER] ¦ measurement ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | CheckLead | _MDC_IDC_ENUM_CHANNEL_STATUS | Status / Measurements | Lead Channel Measurements / Status | Lead Channel Status:
51 | MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_SENSING | [5 empty] | 0..1 |  | Lead Channel Sensing Measurements | The group of sensing measurement attributes per a given chamber. |  | Channel | GROUP
52 | MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_SENSING_INTR_AMPL_[MMM] | 1.0.0 | 1152 | 722048 |  |  |  |  | Lead Channel Sensing Intrinsic Amplitude | The measured amplitude of the intrinsic cardiac signal. | sensing intrinsic amplitude_[MMM] ¦ lead channel_[CHAMBER] ¦ measurement ¦ idc ¦ mdc | Metric | Numeric | F3.2 | mV | MMM |  |  | 2.8 |  | Status / Measurements | Lead Channel Measurements / Status | Intrinsic Amplitude (mean, min, max):
53 | MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_SENSING_POLARITY | 1.0.0 | 1216 | 722112 |  |  |  |  | Lead Channel Sensing Polarity | The type of polarity used for the intrinsic amplitude measurement. | sensing polarity ¦ lead channel_[CHAMBER] ¦ measurement ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | BI | _MDC_IDC_ENUM_POLARITY | Status / Measurements | Lead Channel Measurements / Status | Sensing Polarity:
54 | MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_PACING_THRESHOLD | [5 empty] | 0..1 |  | Lead Channel Pacing Threshold Measurements | The group of pacing threshold measurement attributes per a given chamber. |  | Channel | GROUP
55 | MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_PACING_THRESHOLD_AMPLITUDE | 1.0.0 | 1280 | 722176 |  |  |  |  | Lead Channel Pacing Threshold Amplitude | The minimum pulse amplitude needed for pacing capture for a given pulse width. | amplitude ¦ pacing threshold ¦ lead channel_[CHAMBER] ¦ measurement ¦ idc ¦ mdc | Metric | Numeric | F3.3 | V |  |  |  | 0.8 |  | Status / Measurements | Lead Channel Measurements / Status | Pacing Amplitude:
56 | MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_PACING_THRESHOLD_PULSEWIDTH | 1.0.0 | 1344 | 722240 |  |  |  |  | Lead Channel Pacing Threshold Pulse Width | The pulse width that was applied when determining the pacing threshold amplitude. | pulse width ¦ pacing threshold ¦ lead channel_[CHAMBER] ¦ measurement ¦ idc ¦ mdc | Metric | Numeric | F3.2 | ms |  |  |  | 0.5 |  | Status / Measurements | Lead Channel Measurements / Status | Pacing Pulse Width:
57 | MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_PACING_THRESHOLD_MEASUREMENT_METHOD | 1.0.0 | 1408 | 722304 |  |  |  |  | Lead Channel Pacing Threshold Measurement Method | The method that was used to obtain the pacing threshold. | measurement method ¦ pacing threshold ¦ lead channel_[CHAMBER] ¦ measurement ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | DeviceAutomatic | _MDC_IDC_ENUM_MEASUREMENT_METHOD | Status / Measurements | Lead Channel Measurements / Status | Threshold Measurement Method:
58 | MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_PACING_THRESHOLD_POLARITY | 1.0.0 | 1472 | 722368 |  |  |  |  | Lead Channel Pacing Threshold Polarity | The type of lead polarity that was used for the pacing threshold measurement. | polarity ¦ pacing threshold ¦ lead channel_[CHAMBER] ¦ measurement ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | BI | _MDC_IDC_ENUM_POLARITY | Status / Measurements | Lead Channel Measurements / Status | Threshold Polarity:
59 | MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_IMPEDANCE | [5 empty] | 0..1 |  | Lead Channel Impedance Measurements | The group of impedance measurement attributes per a given chamber. |  | Channel | GROUP
60 | MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_IMPEDANCE_VALUE | 1.0.0 | 1536 | 722432 |  |  |  |  | Lead Channel Impedance Value | The sum of the impedances of the lead wires and the electrode-myocardial interface. | value ¦ impedance ¦ lead channel_[CHAMBER] ¦ measurement ¦ idc ¦ mdc | Metric | Numeric | F6.0 | Ohm |  | Because the term is the child of impedance we added "value" to the reference id to keep it unique. |  | 530 |  | Status / Measurements | Lead Channel Measurements / Status | Impedance Measurement:
61 | MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_IMPEDANCE_POLARITY | 1.0.0 | 1600 | 722496 |  |  |  |  | Lead Channel Impedance Polarity | The type of lead polarity used for measuring the lead impedance. | polarity ¦ impedance ¦ lead channel_[CHAMBER] ¦ measurement ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | BI | _MDC_IDC_ENUM_POLARITY | Status / Measurements | Lead Channel Measurements / Status | Impedance Polarity:
62 | MDC_IDC_MSMT_LEADHVCHNL | [5 empty] | 0..N |  | Lead High Voltage Channel Measurements | The group of high voltage channel measurement attributes. |  | Channel | GROUP
63 | MDC_IDC_MSMT_LEADHVCHNL_DTM_[STRTEND] | 1.0.0 | 1664 | 722560 |  |  |  |  | Lead High Voltage Channel Date Time | The date and time of the high voltage lead channel measurements. | date time_[STRTEND] ¦ lead high voltage channel ¦ measurement ¦ idc ¦ mdc | Metric | DateTime | N/A |  | STRTEND |  |  | 20090501095010 |  | Status / Measurements | Shock Lead Configuration Measurements | Shock Lead Measurement Date:
64 | MDC_IDC_MSMT_LEADHVCHNL_IMPEDANCE | 1.0.0 | 1728 | 722624 |  |  |  |  | Lead High Voltage Channel Impedance | The sum of the impedances of the shock lead wires and the electrode-myocardial interface. | impedance ¦ lead high voltage channel ¦ measurement ¦ idc ¦ mdc | Metric | Numeric | F6.0 | Ohm |  |  |  | 47 |  | Status / Measurements | Shock Lead Configuration Measurements | Shock Lead Impedance:
65 | MDC_IDC_MSMT_LEADHVCHNL_MEASUREMENT_TYPE | 1.0.0 | 1792 | 722688 |  |  |  |  | Lead High Voltage Channel Measurement Type | The electric pulse type used for measuring the shock lead impedance. | measurement type ¦ lead high voltage channel ¦ measurement ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | LowVoltage | _MDC_IDC_ENUM_HVCHNL_MEASUREMENT_TYPE | Status / Measurements | Shock Lead Configuration Measurements | Shock Lead Measurement Type:
66 | MDC_IDC_MSMT_LEADHVCHNL_STATUS | 1.0.0 | 1856 | 722752 |  |  |  |  | Lead High Voltage Channel Status | The indication whether to check the high voltage lead. | status ¦ lead high voltage channel ¦ measurement ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | Null | _MDC_IDC_ENUM_CHANNEL_STATUS | Status / Measurements | Shock Lead Configuration Measurements | Shock Lead Status:
67 | MDC_IDC_SET |  | 8384 |  |  |  | 0..1 |  | Settings | The group of settings. |  | Channel | GROUP
68 | MDC_IDC_SET_CRT | [5 empty] | 0..1 |  | CRT Settings | The group of CRT settings. |  | Channel | GROUP
69 | MDC_IDC_SET_CRT_LVRV_DELAY | 1.0.0 | 8448 | 729344 |  |  |  |  | CRT LV-RV Delay | The time between the LV and RV pacing pulses (positive values indicate LV is first). | cardiac resynchronization therapy left ventricle right ventricle delay ¦ setting ¦ idc ¦ mdc | Metric | Numeric | F4.0 | ms |  |  |  | 10 |  | Device Settings | CRT Settings | CRT LV-RV Delay:
70 | MDC_IDC_SET_CRT_PACED_CHAMBERS | 1.0.0 | 8512 | 729408 |  |  |  |  | Ventricular chambers paced during CRT pacing. | The ventricular chambers paced during CRT pacing. | cardiac resynchronization therapy paced chambers ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | BiV | _MDC_IDC_ENUM_CRT_PACED_CHAMBERS | Device Settings | CRT Settings | CRT Paced Chambers:
71 | MDC_IDC_SET_MAGNET | [5 empty] | 0..1 |  | Magnet Settings | The group of magnet settings. |  | Channel | GROUP
72 | MDC_IDC_SET_MAGNET_RESP | 1.0.0 | 8576 | 729472 |  |  |  |  | Magnet Response | A description of how the device responds to a magnet. | magnet responce ¦ setting ¦ idc ¦ mdc | Metric | String | 100 |  |  | We should limit the length of strings in the profile. |  | "Therapies inactive, VVIR" |  | Device Settings | Magnet Settings | Magnet Mode:
73 | MDC_IDC_SET_LEADCHNL_[CHAMBER] | [5 empty] | 0..1 |  | Lead Channel Settings | The group of settings per a given chamber. |  | Channel | GROUP |  |  | CHAMBER
74 | MDC_IDC_SET_LEADCHNL_[CHAMBER]_SENSING | [5 empty] | 0..1 |  | Lead Channel Settings Sensing | The group of sensing settings per a given chamber. |  | Channel | GROUP
75 | MDC_IDC_SET_LEADCHNL_[CHAMBER]_SENSING_SENSITIVITY | 1.0.0 | 8640 | 729536 |  |  |  |  | Lead Channel Setting Sensing Sensitivity | The smallest electrical signal programmed to be detected by the device's sensing circuitry. | sensitivity ¦ sensing ¦ lead channel_[CHAMBER] ¦ setting ¦ idc ¦ mdc | Metric | Numeric | F3.2 | mV |  |  |  | 0.8 |  | Device Settings | Lead Channel Settings | Sensitivity:
76 | MDC_IDC_SET_LEADCHNL_[CHAMBER]_SENSING_POLARITY | 1.0.0 | 8704 | 729600 |  |  |  |  | Lead Channel Setting Sensing Polarity | The indication of unipolar or bipolar configuration for sensing. | polarity ¦ sensing ¦ lead channel_[CHAMBER] ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | BI | _MDC_IDC_ENUM_POLARITY | Device Settings | Lead Channel Settings | Sensing Polarity:
77 | MDC_IDC_SET_LEADCHNL_[CHAMBER]_SENSING_ANODE_LOCATION_[1..3] | 1.0.0 | 8768 | 729664 |  |  |  | CC1 | Lead Channel Setting Sensing Anode Location | The location of the electrodes that define the sensing vector's anode, noting that sensing is polarity neutral. An anode may consist of more than one electrode. | annode location_[1..3] ¦ sensing ¦ lead channel_[CHAMBER] ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  | 1..3 |  |  | RV | _MDC_IDC_ENUM_ELECTRODE_LOCATION | Device Settings | Lead Channel Settings | Sensing Vector:
78 | MDC_IDC_SET_LEADCHNL_[CHAMBER]_SENSING_ANODE_ELECTRODE_[1..3] | 1.0.0 | 8832 | 729728 |  |  |  | CC1 | Lead Channel Setting Sensing Anode Terminal | The type of the electrodes that define the sensing vector's anode, noting that sensing is polarity neutral. An anode may consist of more than one electrode. | annode electrode_[1..3] ¦ sensing ¦ lead channel_[CHAMBER] ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  | 1..3 |  |  | Tip | _MDC_IDC_ENUM_ELECTRODE_NAME | Device Settings | Lead Channel Settings | Sensing Vector:
79 | MDC_IDC_SET_LEADCHNL_[CHAMBER]_SENSING_CATHODE_LOCATION_[1..3] | 1.0.0 | 8896 | 729792 |  |  |  | CC1 | Lead Channel Setting Sensing Cathode Location | The location of the electrodes that define the sensing vector's cathode, noting that sensing is polarity neutral. A cathode may consist of more than one electrode. | cathode location_[1..3] ¦ sensing ¦ lead channel_[CHAMBER] ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  | 1..3 |  |  | RV | _MDC_IDC_ENUM_ELECTRODE_LOCATION | Device Settings | Lead Channel Settings | Sensing Vector:
80 | MDC_IDC_SET_LEADCHNL_[CHAMBER]_SENSING_CATHODE_ELECTRODE_[1..3] | 1.0.0 | 8960 | 729856 |  |  |  | CC1 | Lead Channel Setting Sensing Cathode Terminal | The type of the electrodes that define the sensing vector's cathode, noting that sensing is polarity neutral. A cathode may consist of more than one electrode. | cathode electrode_[1..3] ¦ sensing ¦ lead channel_[CHAMBER] ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  | 1..3 |  |  | Ring | _MDC_IDC_ENUM_ELECTRODE_NAME | Device Settings | Lead Channel Settings | Sensing Vector:
81 | MDC_IDC_SET_LEADCHNL_[CHAMBER]_SENSING_ADAPTATION_MODE | 1.0.0 | 9024 | 729920 |  |  |  |  | Lead Channel Setting Sensing Adaptation Mode | Specifies whether the sensitivity is fixed or adapting. | adaptation mode ¦ sensing ¦ lead channel_[CHAMBER] ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | AdaptiveSensing | _MDC_IDC_ENUM_SENSING_ADAPTATION_MODE | Device Settings | Lead Channel Settings | Sensing Adaptation Mode:
82 | MDC_IDC_SET_LEADCHNL_[CHAMBER]_PACING | [5 empty] | 0..1 |  | Lead Channel Settings Pacing | The group of pacing settings per a given chamber. |  | Channel | GROUP
83 | MDC_IDC_SET_LEADCHNL_[CHAMBER]_PACING_AMPLITUDE | 1.0.0 | 9088 | 729984 |  |  |  |  | Lead Channel Setting Pacing Amplitude | The pacing output amplitude. | amplitude ¦ pacing ¦ lead channel_[CHAMBER] ¦ setting ¦ idc ¦ mdc | Metric | Numeric | F3.3 | V |  |  |  | 3.2 |  | Device Settings | Lead Channel Settings | Pacing Output:
84 | MDC_IDC_SET_LEADCHNL_[CHAMBER]_PACING_PULSEWIDTH | 1.0.0 | 9152 | 730048 |  |  |  |  | Lead Channel Setting Pacing Pulse Width | The pacing output pulse width. | pulse width ¦ pacing ¦ lead channel_[CHAMBER] ¦ setting ¦ idc ¦ mdc | Metric | Numeric | F3.2 | ms |  |  |  | 0.5 |  | Device Settings | Lead Channel Settings | Pacing Pulse Width:
85 | MDC_IDC_SET_LEADCHNL_[CHAMBER]_PACING_POLARITY | 1.0.0 | 9216 | 730112 |  |  |  |  | Lead Channel Setting Pacing Polarity | The indication of unipolar or bipolar configuration for pacing. | polarity ¦ pacing ¦ lead channel_[CHAMBER] ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | BI | _MDC_IDC_ENUM_POLARITY | Device Settings | Lead Channel Settings | Pacing Polarity:
86 | MDC_IDC_SET_LEADCHNL_[CHAMBER]_PACING_ANODE_LOCATION_[1..3] | 1.0.0 | 9280 | 730176 |  |  |  | CC1 | Lead Channel Setting Pacing Anode Location | The location of the electrodes that define the pacing vector's anode. An anode may consist of more than one electrode. | annode location_[1..3] ¦ pacing ¦ lead channel_[CHAMBER] ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  | 1..3 |  |  | RV | _MDC_IDC_ENUM_ELECTRODE_LOCATION | Device Settings | Lead Channel Settings | Pacing Vector:
87 | MDC_IDC_SET_LEADCHNL_[CHAMBER]_PACING_ANODE_ELECTRODE_[1..3] | 1.0.0 | 9344 | 730240 |  |  |  | CC1 | Lead Channel Setting Pacing Anode Terminal | The type of the electrodes that define the pacing vector's anode. An anode may consist of more than one electrode. | annode electrode_[1..3] ¦ pacing ¦ lead channel_[CHAMBER] ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  | 1..3 |  |  | Tip | _MDC_IDC_ENUM_ELECTRODE_NAME | Device Settings | Lead Channel Settings | Pacing Vector:
88 | MDC_IDC_SET_LEADCHNL_[CHAMBER]_PACING_CATHODE_LOCATION_[1..3] | 1.0.0 | 9408 | 730304 |  |  |  | CC1 | Lead Channel Setting Pacing Cathode Location | The location of the electrodes that define the pacing vector's cathode. A cathode may consist of more than one electrode. | cathode location_[1..3] ¦ pacing ¦ lead channel_[CHAMBER] ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  | 1..3 |  |  | RV | _MDC_IDC_ENUM_ELECTRODE_LOCATION | Device Settings | Lead Channel Settings | Pacing Vector:
89 | MDC_IDC_SET_LEADCHNL_[CHAMBER]_PACING_CATHODE_ELECTRODE_[1..3] | 1.0.0 | 9472 | 730368 |  |  |  | CC1 | Lead Channel Setting Pacing Cathode Terminal | The type of the electrodes that define the pacing vector's cathode. A cathode may consist of more than one electrode. | cathode electrode_[1..3] ¦ pacing ¦ lead channel_[CHAMBER] ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  | 1..3 |  |  | Ring | _MDC_IDC_ENUM_ELECTRODE_NAME | Device Settings | Lead Channel Settings | Pacing Vector:
90 | MDC_IDC_SET_LEADCHNL_[CHAMBER]_PACING_CAPTURE_MODE | 1.0.0 | 9536 | 730432 |  |  |  |  | Lead Channel Setting Pacing Capture Mode | The method the device uses for managing pacing capture. | capture mode ¦ pacing ¦ lead channel_[CHAMBER] ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | AdaptiveCapture | _MDC_IDC_ENUM_PACING_CAPTURE_MODE | Device Settings | Lead Channel Settings | Pacing Capture Mode:
91 | MDC_IDC_SET_LEADHVCHNL | [5 empty] | 0..1 |  | Lead High Voltage Channel Setting | The group of high voltage settings. |  | Channel | GROUP
92 | MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR | [5 empty] | 0..1 |  | Lead High Voltage Channel Setting Shock Vector | The group of high voltage shock vector settings. |  | Channel | GROUP
93 | MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_LOCATION_[1..3] | 1.0.0 | 9600 | 730496 |  |  |  | CC1 | Lead High Voltage Channel Setting Shock Vector Anode Location | The location of the anode electrode that defines the shocking vector for the lead. An anode may consist of more than one electrode. | annode location_[1..3] ¦ vector ¦ shock ¦ lead high voltage channel ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  | 1..3 |  |  | RV | _MDC_IDC_ENUM_ELECTRODE_LOCATION | Device Settings | Shock Vector Settings | Shock Vector:
94 | MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_ELECTRODE_[1..3] | 1.0.0 | 9664 | 730560 |  |  |  | CC1 | Lead High Voltage Channel Setting Shock Vector Anode Terminal | The anode electrode terminal that defines the shocking vector for the lead. An anode may consist of more than one electrode. | annode electrode_[1..3] ¦ vector ¦ shock ¦ lead high voltage channel ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  | 1..3 |  |  | Coil | _MDC_IDC_ENUM_ELECTRODE_NAME | Device Settings | Shock Vector Settings | Shock Vector:
95 | MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_CATHODE_LOCATION_[1..3] | 1.0.0 | 9728 | 730624 |  |  |  | CC1 | Lead High Voltage Channel Setting Shock Vector Cathode Location | The location of the cathode electrode that defines the shocking vector for the lead. A cathode may consist of more than one electrode. | cathode location_[1..3] ¦ vector ¦ shock ¦ lead high voltage channel ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  | 1..3 |  |  | N/A | _MDC_IDC_ENUM_ELECTRODE_LOCATION | Device Settings | Shock Vector Settings | Shock Vector:
96 | MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_CATHODE_ELECTRODE_[1..3] | 1.0.0 | 9792 | 730688 |  |  |  | CC1 | Lead High Voltage Channel Setting Shock Vector Cathode Terminal | The cathode electrode terminal that defines the shocking vector for the lead. A cathode may consist of more than one electrode. | cathode electrode_[1..3] ¦ vector ¦ shock ¦ lead high voltage channel ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  | 1..3 |  |  | Can | _MDC_IDC_ENUM_ELECTRODE_NAME | Device Settings | Shock Vector Settings | Shock Vector:
97 | MDC_IDC_SET_BRADY | [5 empty] | 0..1 |  | Brady Settings | The group of brady therapy settings. |  | Channel | GROUP
98 | MDC_IDC_SET_BRADY_MODE | 1.0.0 | 9856 | 730752 |  |  |  |  | Brady Setting Mode (NBG Code) | The brady pacing mode according to the NBG standard. | mode ¦ brady ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | DDDR | _MDC_IDC_ENUM_BRADY_MODE | Device Settings | Brady Settings | Brady Mode:
99 | MDC_IDC_SET_BRADY_VENDOR_MODE | 1.0.0 | 9920 | 730816 |  |  |  |  | Brady Setting Mode (Vendor Specific) | The brady pacing mode as defined by vendor specific codes. | mode ¦ brady ¦ setting ¦ idc ¦ mdc | Metric | String | 20 |  |  |  |  | DDI-CLS |  | Device Settings | Brady Settings | Brady Mode:
100 | MDC_IDC_SET_BRADY_LOWRATE | 1.0.0 | 9984 | 730880 |  |  |  |  | Brady Setting Lower Rate Limit | The rate at which the device paces in the absence of a cardiac rhythm faster than the lower rate and without influence from features that can affect the pacing rate. | lowrate ¦ brady ¦ setting ¦ idc ¦ mdc | Metric | Numeric | F3.0 | {beats}/min |  |  |  | 60 |  | Device Settings | Brady Settings | Lower Rate:
101 | MDC_IDC_SET_BRADY_HYSTRATE | 1.0.0 | 10048 | 730944 |  |  |  |  | Brady Setting Hysteresis Rate | The lower rate for the hysteresis feature. | hystrate ¦ brady ¦ setting ¦ idc ¦ mdc | Metric | Numeric | F3.0 | {beats}/min |  |  |  | 55 |  | Device Settings | Brady Settings | Hysteresis Rate:
102 | MDC_IDC_SET_BRADY_NIGHT_RATE | 1.0.0 | 10112 | 731008 |  |  |  |  | Brady Setting Night Rate | The lower rate for the night rate feature. | night rate ¦ brady ¦ setting ¦ idc ¦ mdc | Metric | Numeric | F3.0 | {beats}/min |  |  |  | 50 |  | Device Settings | Brady Settings | Night Rate:
103 | MDC_IDC_SET_BRADY_SENSOR_TYPE | 1.0.0 | 10176 | 731072 |  |  |  |  | Brady Setting Sensor Type | The type of sensor applied for rate adaptive pacing. | sensor type ¦ brady ¦ setting ¦ idc ¦ mdc | Metric | String | 20 |  |  |  |  | "Acceleration and CLS" |  | Device Settings | Brady Settings | Sensor Type:
104 | MDC_IDC_SET_BRADY_MAX_TRACKING_RATE | 1.0.0 | 10240 | 731136 |  |  |  |  | Brady Setting Maximum Tracking Rate | The fastest atrial rate at which ventricular pacing occurs with 1:1 synchrony. | maximum tracking rate ¦ brady ¦ setting ¦ idc ¦ mdc | Metric | Numeric | F3.0 | {beats}/min |  |  |  | 140 |  | Device Settings | Brady Settings | Max Tracking Rate:
105 | MDC_IDC_SET_BRADY_MAX_SENSOR_RATE | 1.0.0 | 10304 | 731200 |  |  |  |  | Brady Setting Maximum Sensor Rate | The fastest sensor-driven pacing rate that can be achieved in a rate-adaptive pacing system. | maximum sensor rate ¦ brady ¦ setting ¦ idc ¦ mdc | Metric | Numeric | F3.0 | {beats}/min |  |  |  | 120 |  | Device Settings | Brady Settings | Max Sensor Rate:
106 | MDC_IDC_SET_BRADY_SAV_DELAY_[HIGHLOW] | 1.0.0 | 10368 | 731264 |  |  |  |  | Brady Setting SAV Delay | The interval from an intrinsic (sensed) P wave to a paced ventricular event. | sav delay_[HIGHLOW] ¦ brady ¦ setting ¦ idc ¦ mdc | Metric | Numeric | F4.0 | ms | HIGHLOW | High means the SAV delay for the higher heart rate. |  | 120 |  | Device Settings | Brady Settings | SAV Delay:
107 | MDC_IDC_SET_BRADY_PAV_DELAY_[HIGHLOW] | 1.0.0 | 10432 | 731328 |  |  |  |  | Brady Setting PAV Delay | The interval from a paced atrial event to a paced ventricular event. | pav delay_[HIGHLOW] ¦ brady ¦ setting ¦ idc ¦ mdc | Metric | Numeric | F4.0 | ms | HIGHLOW | High means the PAV delay for the higher heart rate. |  | 150 |  | Device Settings | Brady Settings | PAV Delay:
108 | MDC_IDC_SET_BRADY_AT_MODE_SWITCH_MODE | 1.0.0 | 10496 | 731392 |  |  |  |  | Brady Setting AT Mode Switch Mode | The atrial tachycardia mode switch mode. | atrial tachy mode switch mode ¦ brady ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | DDIR | _MDC_IDC_ENUM_BRADY_MODE | Device Settings | Atrial Tachyarrhythmia Settings | Mode Switch Mode:
109 | MDC_IDC_SET_BRADY_AT_MODE_SWITCH_RATE | 1.0.0 | 10560 | 731456 |  |  |  |  | Brady Setting AT Mode Switch Rate | The atrial tachycardia mode switch rate. | atrail tachy mode switch rate ¦ brady ¦ setting ¦ idc ¦ mdc | Metric | Numeric | F3.0 | {beats}/min |  |  |  | 180 |  | Device Settings | Atrial Tachyarrhythmia Settings | Mode Switch Rate:
110 | MDC_IDC_SET_TACHYTHERAPY | [5 empty] | 0..1 |  | Tachy Therapy Settings | The group of high level tachy therapy settings. |  | Channel | GROUP
111 | MDC_IDC_SET_TACHYTHERAPY_VSTAT | 1.0.0 | 10624 | 731520 |  |  |  |  | Tachy Therapy Setting Ventricular Status | The high level programmed status for ventricular tachy therapies. | ventricular status ¦ tachytherapy ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | On | _MDC_IDC_ENUM_THERAPY_STATUS | Device Settings | Tachyarrhythmia Zone Settings | Ventricular Therapy:
112 | MDC_IDC_SET_TACHYTHERAPY_ASTAT | 1.0.0 | 10688 | 731584 |  |  |  |  | Tachy Therapy Setting Atrial Status | The high level programmed status for atrial tachy therapies. | atrial status ¦ tachytherapy ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | Off | _MDC_IDC_ENUM_THERAPY_STATUS | Device Settings | Tachyarrhythmia Zone Settings | Atrial Therapy:
113 | MDC_IDC_SET_ZONE | [5 empty] | 0..N |  | Zone Settings | The group of zone settings. |  | Channel | GROUP
114 | MDC_IDC_SET_ZONE_TYPE | 1.0.0 | 10752 | 731648 | Y | RE |  |  | Zone Setting Type Category | The general type of the zone for tachy therapy detection. | type ¦ zone ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | Epis_VT | _MDC_IDC_ENUM_ZONE_TYPE | Device Settings | Tachyarrhythmia Zone Settings | Zone:
115 | MDC_IDC_SET_ZONE_VENDOR_TYPE | 1.0.0 | 10816 | 731712 | Y | RE |  |  | Zone Setting Vendor Type Category | The vendor specific type of the zone for tachy therapy detection. | vendor type ¦ zone ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | VT1 | _MDC_IDC_ENUM_ZONE_VENDOR_TYPE | Device Settings | Tachyarrhythmia Zone Settings | Zone:
116 | MDC_IDC_SET_ZONE_STATUS | 1.0.0 | 10880 | 731776 |  |  |  |  | Zone Setting Status | The status of the zone, whether it is is active or not. | status ¦ zone ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | Null | _MDC_IDC_ENUM_ZONE_STATUS | Device Settings | Tachyarrhythmia Zone Settings | Status:
117 | MDC_IDC_SET_ZONE_DETECTION_INTERVAL | 1.0.0 | 10944 | 731840 |  |  |  |  | Zone Setting Detection Interval | The upper interval limit of the zone for tachy therapy detection. | detection interval ¦ zone ¦ setting ¦ idc ¦ mdc | Metric | Numeric | F4.0 | ms |  | The representation at implementation could be converted to rate. |  | 140 |  | Device Settings | Tachyarrhythmia Zone Settings | Limit BPM:
118 | MDC_IDC_SET_ZONE_DETECTION_BEATS_NUMERATOR | 1.0.0 | 11008 | 731904 |  |  |  | CC2 | Zone Setting Detection Beats Numerator | The numerator portion of the tachy detection ratio. | detection beats numerator ¦ zone ¦ setting ¦ idc ¦ mdc | Metric | Numeric | F3.0 | {beats} |  |  |  | 8 |  | Device Settings | Tachyarrhythmia Zone Settings | Detection X of Y
119 | MDC_IDC_SET_ZONE_DETECTION_BEATS_DENOMINATOR | 1.0.0 | 11072 | 731968 |  |  |  | CC2 | Zone Setting Detection Beats Denominator | The denominator portion of the tachy detection ratio. | detection beats denominator ¦ zone ¦ setting ¦ idc ¦ mdc | Metric | Numeric | F3.0 | {beats} |  |  |  | 12 |  | Device Settings | Tachyarrhythmia Zone Settings | Detection X of Y
120 | MDC_IDC_SET_ZONE_DETECTION_DETAILS | 1.0.0 | 11136 | 732032 |  |  |  |  | Zone Setting Detection Details | A text describing arrhythmia detection details. | detection details ¦ zone ¦ setting ¦ idc ¦ mdc | Metric | String | 50 |  |  |  |  | "Onset / Stability ON" |  | Device Settings | Tachyarrhythmia Zone Settings | Details:
121 | MDC_IDC_SET_ZONE_TYPE_ATP_[1..10] | 1.0.0 | 11200 | 732096 |  |  |  | CC3 | Zone Setting ATP Type | The type of anti-tachycardia pacing (ATP) sequences programmed per ATP type. | type anti-tachycardia pacing pulse_[1..10] ¦ zone ¦ setting ¦ idc ¦ mdc | Metric | Enum | N/A |  | 1..10 |  |  | #REF! | _MDC_IDC_ENUM_ATP_TYPE | Device Settings | Tachyarrhythmia Zone Settings | ATP:
122 | MDC_IDC_SET_ZONE_NUM_ATP_SEQS_[1..10] | 1.0.0 | 11264 | 732160 |  |  |  | CC3 | Zone Setting Number of ATP Sequences | The number of anti-tachycardia pacing pulse (ATP) sequences programmed per ATP type. | number of anti-tachycardia pacing pulse sequences_[1..10] ¦ zone ¦ setting ¦ idc ¦ mdc | Metric | Numeric | F3.0 | {seq} | 1..10 |  |  | 3 |  | Device Settings | Tachyarrhythmia Zone Settings | ATP:
123 | MDC_IDC_SET_ZONE_SHOCK_ENERGY_[1..10] | 1.0.0 | 11328 | 732224 |  |  |  | CC4 | Zone Setting Shock Energy | The shock energy delivered at a particular sequence of a zone therapy. | shock energy_[1..10] ¦ zone ¦ setting ¦ idc ¦ mdc | Metric | Numeric | F4.1 | J | 1..10 |  |  | 30 |  | Device Settings | Tachyarrhythmia Zone Settings | Shocks:
124 | MDC_IDC_SET_ZONE_NUM_SHOCKS_[1..10] | 1.0.0 | 11392 | 732288 |  |  |  | CC4 | Zone Setting Number of Shocks | The maximum number of shocks set to be delivered at the given shock energy. | number of shocks_[1..10] ¦ zone ¦ setting ¦ idc ¦ mdc | Metric | Numeric | F3.0 | {shocks} | 1..10 |  |  | 5 |  | Device Settings | Tachyarrhythmia Zone Settings | Shocks:
125 | MDC_IDC_STAT |  | 16576 |  |  |  | 0..1 |  | Statistics | The group of statistics. |  | Channel | GROUP
126 | MDC_IDC_STAT_DTM_[STRTEND] | 1.0.0 | 16592 | 737488 |  | RE |  |  | Statistic Date Time | The date and time period for the aggregated statistics. | date time_[STRTEND] ¦ statistic ¦ idc ¦ mdc | Metric | DateTime | N/A |  | STRTEND |  |  | 20090201095010 |  | Status / Measurements
127 | MDC_IDC_STAT_HEART_RATE | [5 empty] | 0..1 |  | Statistic Heart Rate | The group of heart rate statistics. |  | Channel | GROUP
128 | MDC_IDC_STAT_HEART_RATE_DTM_[STRTEND] | 1.0.0 | 16720 | 737616 |  | RE |  | CC5 | Statistic Heart Rate Date Time | The date and time period for the heart rate statistics. | date time_[STRTEND] ¦ heart rate ¦ statistic ¦ idc ¦ mdc | Metric | DateTime | N/A |  | STRTEND |  |  | 20090201095010 |  | Status / Measurements | Brady Statistics | Heart Rate Measurement Date Time:
129 | MDC_IDC_STAT_HEART_RATE_ATRIAL_[MMM] | 1.0.0 | 16736 | 737632 |  |  |  |  | Statistic Atrial Heart Rate | The mean atrial heart rate over the specified time period. | atrial_[MMM] ¦ heart rate ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F3.0 | {beats}/min | MMM |  |  | 80 |  | Status / Measurements | Brady Statistics | Atrial Heart Rate:
130 | MDC_IDC_STAT_HEART_RATE_VENTRICULAR_[MMM] | 1.0.0 | 16752 | 737648 |  |  |  |  | Statistic Ventricular Heart Rate | The mean ventricular heart rate over the specified time period. | ventricular_[MMM] ¦ heart rate ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F3.0 | {beats}/min | MMM |  |  | 78 |  | Status / Measurements | Brady Statistics | Ventricular Heart Rate:
131 | MDC_IDC_STAT_BRADY | [5 empty] | 0..1 |  | Brady Statistics | The group of brady statistics. |  | Channel | GROUP
132 | MDC_IDC_STAT_BRADY_DTM_[STRTEND] | 1.0.0 | 16608 | 737504 |  | RE |  | CC5 | Brady Statistic Date Time | The date time period of the brady statistics. | date time_[STRTEND] ¦ brady ¦ statistic ¦ idc ¦ mdc | Metric | DateTime | N/A |  | STRTEND |  |  | 20090201095010 |  | Status / Measurements | Brady Statistics | Brady Statistics Measurement Date:
133 | MDC_IDC_STAT_BRADY_RA_PERCENT_PACED | 1.0.0 | 16624 | 737520 |  |  |  |  | Brady Statistic RA Percent Paced | The percentage of pacing events in the right atrial chamber over the specified time period. | right atrial percent paced ¦ brady ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F3.1 | % |  |  |  | 0 |  | Status / Measurements | Brady Statistics | RA Pacing:
134 | MDC_IDC_STAT_BRADY_RV_PERCENT_PACED | 1.0.0 | 16640 | 737536 |  |  |  |  | Brady Statistic RV Percent Paced | The percentage of pacing events in the right ventricular chamber over the specified time period. | right ventricle percent paced ¦ brady ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F3.1 | % |  |  |  | 4 |  | Status / Measurements | Brady Statistics | RV Pacing:
135 | MDC_IDC_STAT_BRADY_AP_VP_PERCENT | 1.0.0 | 16656 | 737552 |  |  |  |  | Brady Statistic AP VP Percent | The percentage of atrial pace - ventricular pace sequences in relationship to all AV sequences over the specified time period. | ap-vp percent ¦ brady ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F3.1 | % |  |  |  | 1 |  | Status / Measurements | Brady Statistics | AP-VP:
136 | MDC_IDC_STAT_BRADY_AS_VP_PERCENT | 1.0.0 | 16672 | 737568 |  |  |  |  | Brady Statistic AS VP Percent | The percentage of atrial sense - ventricular pace sequences in relationship to all AV sequences over the specified time period. | as-vp percent ¦ brady ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F3.1 | % |  |  |  | 3 |  | Status / Measurements | Brady Statistics | AS-VP:
137 | MDC_IDC_STAT_BRADY_AP_VS_PERCENT | 1.0.0 | 16688 | 737584 |  |  |  |  | Brady Statistic AP VS Percent | The percentage of atrial pace - ventricular sense sequences in relationship to all AV sequences over the specified time period. | ap-vs percent ¦ brady ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F3.1 | % |  |  |  | 0 |  | Status / Measurements | Brady Statistics | AP-VS:
138 | MDC_IDC_STAT_BRADY_AS_VS_PERCENT | 1.0.0 | 16704 | 737600 |  |  |  |  | Brady Statistic AS VS Percent | The percentage of atrial sense - ventricular sense sequences in relationship to all AV sequences over the specified time period. | as-vs percent ¦ brady ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F3.1 | % |  |  |  | 0 |  | Status / Measurements | Brady Statistics | AS-VS:
139 | MDC_IDC_STAT_AT | [5 empty] | 0..1 |  | Atrial Tachy Statistics | The group of atrial tachyarrhythmia statistics. |  | Channel | GROUP
140 | MDC_IDC_STAT_AT_DTM_[STRTEND] | 1.0.0 | 16768 | 737664 |  | RE |  | CC5 | Atrial Tachy Statistic Date Time | The date and time period for the AT statistics. | date time_[STRTEND] ¦ atrial tachy ¦ statistic ¦ idc ¦ mdc | Metric | DateTime | N/A |  | STRTEND |  |  | 20090201095010 |  | Status / Measurements | Atrial Tachy Statistics | Atrial Tachy Statistics Measurement Date:
141 | MDC_IDC_STAT_AT_MODE_SW_MAX_DURATION | 1.0.0 | 16784 | 737680 |  |  |  |  | Atrial Tachy Statistic Maximum Mode Switch Duration | The maximum contiguous time the device is in AT mode switch over the specified time period. | mode switch maximum duration ¦ atrial tachy ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F8.0 | s |  |  |  | 3 |  | Status / Measurements | Atrial Tachy Statistics | Max ModeSw-Epis Duration:
142 | MDC_IDC_STAT_AT_BURDEN_PERCENT | 1.0.0 | 16800 | 737696 |  |  |  |  | Atrial Tachy Statistic AT/AF Burden Percent | The percent of time in AT AF per day over the specified time period. | burden percent ¦ atrial tachy ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F3.0 | % |  |  |  | 3 |  | Status / Measurements | Atrial Tachy Statistics | AT/AF Burden per day:
143 | MDC_IDC_STAT_AT_MODE_SW_PERCENT_TIME | 1.0.0 | 16816 | 737712 |  |  |  |  | Atrial Tachy Statistic Percent Time In Mode Switch | The percentage of time the device is in AT mode switch over the specified time period. | mode switch percent of time ¦ atrial tachy ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F3.0 | % |  |  |  | 3 |  | Status / Measurements | Atrial Tachy Statistics | Time in ModeSw:
144 | MDC_IDC_STAT_AT_MODE_SW_PERCENT_TIME_PER_DAY | 1.0.0 | 16832 | 737728 |  |  |  |  | Atrial Tachy Statistic Percent Time In Mode Switch Per Day | The percentage of time per day the device is in AT mode switch over the specified time period. | mode switch percent of time per day ¦ atrial tachy ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F3.0 | % |  |  |  | 3 |  | Status / Measurements | Atrial Tachy Statistics | Time in ModeSw per day:
145 | MDC_IDC_STAT_AT_MODE_SW_COUNT | 1.0.0 | 16848 | 737744 |  |  |  |  | Atrial Tachy Statistic Number Of Mode Switches | The number of mode switches over the specified time period. | mode switch count ¦ atrial tachy ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F8.0 | {switches} |  |  |  | 7 |  | Status / Measurements | Atrial Tachy Statistics | Number of ModeSw:
146 | MDC_IDC_STAT_AT_MODE_SW_COUNT_PER_DAY | 1.0.0 | 16864 | 737760 |  |  |  |  | Atrial Tachy Statistic Number Of Mode Switches Per Day | The average number of mode switches per day over the specified time period. | mode switch count per day ¦ atrial tachy ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F8.0 | {switches} |  |  |  | 7 |  | Status / Measurements | Atrial Tachy Statistics | Number of ModeSw per day:
147 | MDC_IDC_STAT_CRT | [5 empty] | 0..1 |  | CRT Statistics | The group of CRT statistics. |  | Channel | GROUP
148 | MDC_IDC_STAT_CRT_DTM_[STRTEND] | 1.0.0 | 16880 | 737776 |  | RE |  | CC5 | CRT Statistic Date Time | The date and time period for the CRT statistics. | date time_[STRTEND] ¦ cardiac resynchronization therapy ¦ statistic ¦ idc ¦ mdc | Metric | DateTime | N/A |  | STRTEND |  |  | 20090201095010 |  | Status / Measurements | CRT Statistics | CRT Statistics Measurement Date:
149 | MDC_IDC_STAT_CRT_LV_PERCENT_PACED | 1.0.0 | 16896 | 737792 |  |  |  |  | CRT Statistic LV Percent Paced | The percentage of LV pacing events in relation to all LV events over the specified time period. | left ventricle percent paced ¦ cardiac resynchronization therapy ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F3.1 | % |  |  |  | 95 |  | Status / Measurements | CRT Statistics | LV Pacing:
150 | MDC_IDC_STAT_CRT_PERCENT_PACED | 1.0.0 | 16912 | 737808 |  |  |  |  | CRT Statistic CRT Percent Paced | The percentage of LV pacing events due to CRT mode in relation to all ventricular events over the specified time period. | percent paced ¦ cardiac resynchronization therapy ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F3.1 | % |  |  |  | 93 |  | Status / Measurements | CRT Statistics | CRT Pacing:
151 | MDC_IDC_STAT_TACHYTHERAPY | [5 empty] | 0..1 |  | Tachy Therapy Statistics | The group of tachy therapy statistics. |  | Channel | GROUP
152 | MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_DELIVERED_RECENT | 1.0.0 | 16928 | 737824 |  |  |  |  | Tachy Therapy Statistic Recent Shocks Delivered | The number of shocks applied over the specified time period. | recent shocks delivered ¦ tachy therapy ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F5.0 | {shocks} |  | Specific definitions for non UCUM units of measure will be described in the profile. |  | 1 |  | Counters / Episodes | Therapy Counts | Shocks Delivered Recent:
153 | MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_DELIVERED_TOTAL | 1.0.0 | 16944 | 737840 |  |  |  |  | Tachy Therapy Statistic Total Shocks Delivered | The total number of shocks applied since implantation or device reset. | total shocks delivered ¦ tachy therapy ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F5.0 | {shocks} |  |  |  | 4 |  | Counters / Episodes | Therapy Counts | Shocks Delivered Total:
154 | MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_ABORTED_RECENT | 1.0.0 | 16960 | 737856 |  |  |  |  | Tachy Therapy Statistic Recent Shocks Aborted | The number of shocks aborted over the specified time period. | recent shocks aborted ¦ tachy therapy ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F5.0 | {shocks} |  |  |  | 0 |  | Counters / Episodes | Therapy Counts | Shocks Aborted Recent:
155 | MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_ABORTED_TOTAL | 1.0.0 | 16976 | 737872 |  |  |  |  | Tachy Therapy Statistic Total Shocks Aborted | The total number of shocks aborted since implantation or device reset. | total shocks aborted ¦ tachy therapy ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F5.0 | {shocks} |  |  |  | 1 |  | Counters / Episodes | Therapy Counts | Shocks Aborted Total:
156 | MDC_IDC_STAT_TACHYTHERAPY_ATP_DELIVERED_RECENT | 1.0.0 | 16992 | 737888 |  |  |  |  | Tachy Therapy Statistic Recent ATP Delivered | The number of anti-tachycardia pacing (ATP) sequences delivered over the specified time period. | recent anti-tachycardia pacing delivered ¦ tachy therapy ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F6.0 | {seq} |  |  |  | 1 |  | Counters / Episodes | Therapy Counts | ATPs Recent:
157 | MDC_IDC_STAT_TACHYTHERAPY_ATP_DELIVERED_TOTAL | 1.0.0 | 17008 | 737904 |  |  |  |  | Tachy Therapy Statistic Total ATP Delivered | The total number of anti-tachycardia pacing (ATP) sequences delivered since implantation or device reset. | total anti-tachycardia pacing delivered ¦ tachy therapy ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F6.0 | {seq} |  |  |  | 3 |  | Counters / Episodes | Therapy Counts | ATPs Total:
158 | MDC_IDC_STAT_TACHYTHERAPY_TOTAL_DTM_[STRTEND] | 1.0.0 | 17024 | 737920 |  |  |  |  | Tachy Therapy Statistic Total Date Time | The date and time period for the total tachy therapy statistics. | date time_[STRTEND] ¦ total ¦ tachy therapy ¦ statistic ¦ idc ¦ mdc | Metric | DateTime | N/A |  | STRTEND |  |  | 20090201095010 |  | Counters / Episodes | Therapy Counts | Counter Measurements Total Dates:
159 | MDC_IDC_STAT_TACHYTHERAPY_RECENT_DTM_[STRTEND] | 1.0.0 | 17040 | 737936 |  | RE |  | CC5 | Tachy Therapy Statistic Recent Date Time | The date and time period for the recent tachy therapy statistics. | date time_[STRTEND] ¦ recent ¦ tachy therapy ¦ statistic ¦ idc ¦ mdc | Metric | DateTime | N/A |  | STRTEND |  |  | 20090201095010 |  | Counters / Episodes | Therapy Counts | Counter Measurements Recent Dates:
160 | MDC_IDC_STAT_EPISODE | [5 empty] | 0..N |  | Episode Statistics | The group of episode statistics. |  | Channel | GROUP
161 | MDC_IDC_STAT_EPISODE_TYPE | 1.0.0 | 17056 | 737952 | Y | RE |  |  | Episode Statistic Type Category | The generic type of the episode being reported. | type ¦ episode ¦ statistic ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | Epis_VT | _MDC_IDC_ENUM_EPISODE_TYPE | Counters / Episodes | Episode Counts | Episode Type:
162 | MDC_IDC_STAT_EPISODE_TYPE_INDUCED | 1.0.0 | 17072 | 737968 |  |  |  |  | Episode Statistic Type Induced | The indication of whether the episode has been induced or not. | type induced ¦ episode ¦ statistic ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | NO | _MDC_IDC_ENUM_EPISODE_TYPE_INDUCED | Counters / Episodes | Episode Counts | N/A
163 | MDC_IDC_STAT_EPISODE_VENDOR_TYPE | 1.0.0 | 17088 | 737984 | Y | RE |  |  | Episode Statistic Vendor Type Category | The vendor type of the episode being reported. | vendor type ¦ episode ¦ statistic ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | VT1 | _MDC_IDC_ENUM_EPISODE_VENDOR_TYPE | Counters / Episodes | Episode Counts | Episode Type:
164 | MDC_IDC_STAT_EPISODE_RECENT_COUNT | 1.0.0 | 17104 | 738000 |  |  |  |  | Episode Statistic Recent Count | The number of episodes per episode type over the specified time period. | recent count ¦ episode ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F7.0 |  |  |  |  | 3 |  | Counters / Episodes | Episode Counts | Episode Count Recent:
165 | MDC_IDC_STAT_EPISODE_RECENT_COUNT_DTM_[STRTEND] | 1.0.0 | 17120 | 738016 |  | RE |  |  | Episode Statistic Recent Date Time | The date and time period for the recent episode counts. | date time_[STRTEND] ¦ recent ¦ episode ¦ statistic ¦ idc ¦ mdc | Metric | DateTime | N/A |  | STRTEND |  |  | 20090201095010 |  | Counters / Episodes | Episode Counts | Episode Count Recent Date Range:
166 | MDC_IDC_STAT_EPISODE_TOTAL_COUNT | 1.0.0 | 17136 | 738032 |  |  |  |  | Episode Statistic Total Count | The number of episodes per episode type since implantation or device reset. | total count ¦ episode ¦ statistic ¦ idc ¦ mdc | Metric | Numeric | F7.0 |  |  |  |  | 6 |  | Counters / Episodes | Episode Counts | Episode Count Total:
167 | MDC_IDC_STAT_EPISODE_TOTAL_COUNT_DTM_[STRTEND] | 1.0.0 | 17152 | 738048 |  |  |  |  | Episode Statistic Total Date Time | The date and time period for the total episode counts. | date time_[STRTEND] ¦ total ¦ episode ¦ statistic ¦ idc ¦ mdc | Metric | DateTime | N/A |  | STRTEND |  |  | 20090201095010 |  | Counters / Episodes | Episode Counts | Episode Count Total Date Range:
168 | MDC_IDC_EPISODE |  | 18624 |  |  |  | 0..N |  | Episode | The group of episodes. |  | Channel | GROUP
169 | MDC_IDC_EPISODE_ID | 1.0.0 | 18640 | 739536 | Y | RE |  |  | Episode Identifier | The unique identifier for the episode. | identifier ¦ episode ¦ idc ¦ mdc | Metric | String | 100 |  |  |  |  | 1 |  | Counters / Episodes | Episode List | Episode ID:
170 | MDC_IDC_EPISODE_DTM | 1.0.0 | 18656 | 739552 |  |  |  |  | Episode Date Time | The detection date and time of the episode. | date time ¦ episode ¦ idc ¦ mdc | Metric | DateTime | N/A |  |  |  |  | 20080222063010 |  | Counters / Episodes | Episode List | Episode Date Time:
171 | MDC_IDC_EPISODE_TYPE | 1.0.0 | 18672 | 739568 |  | RE |  |  | Episode Type Category | The generic type of the episode being reported. | type ¦ episode ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | Epis_VT | _MDC_IDC_ENUM_EPISODE_TYPE | Counters / Episodes | Episode List | Episode Type:
172 | MDC_IDC_EPISODE_TYPE_INDUCED | 1.0.0 | 18688 | 739584 |  |  |  |  | Episode Type Induced Flag | The indication of whether the episode has been induced or not. | type induced ¦ episode ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | NO | _MDC_IDC_ENUM_EPISODE_TYPE_INDUCED | Counters / Episodes | Episode List | Episode Induced Flag:
173 | MDC_IDC_EPISODE_VENDOR_TYPE | 1.0.0 | 18704 | 739600 |  | RE |  |  | Episode Vendor Type Category | The vendor type of the episode being reported. | vendor type ¦ episode ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | VT1 | _MDC_IDC_ENUM_EPISODE_VENDOR_TYPE | Counters / Episodes | Episode List | Episode Type:
174 | MDC_IDC_EPISODE_ATRIAL_INTERVAL_AT_DETECTION | 1.0.0 | 18720 | 739616 |  |  |  |  | Episode Detection Interval Atrial | The atrial interval at the time of detection. | atrial interval at detection ¦ episode ¦ idc ¦ mdc | Metric | Numeric | F4.0 | ms |  |  |  | not used in this example |  | Counters / Episodes | Episode List | Atrial Detection Interval:
175 | MDC_IDC_EPISODE_ATRIAL_INTERVAL_AT_TERMINATION | 1.0.0 | 18736 | 739632 |  |  |  |  | Episode Termination Interval Atrial | The atrial interval at the time of termination. | atrial interval at termination ¦ episode ¦ idc ¦ mdc | Metric | Numeric | F4.0 | ms |  |  |  | not used in this example |  | Counters / Episodes | Episode List | Atrial Termination Interval:
176 | MDC_IDC_EPISODE_VENTRICULAR_INTERVAL_AT_DETECTION | 1.0.0 | 18752 | 739648 |  |  |  |  | Episode Detection Interval Ventricular | The ventricular interval at the time of detection. | ventricular interval at detection ¦ episode ¦ idc ¦ mdc | Metric | Numeric | F4.0 | ms |  |  |  | 147 |  | Counters / Episodes | Episode List | Ventricular Detection Interval:
177 | MDC_IDC_EPISODE_VENTRICULAR_INTERVAL_AT_TERMINATION | 1.0.0 | 18768 | 739664 |  |  |  |  | Episode Termination Interval Ventricular | The ventricular interval at the time of termination. | ventricular interval at termination ¦ episode ¦ idc ¦ mdc | Metric | Numeric | F4.0 | ms |  |  |  | 95 |  | Counters / Episodes | Episode List | Ventricular Termination Interval:
178 | MDC_IDC_EPISODE_DETECTION_THERAPY_DETAILS | 1.0.0 | 18784 | 739680 |  |  |  |  | Episode Detection And Therapy Details | A text describing details about the detection and therapy delivered. | detection therapy details ¦ episode ¦ idc ¦ mdc | Metric | String | 100 |  |  |  |  | "2 ATPs applied, 5x 30J" |  | Counters / Episodes | Episode List | Therapy Details:
179 | MDC_IDC_EPISODE_THERAPY_RESULT | 1.0.0 | 18800 | 739696 |  |  |  |  | Episode Therapy Result | The indicator of therapy success. | therapy result ¦ episode ¦ idc ¦ mdc | Metric | Enum | N/A |  |  |  |  | Successful | _MDC_IDC_ENUM_EPISODE_THERAPY_RESULT | Counters / Episodes | Episode List | Therapy Results:
180 | MDC_IDC_EPISODE_DURATION | 1.0.0 | 18816 | 739712 |  |  |  |  | Episode Duration | The duration of the episode in seconds. | duration ¦ episode ¦ idc ¦ mdc | Metric | Numeric | F8.0 | s |  |  |  | 27 |  | Counters / Episodes | Episode List | Episode Duration:

## Expanded Terms
Reference ID | Systematic Name | Code | Enumeration
MDC_IDC_DEV_TYPE | type ¦ cardiac device ¦ idc ¦ mdc | 720897 | __MDC_IDC_ENUM_PG_TYPE
MDC_IDC_DEV_MODEL | model ¦ cardiac device ¦ idc ¦ mdc | 720898
MDC_IDC_DEV_SERIAL | serial number ¦ cardiac device ¦ idc ¦ mdc | 720899
MDC_IDC_DEV_MFG | manufacturer ¦ cardiac device ¦ idc ¦ mdc | 720900 | __MDC_IDC_ENUM_MFG
MDC_IDC_DEV_IMPLANT_DT | implant date ¦ cardiac device ¦ idc ¦ mdc | 720901
MDC_IDC_DEV_IMPLANTER | implanter ¦ cardiac device ¦ idc ¦ mdc | 720902
MDC_IDC_DEV_IMPLANTER_CONTACT_INFO | implanter contact information ¦ cardiac device ¦ idc ¦ mdc | 720903
MDC_IDC_DEV_IMPLANTING_FACILITY | implanting facility ¦ cardiac device ¦ idc ¦ mdc | 720904
MDC_IDC_LEAD_MODEL | model ¦ lead ¦ idc ¦ mdc | 720961
MDC_IDC_LEAD_SERIAL | serial number ¦ lead ¦ idc ¦ mdc | 720962
MDC_IDC_LEAD_MFG | manufacturer ¦ lead ¦ idc ¦ mdc | 720963 | __MDC_IDC_ENUM_MFG
MDC_IDC_LEAD_IMPLANT_DT | implant date ¦ lead ¦ idc ¦ mdc | 720964
MDC_IDC_LEAD_POLARITY_TYPE | polarity type ¦ lead ¦ idc ¦ mdc | 720965 | __MDC_IDC_ENUM_LEAD_POLARITY_TYPE
MDC_IDC_LEAD_LOCATION | location ¦ lead ¦ idc ¦ mdc | 720966 | __MDC_IDC_ENUM_LEAD_LOCATION_CHAMBER
MDC_IDC_LEAD_LOCATION_DETAIL_1 | detail 1 ¦ location ¦ lead ¦ idc ¦ mdc | 720967 | __MDC_IDC_ENUM_LEAD_LOCATION_DETAIL
MDC_IDC_LEAD_LOCATION_DETAIL_2 | detail 2 ¦ location ¦ lead ¦ idc ¦ mdc | 720968 | __MDC_IDC_ENUM_LEAD_LOCATION_DETAIL
MDC_IDC_LEAD_LOCATION_DETAIL_3 | detail 3 ¦ location ¦ lead ¦ idc ¦ mdc | 720969 | __MDC_IDC_ENUM_LEAD_LOCATION_DETAIL
MDC_IDC_LEAD_CONNECTION_STATUS | connection status ¦ lead ¦ idc ¦ mdc | 720970 | __MDC_IDC_ENUM_LEAD_STATUS
MDC_IDC_LEAD_SPECIAL_FUNCTION | special function ¦ lead ¦ idc ¦ mdc | 720971
MDC_IDC_SESS_DTM | date time ¦ session ¦ idc ¦ mdc | 721025
MDC_IDC_SESS_TYPE | type ¦ session ¦ idc ¦ mdc | 721026 | __MDC_IDC_ENUM_SESS_TYPE
MDC_IDC_SESS_REPROGRAMMED | reprogrammed ¦ session ¦ idc ¦ mdc | 721027 | __MDC_IDC_ENUM_SESS_REPROGRAMMED
MDC_IDC_SESS_DTM_PREVIOUS | previous date time ¦ session ¦ idc ¦ mdc | 721028
MDC_IDC_SESS_TYPE_PREVIOUS | previous type ¦ session ¦ idc ¦ mdc | 721029 | __MDC_IDC_ENUM_SESS_TYPE
MDC_IDC_SESS_REPROGRAMMED_PREVIOUS | reprogrammed previous ¦ session ¦ idc ¦ mdc | 721030 | __MDC_IDC_ENUM_SESS_REPROGRAMMED
MDC_IDC_SESS_CLINICIAN_NAME | clinician name ¦ session ¦ idc ¦ mdc | 721031
MDC_IDC_SESS_CLINICIAN_CONTACT_INFORMATION | clinician contact information ¦ session ¦ idc ¦ mdc | 721032
MDC_IDC_SESS_CLINIC_NAME | clinic name ¦ session ¦ idc ¦ mdc | 721033
MDC_IDC_MSMT_DTM | date time ¦ measurement ¦ idc ¦ mdc | 721152
MDC_IDC_MSMT_DTM_START | date time_start ¦ measurement ¦ idc ¦ mdc | 721153
MDC_IDC_MSMT_DTM_END | date time_end ¦ measurement ¦ idc ¦ mdc | 721154
MDC_IDC_MSMT_BATTERY_DTM | date time ¦ battery ¦ measurement ¦ idc ¦ mdc | 721216
MDC_IDC_MSMT_BATTERY_STATUS | status ¦ battery ¦ measurement ¦ idc ¦ mdc | 721280 | __MDC_IDC_ENUM_BATTERY_STATUS
MDC_IDC_MSMT_BATTERY_VOLTAGE | voltage ¦ battery ¦ measurement ¦ idc ¦ mdc | 721344
MDC_IDC_MSMT_BATTERY_IMPEDANCE | impedance ¦ battery ¦ measurement ¦ idc ¦ mdc | 721408
MDC_IDC_MSMT_BATTERY_REMAINING_LONGEVITY | remaining longevity ¦ battery ¦ measurement ¦ idc ¦ mdc | 721472
MDC_IDC_MSMT_BATTERY_REMAINING_PERCENTAGE | remaining percentage ¦ battery ¦ measurement ¦ idc ¦ mdc | 721536
MDC_IDC_MSMT_BATTERY_RRT_TRIGGER | recommended replacement time trigger ¦ battery ¦ measurement ¦ idc ¦ mdc | 721600
MDC_IDC_MSMT_CAP_CHARGE_DTM | last charge date time ¦ capacitor ¦ measurement ¦ idc ¦ mdc | 721664
MDC_IDC_MSMT_CAP_CHARGE_TIME | charge time ¦ capacitor ¦ measurement ¦ idc ¦ mdc | 721728
MDC_IDC_MSMT_CAP_CHARGE_ENERGY | charge energy ¦ capacitor ¦ measurement ¦ idc ¦ mdc | 721792
MDC_IDC_MSMT_CAP_CHARGE_TYPE | charge type ¦ capacitor ¦ measurement ¦ idc ¦ mdc | 721856 | __MDC_IDC_ENUM_CHARGE_TYPE
MDC_IDC_MSMT_LEADCHNL_RA_DTM | date time ¦ lead channel_ra ¦ measurement ¦ idc ¦ mdc | 721920
MDC_IDC_MSMT_LEADCHNL_RA_DTM_START | date time_start ¦ lead channel_ra ¦ measurement ¦ idc ¦ mdc | 721921
MDC_IDC_MSMT_LEADCHNL_RA_DTM_END | date time_end ¦ lead channel_ra ¦ measurement ¦ idc ¦ mdc | 721922
MDC_IDC_MSMT_LEADCHNL_RV_DTM | date time ¦ lead channel_rv ¦ measurement ¦ idc ¦ mdc | 721924
MDC_IDC_MSMT_LEADCHNL_RV_DTM_START | date time_start ¦ lead channel_rv ¦ measurement ¦ idc ¦ mdc | 721925
MDC_IDC_MSMT_LEADCHNL_RV_DTM_END | date time_end ¦ lead channel_rv ¦ measurement ¦ idc ¦ mdc | 721926
MDC_IDC_MSMT_LEADCHNL_LA_DTM | date time ¦ lead channel_la ¦ measurement ¦ idc ¦ mdc | 721928
MDC_IDC_MSMT_LEADCHNL_LA_DTM_START | date time_start ¦ lead channel_la ¦ measurement ¦ idc ¦ mdc | 721929
MDC_IDC_MSMT_LEADCHNL_LA_DTM_END | date time_end ¦ lead channel_la ¦ measurement ¦ idc ¦ mdc | 721930
MDC_IDC_MSMT_LEADCHNL_LV_DTM | date time ¦ lead channel_lv ¦ measurement ¦ idc ¦ mdc | 721932
MDC_IDC_MSMT_LEADCHNL_LV_DTM_START | date time_start ¦ lead channel_lv ¦ measurement ¦ idc ¦ mdc | 721933
MDC_IDC_MSMT_LEADCHNL_LV_DTM_END | date time_end ¦ lead channel_lv ¦ measurement ¦ idc ¦ mdc | 721934
MDC_IDC_MSMT_LEADCHNL_MIXED_DTM | date time ¦ lead channel_mixed ¦ measurement ¦ idc ¦ mdc | 721936
MDC_IDC_MSMT_LEADCHNL_MIXED_DTM_START | date time_start ¦ lead channel_mixed ¦ measurement ¦ idc ¦ mdc | 721937
MDC_IDC_MSMT_LEADCHNL_MIXED_DTM_END | date time_end ¦ lead channel_mixed ¦ measurement ¦ idc ¦ mdc | 721938
MDC_IDC_MSMT_LEADCHNL_RA_LEAD_CHANNEL_STATUS | lead channel status ¦ lead channel_ra ¦ measurement ¦ idc ¦ mdc | 721984 | __MDC_IDC_ENUM_CHANNEL_STATUS
MDC_IDC_MSMT_LEADCHNL_RV_LEAD_CHANNEL_STATUS | lead channel status ¦ lead channel_rv ¦ measurement ¦ idc ¦ mdc | 721985 | __MDC_IDC_ENUM_CHANNEL_STATUS
MDC_IDC_MSMT_LEADCHNL_LA_LEAD_CHANNEL_STATUS | lead channel status ¦ lead channel_la ¦ measurement ¦ idc ¦ mdc | 721986 | __MDC_IDC_ENUM_CHANNEL_STATUS
MDC_IDC_MSMT_LEADCHNL_LV_LEAD_CHANNEL_STATUS | lead channel status ¦ lead channel_lv ¦ measurement ¦ idc ¦ mdc | 721987 | __MDC_IDC_ENUM_CHANNEL_STATUS
MDC_IDC_MSMT_LEADCHNL_MIXED_LEAD_CHANNEL_STATUS | lead channel status ¦ lead channel_mixed ¦ measurement ¦ idc ¦ mdc | 721988 | __MDC_IDC_ENUM_CHANNEL_STATUS
MDC_IDC_MSMT_LEADCHNL_RA_SENSING_INTR_AMPL | sensing intrinsic amplitude ¦ lead channel_ra ¦ measurement ¦ idc ¦ mdc | 722048
MDC_IDC_MSMT_LEADCHNL_RA_SENSING_INTR_AMPL_MAX | sensing intrinsic amplitude_max ¦ lead channel_ra ¦ measurement ¦ idc ¦ mdc | 722049
MDC_IDC_MSMT_LEADCHNL_RA_SENSING_INTR_AMPL_MIN | sensing intrinsic amplitude_min ¦ lead channel_ra ¦ measurement ¦ idc ¦ mdc | 722050
MDC_IDC_MSMT_LEADCHNL_RA_SENSING_INTR_AMPL_MEAN | sensing intrinsic amplitude_mean ¦ lead channel_ra ¦ measurement ¦ idc ¦ mdc | 722051
MDC_IDC_MSMT_LEADCHNL_RV_SENSING_INTR_AMPL | sensing intrinsic amplitude ¦ lead channel_rv ¦ measurement ¦ idc ¦ mdc | 722052
MDC_IDC_MSMT_LEADCHNL_RV_SENSING_INTR_AMPL_MAX | sensing intrinsic amplitude_max ¦ lead channel_rv ¦ measurement ¦ idc ¦ mdc | 722053
MDC_IDC_MSMT_LEADCHNL_RV_SENSING_INTR_AMPL_MIN | sensing intrinsic amplitude_min ¦ lead channel_rv ¦ measurement ¦ idc ¦ mdc | 722054
MDC_IDC_MSMT_LEADCHNL_RV_SENSING_INTR_AMPL_MEAN | sensing intrinsic amplitude_mean ¦ lead channel_rv ¦ measurement ¦ idc ¦ mdc | 722055
MDC_IDC_MSMT_LEADCHNL_LA_SENSING_INTR_AMPL | sensing intrinsic amplitude ¦ lead channel_la ¦ measurement ¦ idc ¦ mdc | 722056
MDC_IDC_MSMT_LEADCHNL_LA_SENSING_INTR_AMPL_MAX | sensing intrinsic amplitude_max ¦ lead channel_la ¦ measurement ¦ idc ¦ mdc | 722057
MDC_IDC_MSMT_LEADCHNL_LA_SENSING_INTR_AMPL_MIN | sensing intrinsic amplitude_min ¦ lead channel_la ¦ measurement ¦ idc ¦ mdc | 722058
MDC_IDC_MSMT_LEADCHNL_LA_SENSING_INTR_AMPL_MEAN | sensing intrinsic amplitude_mean ¦ lead channel_la ¦ measurement ¦ idc ¦ mdc | 722059
MDC_IDC_MSMT_LEADCHNL_LV_SENSING_INTR_AMPL | sensing intrinsic amplitude ¦ lead channel_lv ¦ measurement ¦ idc ¦ mdc | 722060
MDC_IDC_MSMT_LEADCHNL_LV_SENSING_INTR_AMPL_MAX | sensing intrinsic amplitude_max ¦ lead channel_lv ¦ measurement ¦ idc ¦ mdc | 722061
MDC_IDC_MSMT_LEADCHNL_LV_SENSING_INTR_AMPL_MIN | sensing intrinsic amplitude_min ¦ lead channel_lv ¦ measurement ¦ idc ¦ mdc | 722062
MDC_IDC_MSMT_LEADCHNL_LV_SENSING_INTR_AMPL_MEAN | sensing intrinsic amplitude_mean ¦ lead channel_lv ¦ measurement ¦ idc ¦ mdc | 722063
MDC_IDC_MSMT_LEADCHNL_MIXED_SENSING_INTR_AMPL | sensing intrinsic amplitude ¦ lead channel_mixed ¦ measurement ¦ idc ¦ mdc | 722064
MDC_IDC_MSMT_LEADCHNL_MIXED_SENSING_INTR_AMPL_MAX | sensing intrinsic amplitude_max ¦ lead channel_mixed ¦ measurement ¦ idc ¦ mdc | 722065
MDC_IDC_MSMT_LEADCHNL_MIXED_SENSING_INTR_AMPL_MIN | sensing intrinsic amplitude_min ¦ lead channel_mixed ¦ measurement ¦ idc ¦ mdc | 722066
MDC_IDC_MSMT_LEADCHNL_MIXED_SENSING_INTR_AMPL_MEAN | sensing intrinsic amplitude_mean ¦ lead channel_mixed ¦ measurement ¦ idc ¦ mdc | 722067
MDC_IDC_MSMT_LEADCHNL_RA_SENSING_POLARITY | sensing polarity ¦ lead channel_ra ¦ measurement ¦ idc ¦ mdc | 722112 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_MSMT_LEADCHNL_RV_SENSING_POLARITY | sensing polarity ¦ lead channel_rv ¦ measurement ¦ idc ¦ mdc | 722113 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_MSMT_LEADCHNL_LA_SENSING_POLARITY | sensing polarity ¦ lead channel_la ¦ measurement ¦ idc ¦ mdc | 722114 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_MSMT_LEADCHNL_LV_SENSING_POLARITY | sensing polarity ¦ lead channel_lv ¦ measurement ¦ idc ¦ mdc | 722115 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_MSMT_LEADCHNL_MIXED_SENSING_POLARITY | sensing polarity ¦ lead channel_mixed ¦ measurement ¦ idc ¦ mdc | 722116 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_MSMT_LEADCHNL_RA_PACING_THRESHOLD_AMPLITUDE | amplitude ¦ pacing threshold ¦ lead channel_ra ¦ measurement ¦ idc ¦ mdc | 722176
MDC_IDC_MSMT_LEADCHNL_RV_PACING_THRESHOLD_AMPLITUDE | amplitude ¦ pacing threshold ¦ lead channel_rv ¦ measurement ¦ idc ¦ mdc | 722177
MDC_IDC_MSMT_LEADCHNL_LA_PACING_THRESHOLD_AMPLITUDE | amplitude ¦ pacing threshold ¦ lead channel_la ¦ measurement ¦ idc ¦ mdc | 722178
MDC_IDC_MSMT_LEADCHNL_LV_PACING_THRESHOLD_AMPLITUDE | amplitude ¦ pacing threshold ¦ lead channel_lv ¦ measurement ¦ idc ¦ mdc | 722179
MDC_IDC_MSMT_LEADCHNL_MIXED_PACING_THRESHOLD_AMPLITUDE | amplitude ¦ pacing threshold ¦ lead channel_mixed ¦ measurement ¦ idc ¦ mdc | 722180
MDC_IDC_MSMT_LEADCHNL_RA_PACING_THRESHOLD_PULSEWIDTH | pulse width ¦ pacing threshold ¦ lead channel_ra ¦ measurement ¦ idc ¦ mdc | 722240
MDC_IDC_MSMT_LEADCHNL_RV_PACING_THRESHOLD_PULSEWIDTH | pulse width ¦ pacing threshold ¦ lead channel_rv ¦ measurement ¦ idc ¦ mdc | 722241
MDC_IDC_MSMT_LEADCHNL_LA_PACING_THRESHOLD_PULSEWIDTH | pulse width ¦ pacing threshold ¦ lead channel_la ¦ measurement ¦ idc ¦ mdc | 722242
MDC_IDC_MSMT_LEADCHNL_LV_PACING_THRESHOLD_PULSEWIDTH | pulse width ¦ pacing threshold ¦ lead channel_lv ¦ measurement ¦ idc ¦ mdc | 722243
MDC_IDC_MSMT_LEADCHNL_MIXED_PACING_THRESHOLD_PULSEWIDTH | pulse width ¦ pacing threshold ¦ lead channel_mixed ¦ measurement ¦ idc ¦ mdc | 722244
MDC_IDC_MSMT_LEADCHNL_RA_PACING_THRESHOLD_MEASUREMENT_METHOD | measurement method ¦ pacing threshold ¦ lead channel_ra ¦ measurement ¦ idc ¦ mdc | 722304 | __MDC_IDC_ENUM_MEASUREMENT_METHOD
MDC_IDC_MSMT_LEADCHNL_RV_PACING_THRESHOLD_MEASUREMENT_METHOD | measurement method ¦ pacing threshold ¦ lead channel_rv ¦ measurement ¦ idc ¦ mdc | 722305 | __MDC_IDC_ENUM_MEASUREMENT_METHOD
MDC_IDC_MSMT_LEADCHNL_LA_PACING_THRESHOLD_MEASUREMENT_METHOD | measurement method ¦ pacing threshold ¦ lead channel_la ¦ measurement ¦ idc ¦ mdc | 722306 | __MDC_IDC_ENUM_MEASUREMENT_METHOD
MDC_IDC_MSMT_LEADCHNL_LV_PACING_THRESHOLD_MEASUREMENT_METHOD | measurement method ¦ pacing threshold ¦ lead channel_lv ¦ measurement ¦ idc ¦ mdc | 722307 | __MDC_IDC_ENUM_MEASUREMENT_METHOD
MDC_IDC_MSMT_LEADCHNL_MIXED_PACING_THRESHOLD_MEASUREMENT_METHOD | measurement method ¦ pacing threshold ¦ lead channel_mixed ¦ measurement ¦ idc ¦ mdc | 722308 | __MDC_IDC_ENUM_MEASUREMENT_METHOD
MDC_IDC_MSMT_LEADCHNL_RA_PACING_THRESHOLD_POLARITY | polarity ¦ pacing threshold ¦ lead channel_ra ¦ measurement ¦ idc ¦ mdc | 722368 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_MSMT_LEADCHNL_RV_PACING_THRESHOLD_POLARITY | polarity ¦ pacing threshold ¦ lead channel_rv ¦ measurement ¦ idc ¦ mdc | 722369 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_MSMT_LEADCHNL_LA_PACING_THRESHOLD_POLARITY | polarity ¦ pacing threshold ¦ lead channel_la ¦ measurement ¦ idc ¦ mdc | 722370 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_MSMT_LEADCHNL_LV_PACING_THRESHOLD_POLARITY | polarity ¦ pacing threshold ¦ lead channel_lv ¦ measurement ¦ idc ¦ mdc | 722371 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_MSMT_LEADCHNL_MIXED_PACING_THRESHOLD_POLARITY | polarity ¦ pacing threshold ¦ lead channel_mixed ¦ measurement ¦ idc ¦ mdc | 722372 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_MSMT_LEADCHNL_RA_IMPEDANCE_VALUE | value ¦ impedance ¦ lead channel_ra ¦ measurement ¦ idc ¦ mdc | 722432
MDC_IDC_MSMT_LEADCHNL_RV_IMPEDANCE_VALUE | value ¦ impedance ¦ lead channel_rv ¦ measurement ¦ idc ¦ mdc | 722433
MDC_IDC_MSMT_LEADCHNL_LA_IMPEDANCE_VALUE | value ¦ impedance ¦ lead channel_la ¦ measurement ¦ idc ¦ mdc | 722434
MDC_IDC_MSMT_LEADCHNL_LV_IMPEDANCE_VALUE | value ¦ impedance ¦ lead channel_lv ¦ measurement ¦ idc ¦ mdc | 722435
MDC_IDC_MSMT_LEADCHNL_MIXED_IMPEDANCE_VALUE | value ¦ impedance ¦ lead channel_mixed ¦ measurement ¦ idc ¦ mdc | 722436
MDC_IDC_MSMT_LEADCHNL_RA_IMPEDANCE_POLARITY | polarity ¦ impedance ¦ lead channel_ra ¦ measurement ¦ idc ¦ mdc | 722496 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_MSMT_LEADCHNL_RV_IMPEDANCE_POLARITY | polarity ¦ impedance ¦ lead channel_rv ¦ measurement ¦ idc ¦ mdc | 722497 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_MSMT_LEADCHNL_LA_IMPEDANCE_POLARITY | polarity ¦ impedance ¦ lead channel_la ¦ measurement ¦ idc ¦ mdc | 722498 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_MSMT_LEADCHNL_LV_IMPEDANCE_POLARITY | polarity ¦ impedance ¦ lead channel_lv ¦ measurement ¦ idc ¦ mdc | 722499 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_MSMT_LEADCHNL_MIXED_IMPEDANCE_POLARITY | polarity ¦ impedance ¦ lead channel_mixed ¦ measurement ¦ idc ¦ mdc | 722500 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_MSMT_LEADHVCHNL_DTM | date time ¦ lead high voltage channel ¦ measurement ¦ idc ¦ mdc | 722560
MDC_IDC_MSMT_LEADHVCHNL_DTM_START | date time_start ¦ lead high voltage channel ¦ measurement ¦ idc ¦ mdc | 722561
MDC_IDC_MSMT_LEADHVCHNL_DTM_END | date time_end ¦ lead high voltage channel ¦ measurement ¦ idc ¦ mdc | 722562
MDC_IDC_MSMT_LEADHVCHNL_IMPEDANCE | impedance ¦ lead high voltage channel ¦ measurement ¦ idc ¦ mdc | 722624
MDC_IDC_MSMT_LEADHVCHNL_MEASUREMENT_TYPE | measurement type ¦ lead high voltage channel ¦ measurement ¦ idc ¦ mdc | 722688 | __MDC_IDC_ENUM_HVCHNL_MEASUREMENT_TYPE
MDC_IDC_MSMT_LEADHVCHNL_STATUS | status ¦ lead high voltage channel ¦ measurement ¦ idc ¦ mdc | 722752 | __MDC_IDC_ENUM_CHANNEL_STATUS
MDC_IDC_SET_CRT_LVRV_DELAY | cardiac resynchronization therapy left ventricle right ventricle delay ¦ setting ¦ idc ¦ mdc | 729344
MDC_IDC_SET_CRT_PACED_CHAMBERS | cardiac resynchronization therapy paced chambers ¦ setting ¦ idc ¦ mdc | 729408 | __MDC_IDC_ENUM_CRT_PACED_CHAMBERS
MDC_IDC_SET_MAGNET_RESP | magnet responce ¦ setting ¦ idc ¦ mdc | 729472
MDC_IDC_SET_LEADCHNL_RA_SENSING_SENSITIVITY | sensitivity ¦ sensing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 729536
MDC_IDC_SET_LEADCHNL_RV_SENSING_SENSITIVITY | sensitivity ¦ sensing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 729537
MDC_IDC_SET_LEADCHNL_LA_SENSING_SENSITIVITY | sensitivity ¦ sensing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 729538
MDC_IDC_SET_LEADCHNL_LV_SENSING_SENSITIVITY | sensitivity ¦ sensing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 729539
MDC_IDC_SET_LEADCHNL_MIXED_SENSING_SENSITIVITY | sensitivity ¦ sensing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 729540
MDC_IDC_SET_LEADCHNL_RA_SENSING_POLARITY | polarity ¦ sensing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 729600 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_SET_LEADCHNL_RV_SENSING_POLARITY | polarity ¦ sensing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 729601 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_SET_LEADCHNL_LA_SENSING_POLARITY | polarity ¦ sensing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 729602 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_SET_LEADCHNL_LV_SENSING_POLARITY | polarity ¦ sensing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 729603 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_SET_LEADCHNL_MIXED_SENSING_POLARITY | polarity ¦ sensing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 729604 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_SET_LEADCHNL_RA_SENSING_ANODE_LOCATION | annode location ¦ sensing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 729664
MDC_IDC_SET_LEADCHNL_RA_SENSING_ANODE_LOCATION_1 | annode location_1 ¦ sensing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 729665
MDC_IDC_SET_LEADCHNL_RA_SENSING_ANODE_LOCATION_2 | annode location_2 ¦ sensing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 729666
MDC_IDC_SET_LEADCHNL_RA_SENSING_ANODE_LOCATION_3 | annode location_3 ¦ sensing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 729667
MDC_IDC_SET_LEADCHNL_RV_SENSING_ANODE_LOCATION | annode location ¦ sensing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 729668
MDC_IDC_SET_LEADCHNL_RV_SENSING_ANODE_LOCATION_1 | annode location_1 ¦ sensing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 729669
MDC_IDC_SET_LEADCHNL_RV_SENSING_ANODE_LOCATION_2 | annode location_2 ¦ sensing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 729670
MDC_IDC_SET_LEADCHNL_RV_SENSING_ANODE_LOCATION_3 | annode location_3 ¦ sensing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 729671
MDC_IDC_SET_LEADCHNL_LA_SENSING_ANODE_LOCATION | annode location ¦ sensing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 729672
MDC_IDC_SET_LEADCHNL_LA_SENSING_ANODE_LOCATION_1 | annode location_1 ¦ sensing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 729673
MDC_IDC_SET_LEADCHNL_LA_SENSING_ANODE_LOCATION_2 | annode location_2 ¦ sensing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 729674
MDC_IDC_SET_LEADCHNL_LA_SENSING_ANODE_LOCATION_3 | annode location_3 ¦ sensing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 729675
MDC_IDC_SET_LEADCHNL_LV_SENSING_ANODE_LOCATION | annode location ¦ sensing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 729676
MDC_IDC_SET_LEADCHNL_LV_SENSING_ANODE_LOCATION_1 | annode location_1 ¦ sensing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 729677
MDC_IDC_SET_LEADCHNL_LV_SENSING_ANODE_LOCATION_2 | annode location_2 ¦ sensing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 729678
MDC_IDC_SET_LEADCHNL_LV_SENSING_ANODE_LOCATION_3 | annode location_3 ¦ sensing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 729679
MDC_IDC_SET_LEADCHNL_MIXED_SENSING_ANODE_LOCATION | annode location ¦ sensing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 729680
MDC_IDC_SET_LEADCHNL_MIXED_SENSING_ANODE_LOCATION_1 | annode location_1 ¦ sensing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 729681
MDC_IDC_SET_LEADCHNL_MIXED_SENSING_ANODE_LOCATION_2 | annode location_2 ¦ sensing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 729682
MDC_IDC_SET_LEADCHNL_MIXED_SENSING_ANODE_LOCATION_3 | annode location_3 ¦ sensing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 729683
MDC_IDC_SET_LEADCHNL_RA_SENSING_ANODE_ELECTRODE | annode electrode ¦ sensing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 729728
MDC_IDC_SET_LEADCHNL_RA_SENSING_ANODE_ELECTRODE_1 | annode electrode_1 ¦ sensing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 729729
MDC_IDC_SET_LEADCHNL_RA_SENSING_ANODE_ELECTRODE_2 | annode electrode_2 ¦ sensing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 729730
MDC_IDC_SET_LEADCHNL_RA_SENSING_ANODE_ELECTRODE_3 | annode electrode_3 ¦ sensing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 729731
MDC_IDC_SET_LEADCHNL_RV_SENSING_ANODE_ELECTRODE | annode electrode ¦ sensing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 729732
MDC_IDC_SET_LEADCHNL_RV_SENSING_ANODE_ELECTRODE_1 | annode electrode_1 ¦ sensing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 729733
MDC_IDC_SET_LEADCHNL_RV_SENSING_ANODE_ELECTRODE_2 | annode electrode_2 ¦ sensing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 729734
MDC_IDC_SET_LEADCHNL_RV_SENSING_ANODE_ELECTRODE_3 | annode electrode_3 ¦ sensing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 729735
MDC_IDC_SET_LEADCHNL_LA_SENSING_ANODE_ELECTRODE | annode electrode ¦ sensing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 729736
MDC_IDC_SET_LEADCHNL_LA_SENSING_ANODE_ELECTRODE_1 | annode electrode_1 ¦ sensing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 729737
MDC_IDC_SET_LEADCHNL_LA_SENSING_ANODE_ELECTRODE_2 | annode electrode_2 ¦ sensing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 729738
MDC_IDC_SET_LEADCHNL_LA_SENSING_ANODE_ELECTRODE_3 | annode electrode_3 ¦ sensing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 729739
MDC_IDC_SET_LEADCHNL_LV_SENSING_ANODE_ELECTRODE | annode electrode ¦ sensing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 729740
MDC_IDC_SET_LEADCHNL_LV_SENSING_ANODE_ELECTRODE_1 | annode electrode_1 ¦ sensing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 729741
MDC_IDC_SET_LEADCHNL_LV_SENSING_ANODE_ELECTRODE_2 | annode electrode_2 ¦ sensing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 729742
MDC_IDC_SET_LEADCHNL_LV_SENSING_ANODE_ELECTRODE_3 | annode electrode_3 ¦ sensing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 729743
MDC_IDC_SET_LEADCHNL_MIXED_SENSING_ANODE_ELECTRODE | annode electrode ¦ sensing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 729744
MDC_IDC_SET_LEADCHNL_MIXED_SENSING_ANODE_ELECTRODE_1 | annode electrode_1 ¦ sensing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 729745
MDC_IDC_SET_LEADCHNL_MIXED_SENSING_ANODE_ELECTRODE_2 | annode electrode_2 ¦ sensing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 729746
MDC_IDC_SET_LEADCHNL_MIXED_SENSING_ANODE_ELECTRODE_3 | annode electrode_3 ¦ sensing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 729747
MDC_IDC_SET_LEADCHNL_RA_SENSING_CATHODE_LOCATION | cathode location ¦ sensing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 729792
MDC_IDC_SET_LEADCHNL_RA_SENSING_CATHODE_LOCATION_1 | cathode location_1 ¦ sensing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 729793
MDC_IDC_SET_LEADCHNL_RA_SENSING_CATHODE_LOCATION_2 | cathode location_2 ¦ sensing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 729794
MDC_IDC_SET_LEADCHNL_RA_SENSING_CATHODE_LOCATION_3 | cathode location_3 ¦ sensing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 729795
MDC_IDC_SET_LEADCHNL_RV_SENSING_CATHODE_LOCATION | cathode location ¦ sensing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 729796
MDC_IDC_SET_LEADCHNL_RV_SENSING_CATHODE_LOCATION_1 | cathode location_1 ¦ sensing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 729797
MDC_IDC_SET_LEADCHNL_RV_SENSING_CATHODE_LOCATION_2 | cathode location_2 ¦ sensing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 729798
MDC_IDC_SET_LEADCHNL_RV_SENSING_CATHODE_LOCATION_3 | cathode location_3 ¦ sensing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 729799
MDC_IDC_SET_LEADCHNL_LA_SENSING_CATHODE_LOCATION | cathode location ¦ sensing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 729800
MDC_IDC_SET_LEADCHNL_LA_SENSING_CATHODE_LOCATION_1 | cathode location_1 ¦ sensing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 729801
MDC_IDC_SET_LEADCHNL_LA_SENSING_CATHODE_LOCATION_2 | cathode location_2 ¦ sensing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 729802
MDC_IDC_SET_LEADCHNL_LA_SENSING_CATHODE_LOCATION_3 | cathode location_3 ¦ sensing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 729803
MDC_IDC_SET_LEADCHNL_LV_SENSING_CATHODE_LOCATION | cathode location ¦ sensing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 729804
MDC_IDC_SET_LEADCHNL_LV_SENSING_CATHODE_LOCATION_1 | cathode location_1 ¦ sensing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 729805
MDC_IDC_SET_LEADCHNL_LV_SENSING_CATHODE_LOCATION_2 | cathode location_2 ¦ sensing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 729806
MDC_IDC_SET_LEADCHNL_LV_SENSING_CATHODE_LOCATION_3 | cathode location_3 ¦ sensing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 729807
MDC_IDC_SET_LEADCHNL_MIXED_SENSING_CATHODE_LOCATION | cathode location ¦ sensing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 729808
MDC_IDC_SET_LEADCHNL_MIXED_SENSING_CATHODE_LOCATION_1 | cathode location_1 ¦ sensing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 729809
MDC_IDC_SET_LEADCHNL_MIXED_SENSING_CATHODE_LOCATION_2 | cathode location_2 ¦ sensing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 729810
MDC_IDC_SET_LEADCHNL_MIXED_SENSING_CATHODE_LOCATION_3 | cathode location_3 ¦ sensing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 729811
MDC_IDC_SET_LEADCHNL_RA_SENSING_CATHODE_ELECTRODE | cathode electrode ¦ sensing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 729856
MDC_IDC_SET_LEADCHNL_RA_SENSING_CATHODE_ELECTRODE_1 | cathode electrode_1 ¦ sensing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 729857
MDC_IDC_SET_LEADCHNL_RA_SENSING_CATHODE_ELECTRODE_2 | cathode electrode_2 ¦ sensing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 729858
MDC_IDC_SET_LEADCHNL_RA_SENSING_CATHODE_ELECTRODE_3 | cathode electrode_3 ¦ sensing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 729859
MDC_IDC_SET_LEADCHNL_RV_SENSING_CATHODE_ELECTRODE | cathode electrode ¦ sensing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 729860
MDC_IDC_SET_LEADCHNL_RV_SENSING_CATHODE_ELECTRODE_1 | cathode electrode_1 ¦ sensing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 729861
MDC_IDC_SET_LEADCHNL_RV_SENSING_CATHODE_ELECTRODE_2 | cathode electrode_2 ¦ sensing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 729862
MDC_IDC_SET_LEADCHNL_RV_SENSING_CATHODE_ELECTRODE_3 | cathode electrode_3 ¦ sensing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 729863
MDC_IDC_SET_LEADCHNL_LA_SENSING_CATHODE_ELECTRODE | cathode electrode ¦ sensing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 729864
MDC_IDC_SET_LEADCHNL_LA_SENSING_CATHODE_ELECTRODE_1 | cathode electrode_1 ¦ sensing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 729865
MDC_IDC_SET_LEADCHNL_LA_SENSING_CATHODE_ELECTRODE_2 | cathode electrode_2 ¦ sensing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 729866
MDC_IDC_SET_LEADCHNL_LA_SENSING_CATHODE_ELECTRODE_3 | cathode electrode_3 ¦ sensing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 729867
MDC_IDC_SET_LEADCHNL_LV_SENSING_CATHODE_ELECTRODE | cathode electrode ¦ sensing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 729868
MDC_IDC_SET_LEADCHNL_LV_SENSING_CATHODE_ELECTRODE_1 | cathode electrode_1 ¦ sensing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 729869
MDC_IDC_SET_LEADCHNL_LV_SENSING_CATHODE_ELECTRODE_2 | cathode electrode_2 ¦ sensing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 729870
MDC_IDC_SET_LEADCHNL_LV_SENSING_CATHODE_ELECTRODE_3 | cathode electrode_3 ¦ sensing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 729871
MDC_IDC_SET_LEADCHNL_MIXED_SENSING_CATHODE_ELECTRODE | cathode electrode ¦ sensing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 729872
MDC_IDC_SET_LEADCHNL_MIXED_SENSING_CATHODE_ELECTRODE_1 | cathode electrode_1 ¦ sensing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 729873
MDC_IDC_SET_LEADCHNL_MIXED_SENSING_CATHODE_ELECTRODE_2 | cathode electrode_2 ¦ sensing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 729874
MDC_IDC_SET_LEADCHNL_MIXED_SENSING_CATHODE_ELECTRODE_3 | cathode electrode_3 ¦ sensing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 729875
MDC_IDC_SET_LEADCHNL_RA_SENSING_ADAPTATION_MODE | adaptation mode ¦ sensing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 729920 | __MDC_IDC_ENUM_SENSING_ADAPTATION_MODE
MDC_IDC_SET_LEADCHNL_RV_SENSING_ADAPTATION_MODE | adaptation mode ¦ sensing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 729921 | __MDC_IDC_ENUM_SENSING_ADAPTATION_MODE
MDC_IDC_SET_LEADCHNL_LA_SENSING_ADAPTATION_MODE | adaptation mode ¦ sensing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 729922 | __MDC_IDC_ENUM_SENSING_ADAPTATION_MODE
MDC_IDC_SET_LEADCHNL_LV_SENSING_ADAPTATION_MODE | adaptation mode ¦ sensing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 729923 | __MDC_IDC_ENUM_SENSING_ADAPTATION_MODE
MDC_IDC_SET_LEADCHNL_MIXED_SENSING_ADAPTATION_MODE | adaptation mode ¦ sensing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 729924 | __MDC_IDC_ENUM_SENSING_ADAPTATION_MODE
MDC_IDC_SET_LEADCHNL_RA_PACING_AMPLITUDE | amplitude ¦ pacing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 729984
MDC_IDC_SET_LEADCHNL_RV_PACING_AMPLITUDE | amplitude ¦ pacing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 729985
MDC_IDC_SET_LEADCHNL_LA_PACING_AMPLITUDE | amplitude ¦ pacing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 729986
MDC_IDC_SET_LEADCHNL_LV_PACING_AMPLITUDE | amplitude ¦ pacing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 729987
MDC_IDC_SET_LEADCHNL_MIXED_PACING_AMPLITUDE | amplitude ¦ pacing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 729988
MDC_IDC_SET_LEADCHNL_RA_PACING_PULSEWIDTH | pulse width ¦ pacing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 730048
MDC_IDC_SET_LEADCHNL_RV_PACING_PULSEWIDTH | pulse width ¦ pacing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 730049
MDC_IDC_SET_LEADCHNL_LA_PACING_PULSEWIDTH | pulse width ¦ pacing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 730050
MDC_IDC_SET_LEADCHNL_LV_PACING_PULSEWIDTH | pulse width ¦ pacing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 730051
MDC_IDC_SET_LEADCHNL_MIXED_PACING_PULSEWIDTH | pulse width ¦ pacing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 730052
MDC_IDC_SET_LEADCHNL_RA_PACING_POLARITY | polarity ¦ pacing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 730112 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_SET_LEADCHNL_RV_PACING_POLARITY | polarity ¦ pacing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 730113 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_SET_LEADCHNL_LA_PACING_POLARITY | polarity ¦ pacing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 730114 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_SET_LEADCHNL_LV_PACING_POLARITY | polarity ¦ pacing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 730115 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_SET_LEADCHNL_MIXED_PACING_POLARITY | polarity ¦ pacing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 730116 | __MDC_IDC_ENUM_POLARITY
MDC_IDC_SET_LEADCHNL_RA_PACING_ANODE_LOCATION | annode location ¦ pacing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 730176
MDC_IDC_SET_LEADCHNL_RA_PACING_ANODE_LOCATION_1 | annode location_1 ¦ pacing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 730177
MDC_IDC_SET_LEADCHNL_RA_PACING_ANODE_LOCATION_2 | annode location_2 ¦ pacing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 730178
MDC_IDC_SET_LEADCHNL_RA_PACING_ANODE_LOCATION_3 | annode location_3 ¦ pacing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 730179
MDC_IDC_SET_LEADCHNL_RV_PACING_ANODE_LOCATION | annode location ¦ pacing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 730180
MDC_IDC_SET_LEADCHNL_RV_PACING_ANODE_LOCATION_1 | annode location_1 ¦ pacing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 730181
MDC_IDC_SET_LEADCHNL_RV_PACING_ANODE_LOCATION_2 | annode location_2 ¦ pacing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 730182
MDC_IDC_SET_LEADCHNL_RV_PACING_ANODE_LOCATION_3 | annode location_3 ¦ pacing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 730183
MDC_IDC_SET_LEADCHNL_LA_PACING_ANODE_LOCATION | annode location ¦ pacing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 730184
MDC_IDC_SET_LEADCHNL_LA_PACING_ANODE_LOCATION_1 | annode location_1 ¦ pacing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 730185
MDC_IDC_SET_LEADCHNL_LA_PACING_ANODE_LOCATION_2 | annode location_2 ¦ pacing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 730186
MDC_IDC_SET_LEADCHNL_LA_PACING_ANODE_LOCATION_3 | annode location_3 ¦ pacing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 730187
MDC_IDC_SET_LEADCHNL_LV_PACING_ANODE_LOCATION | annode location ¦ pacing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 730188
MDC_IDC_SET_LEADCHNL_LV_PACING_ANODE_LOCATION_1 | annode location_1 ¦ pacing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 730189
MDC_IDC_SET_LEADCHNL_LV_PACING_ANODE_LOCATION_2 | annode location_2 ¦ pacing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 730190
MDC_IDC_SET_LEADCHNL_LV_PACING_ANODE_LOCATION_3 | annode location_3 ¦ pacing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 730191
MDC_IDC_SET_LEADCHNL_MIXED_PACING_ANODE_LOCATION | annode location ¦ pacing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 730192
MDC_IDC_SET_LEADCHNL_MIXED_PACING_ANODE_LOCATION_1 | annode location_1 ¦ pacing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 730193
MDC_IDC_SET_LEADCHNL_MIXED_PACING_ANODE_LOCATION_2 | annode location_2 ¦ pacing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 730194
MDC_IDC_SET_LEADCHNL_MIXED_PACING_ANODE_LOCATION_3 | annode location_3 ¦ pacing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 730195
MDC_IDC_SET_LEADCHNL_RA_PACING_ANODE_ELECTRODE | annode electrode ¦ pacing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 730240
MDC_IDC_SET_LEADCHNL_RA_PACING_ANODE_ELECTRODE_1 | annode electrode_1 ¦ pacing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 730241
MDC_IDC_SET_LEADCHNL_RA_PACING_ANODE_ELECTRODE_2 | annode electrode_2 ¦ pacing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 730242
MDC_IDC_SET_LEADCHNL_RA_PACING_ANODE_ELECTRODE_3 | annode electrode_3 ¦ pacing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 730243
MDC_IDC_SET_LEADCHNL_RV_PACING_ANODE_ELECTRODE | annode electrode ¦ pacing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 730244
MDC_IDC_SET_LEADCHNL_RV_PACING_ANODE_ELECTRODE_1 | annode electrode_1 ¦ pacing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 730245
MDC_IDC_SET_LEADCHNL_RV_PACING_ANODE_ELECTRODE_2 | annode electrode_2 ¦ pacing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 730246
MDC_IDC_SET_LEADCHNL_RV_PACING_ANODE_ELECTRODE_3 | annode electrode_3 ¦ pacing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 730247
MDC_IDC_SET_LEADCHNL_LA_PACING_ANODE_ELECTRODE | annode electrode ¦ pacing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 730248
MDC_IDC_SET_LEADCHNL_LA_PACING_ANODE_ELECTRODE_1 | annode electrode_1 ¦ pacing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 730249
MDC_IDC_SET_LEADCHNL_LA_PACING_ANODE_ELECTRODE_2 | annode electrode_2 ¦ pacing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 730250
MDC_IDC_SET_LEADCHNL_LA_PACING_ANODE_ELECTRODE_3 | annode electrode_3 ¦ pacing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 730251
MDC_IDC_SET_LEADCHNL_LV_PACING_ANODE_ELECTRODE | annode electrode ¦ pacing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 730252
MDC_IDC_SET_LEADCHNL_LV_PACING_ANODE_ELECTRODE_1 | annode electrode_1 ¦ pacing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 730253
MDC_IDC_SET_LEADCHNL_LV_PACING_ANODE_ELECTRODE_2 | annode electrode_2 ¦ pacing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 730254
MDC_IDC_SET_LEADCHNL_LV_PACING_ANODE_ELECTRODE_3 | annode electrode_3 ¦ pacing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 730255
MDC_IDC_SET_LEADCHNL_MIXED_PACING_ANODE_ELECTRODE | annode electrode ¦ pacing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 730256
MDC_IDC_SET_LEADCHNL_MIXED_PACING_ANODE_ELECTRODE_1 | annode electrode_1 ¦ pacing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 730257
MDC_IDC_SET_LEADCHNL_MIXED_PACING_ANODE_ELECTRODE_2 | annode electrode_2 ¦ pacing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 730258
MDC_IDC_SET_LEADCHNL_MIXED_PACING_ANODE_ELECTRODE_3 | annode electrode_3 ¦ pacing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 730259
MDC_IDC_SET_LEADCHNL_RA_PACING_CATHODE_LOCATION | cathode location ¦ pacing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 730304
MDC_IDC_SET_LEADCHNL_RA_PACING_CATHODE_LOCATION_1 | cathode location_1 ¦ pacing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 730305
MDC_IDC_SET_LEADCHNL_RA_PACING_CATHODE_LOCATION_2 | cathode location_2 ¦ pacing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 730306
MDC_IDC_SET_LEADCHNL_RA_PACING_CATHODE_LOCATION_3 | cathode location_3 ¦ pacing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 730307
MDC_IDC_SET_LEADCHNL_RV_PACING_CATHODE_LOCATION | cathode location ¦ pacing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 730308
MDC_IDC_SET_LEADCHNL_RV_PACING_CATHODE_LOCATION_1 | cathode location_1 ¦ pacing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 730309
MDC_IDC_SET_LEADCHNL_RV_PACING_CATHODE_LOCATION_2 | cathode location_2 ¦ pacing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 730310
MDC_IDC_SET_LEADCHNL_RV_PACING_CATHODE_LOCATION_3 | cathode location_3 ¦ pacing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 730311
MDC_IDC_SET_LEADCHNL_LA_PACING_CATHODE_LOCATION | cathode location ¦ pacing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 730312
MDC_IDC_SET_LEADCHNL_LA_PACING_CATHODE_LOCATION_1 | cathode location_1 ¦ pacing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 730313
MDC_IDC_SET_LEADCHNL_LA_PACING_CATHODE_LOCATION_2 | cathode location_2 ¦ pacing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 730314
MDC_IDC_SET_LEADCHNL_LA_PACING_CATHODE_LOCATION_3 | cathode location_3 ¦ pacing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 730315
MDC_IDC_SET_LEADCHNL_LV_PACING_CATHODE_LOCATION | cathode location ¦ pacing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 730316
MDC_IDC_SET_LEADCHNL_LV_PACING_CATHODE_LOCATION_1 | cathode location_1 ¦ pacing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 730317
MDC_IDC_SET_LEADCHNL_LV_PACING_CATHODE_LOCATION_2 | cathode location_2 ¦ pacing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 730318
MDC_IDC_SET_LEADCHNL_LV_PACING_CATHODE_LOCATION_3 | cathode location_3 ¦ pacing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 730319
MDC_IDC_SET_LEADCHNL_MIXED_PACING_CATHODE_LOCATION | cathode location ¦ pacing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 730320
MDC_IDC_SET_LEADCHNL_MIXED_PACING_CATHODE_LOCATION_1 | cathode location_1 ¦ pacing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 730321
MDC_IDC_SET_LEADCHNL_MIXED_PACING_CATHODE_LOCATION_2 | cathode location_2 ¦ pacing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 730322
MDC_IDC_SET_LEADCHNL_MIXED_PACING_CATHODE_LOCATION_3 | cathode location_3 ¦ pacing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 730323
MDC_IDC_SET_LEADCHNL_RA_PACING_CATHODE_ELECTRODE | cathode electrode ¦ pacing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 730368
MDC_IDC_SET_LEADCHNL_RA_PACING_CATHODE_ELECTRODE_1 | cathode electrode_1 ¦ pacing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 730369
MDC_IDC_SET_LEADCHNL_RA_PACING_CATHODE_ELECTRODE_2 | cathode electrode_2 ¦ pacing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 730370
MDC_IDC_SET_LEADCHNL_RA_PACING_CATHODE_ELECTRODE_3 | cathode electrode_3 ¦ pacing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 730371
MDC_IDC_SET_LEADCHNL_RV_PACING_CATHODE_ELECTRODE | cathode electrode ¦ pacing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 730372
MDC_IDC_SET_LEADCHNL_RV_PACING_CATHODE_ELECTRODE_1 | cathode electrode_1 ¦ pacing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 730373
MDC_IDC_SET_LEADCHNL_RV_PACING_CATHODE_ELECTRODE_2 | cathode electrode_2 ¦ pacing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 730374
MDC_IDC_SET_LEADCHNL_RV_PACING_CATHODE_ELECTRODE_3 | cathode electrode_3 ¦ pacing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 730375
MDC_IDC_SET_LEADCHNL_LA_PACING_CATHODE_ELECTRODE | cathode electrode ¦ pacing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 730376
MDC_IDC_SET_LEADCHNL_LA_PACING_CATHODE_ELECTRODE_1 | cathode electrode_1 ¦ pacing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 730377
MDC_IDC_SET_LEADCHNL_LA_PACING_CATHODE_ELECTRODE_2 | cathode electrode_2 ¦ pacing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 730378
MDC_IDC_SET_LEADCHNL_LA_PACING_CATHODE_ELECTRODE_3 | cathode electrode_3 ¦ pacing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 730379
MDC_IDC_SET_LEADCHNL_LV_PACING_CATHODE_ELECTRODE | cathode electrode ¦ pacing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 730380
MDC_IDC_SET_LEADCHNL_LV_PACING_CATHODE_ELECTRODE_1 | cathode electrode_1 ¦ pacing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 730381
MDC_IDC_SET_LEADCHNL_LV_PACING_CATHODE_ELECTRODE_2 | cathode electrode_2 ¦ pacing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 730382
MDC_IDC_SET_LEADCHNL_LV_PACING_CATHODE_ELECTRODE_3 | cathode electrode_3 ¦ pacing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 730383
MDC_IDC_SET_LEADCHNL_MIXED_PACING_CATHODE_ELECTRODE | cathode electrode ¦ pacing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 730384
MDC_IDC_SET_LEADCHNL_MIXED_PACING_CATHODE_ELECTRODE_1 | cathode electrode_1 ¦ pacing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 730385
MDC_IDC_SET_LEADCHNL_MIXED_PACING_CATHODE_ELECTRODE_2 | cathode electrode_2 ¦ pacing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 730386
MDC_IDC_SET_LEADCHNL_MIXED_PACING_CATHODE_ELECTRODE_3 | cathode electrode_3 ¦ pacing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 730387
MDC_IDC_SET_LEADCHNL_RA_PACING_CAPTURE_MODE | capture mode ¦ pacing ¦ lead channel_ra ¦ setting ¦ idc ¦ mdc | 730432 | __MDC_IDC_ENUM_PACING_CAPTURE_MODE
MDC_IDC_SET_LEADCHNL_RV_PACING_CAPTURE_MODE | capture mode ¦ pacing ¦ lead channel_rv ¦ setting ¦ idc ¦ mdc | 730433 | __MDC_IDC_ENUM_PACING_CAPTURE_MODE
MDC_IDC_SET_LEADCHNL_LA_PACING_CAPTURE_MODE | capture mode ¦ pacing ¦ lead channel_la ¦ setting ¦ idc ¦ mdc | 730434 | __MDC_IDC_ENUM_PACING_CAPTURE_MODE
MDC_IDC_SET_LEADCHNL_LV_PACING_CAPTURE_MODE | capture mode ¦ pacing ¦ lead channel_lv ¦ setting ¦ idc ¦ mdc | 730435 | __MDC_IDC_ENUM_PACING_CAPTURE_MODE
MDC_IDC_SET_LEADCHNL_MIXED_PACING_CAPTURE_MODE | capture mode ¦ pacing ¦ lead channel_mixed ¦ setting ¦ idc ¦ mdc | 730436 | __MDC_IDC_ENUM_PACING_CAPTURE_MODE
MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_LOCATION | annode location ¦ vector ¦ shock ¦ lead high voltage channel ¦ setting ¦ idc ¦ mdc | 730496 | __MDC_IDC_ENUM_ELECTRODE_LOCATION
MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_LOCATION_1 | annode location_1 ¦ vector ¦ shock ¦ lead high voltage channel ¦ setting ¦ idc ¦ mdc | 730497 | __MDC_IDC_ENUM_ELECTRODE_LOCATION
MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_LOCATION_2 | annode location_2 ¦ vector ¦ shock ¦ lead high voltage channel ¦ setting ¦ idc ¦ mdc | 730498 | __MDC_IDC_ENUM_ELECTRODE_LOCATION
MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_LOCATION_3 | annode location_3 ¦ vector ¦ shock ¦ lead high voltage channel ¦ setting ¦ idc ¦ mdc | 730499 | __MDC_IDC_ENUM_ELECTRODE_LOCATION
MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_ELECTRODE | annode electrode ¦ vector ¦ shock ¦ lead high voltage channel ¦ setting ¦ idc ¦ mdc | 730560 | __MDC_IDC_ENUM_ELECTRODE_NAME
MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_ELECTRODE_1 | annode electrode_1 ¦ vector ¦ shock ¦ lead high voltage channel ¦ setting ¦ idc ¦ mdc | 730561 | __MDC_IDC_ENUM_ELECTRODE_NAME
MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_ELECTRODE_2 | annode electrode_2 ¦ vector ¦ shock ¦ lead high voltage channel ¦ setting ¦ idc ¦ mdc | 730562 | __MDC_IDC_ENUM_ELECTRODE_NAME
MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_ELECTRODE_3 | annode electrode_3 ¦ vector ¦ shock ¦ lead high voltage channel ¦ setting ¦ idc ¦ mdc | 730563 | __MDC_IDC_ENUM_ELECTRODE_NAME
MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_CATHODE_LOCATION | cathode location ¦ vector ¦ shock ¦ lead high voltage channel ¦ setting ¦ idc ¦ mdc | 730624 | __MDC_IDC_ENUM_ELECTRODE_LOCATION
MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_CATHODE_LOCATION_1 | cathode location_1 ¦ vector ¦ shock ¦ lead high voltage channel ¦ setting ¦ idc ¦ mdc | 730625 | __MDC_IDC_ENUM_ELECTRODE_LOCATION
MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_CATHODE_LOCATION_2 | cathode location_2 ¦ vector ¦ shock ¦ lead high voltage channel ¦ setting ¦ idc ¦ mdc | 730626 | __MDC_IDC_ENUM_ELECTRODE_LOCATION
MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_CATHODE_LOCATION_3 | cathode location_3 ¦ vector ¦ shock ¦ lead high voltage channel ¦ setting ¦ idc ¦ mdc | 730627 | __MDC_IDC_ENUM_ELECTRODE_LOCATION
MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_CATHODE_ELECTRODE | cathode electrode ¦ vector ¦ shock ¦ lead high voltage channel ¦ setting ¦ idc ¦ mdc | 730688 | __MDC_IDC_ENUM_ELECTRODE_NAME
MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_CATHODE_ELECTRODE_1 | cathode electrode_1 ¦ vector ¦ shock ¦ lead high voltage channel ¦ setting ¦ idc ¦ mdc | 730689 | __MDC_IDC_ENUM_ELECTRODE_NAME
MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_CATHODE_ELECTRODE_2 | cathode electrode_2 ¦ vector ¦ shock ¦ lead high voltage channel ¦ setting ¦ idc ¦ mdc | 730690 | __MDC_IDC_ENUM_ELECTRODE_NAME
MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_CATHODE_ELECTRODE_3 | cathode electrode_3 ¦ vector ¦ shock ¦ lead high voltage channel ¦ setting ¦ idc ¦ mdc | 730691 | __MDC_IDC_ENUM_ELECTRODE_NAME
MDC_IDC_SET_BRADY_MODE | mode ¦ brady ¦ setting ¦ idc ¦ mdc | 730752 | __MDC_IDC_ENUM_BRADY_MODE
MDC_IDC_SET_BRADY_VENDOR_MODE | mode ¦ brady ¦ setting ¦ idc ¦ mdc | 730816
MDC_IDC_SET_BRADY_LOWRATE | lowrate ¦ brady ¦ setting ¦ idc ¦ mdc | 730880
MDC_IDC_SET_BRADY_HYSTRATE | hystrate ¦ brady ¦ setting ¦ idc ¦ mdc | 730944
MDC_IDC_SET_BRADY_NIGHT_RATE | night rate ¦ brady ¦ setting ¦ idc ¦ mdc | 731008
MDC_IDC_SET_BRADY_SENSOR_TYPE | sensor type ¦ brady ¦ setting ¦ idc ¦ mdc | 731072
MDC_IDC_SET_BRADY_MAX_TRACKING_RATE | maximum tracking rate ¦ brady ¦ setting ¦ idc ¦ mdc | 731136
MDC_IDC_SET_BRADY_MAX_SENSOR_RATE | maximum sensor rate ¦ brady ¦ setting ¦ idc ¦ mdc | 731200
MDC_IDC_SET_BRADY_SAV_DELAY | sav delay ¦ brady ¦ setting ¦ idc ¦ mdc | 731264
MDC_IDC_SET_BRADY_SAV_DELAY_HIGH | sav delay_high ¦ brady ¦ setting ¦ idc ¦ mdc | 731265
MDC_IDC_SET_BRADY_SAV_DELAY_LOW | sav delay_low ¦ brady ¦ setting ¦ idc ¦ mdc | 731266
MDC_IDC_SET_BRADY_PAV_DELAY | pav delay ¦ brady ¦ setting ¦ idc ¦ mdc | 731328
MDC_IDC_SET_BRADY_PAV_DELAY_HIGH | pav delay_high ¦ brady ¦ setting ¦ idc ¦ mdc | 731329
MDC_IDC_SET_BRADY_PAV_DELAY_LOW | pav delay_low ¦ brady ¦ setting ¦ idc ¦ mdc | 731330
MDC_IDC_SET_BRADY_AT_MODE_SWITCH_MODE | atrial tachy mode switch mode ¦ brady ¦ setting ¦ idc ¦ mdc | 731392 | __MDC_IDC_ENUM_BRADY_MODE
MDC_IDC_SET_BRADY_AT_MODE_SWITCH_RATE | atrail tachy mode switch rate ¦ brady ¦ setting ¦ idc ¦ mdc | 731456
MDC_IDC_SET_TACHYTHERAPY_VSTAT | ventricular status ¦ tachytherapy ¦ setting ¦ idc ¦ mdc | 731520 | __MDC_IDC_ENUM_THERAPY_STATUS
MDC_IDC_SET_TACHYTHERAPY_ASTAT | atrial status ¦ tachytherapy ¦ setting ¦ idc ¦ mdc | 731584 | __MDC_IDC_ENUM_THERAPY_STATUS
MDC_IDC_SET_ZONE_TYPE | type ¦ zone ¦ setting ¦ idc ¦ mdc | 731648 | __MDC_IDC_ENUM_ZONE_TYPE
MDC_IDC_SET_ZONE_VENDOR_TYPE | vendor type ¦ zone ¦ setting ¦ idc ¦ mdc | 731712 | __MDC_IDC_ENUM_ZONE_VENDOR_TYPE
MDC_IDC_SET_ZONE_STATUS | status ¦ zone ¦ setting ¦ idc ¦ mdc | 731776 | __MDC_IDC_ENUM_ZONE_STATUS
MDC_IDC_SET_ZONE_DETECTION_INTERVAL | detection interval ¦ zone ¦ setting ¦ idc ¦ mdc | 731840
MDC_IDC_SET_ZONE_DETECTION_BEATS_NUMERATOR | detection beats numerator ¦ zone ¦ setting ¦ idc ¦ mdc | 731904
MDC_IDC_SET_ZONE_DETECTION_BEATS_DENOMINATOR | detection beats denominator ¦ zone ¦ setting ¦ idc ¦ mdc | 731968
MDC_IDC_SET_ZONE_DETECTION_DETAILS | detection details ¦ zone ¦ setting ¦ idc ¦ mdc | 732032
MDC_IDC_SET_ZONE_TYPE_ATP | type anti-tachycardia pacing pulse ¦ zone ¦ setting ¦ idc ¦ mdc | 732096 | __MDC_IDC_ENUM_ATP_TYPE
MDC_IDC_SET_ZONE_TYPE_ATP_1 | type anti-tachycardia pacing pulse_1 ¦ zone ¦ setting ¦ idc ¦ mdc | 732097 | __MDC_IDC_ENUM_ATP_TYPE
MDC_IDC_SET_ZONE_TYPE_ATP_2 | type anti-tachycardia pacing pulse_2 ¦ zone ¦ setting ¦ idc ¦ mdc | 732098 | __MDC_IDC_ENUM_ATP_TYPE
MDC_IDC_SET_ZONE_TYPE_ATP_3 | type anti-tachycardia pacing pulse_3 ¦ zone ¦ setting ¦ idc ¦ mdc | 732099 | __MDC_IDC_ENUM_ATP_TYPE
MDC_IDC_SET_ZONE_TYPE_ATP_4 | type anti-tachycardia pacing pulse_4 ¦ zone ¦ setting ¦ idc ¦ mdc | 732100 | __MDC_IDC_ENUM_ATP_TYPE
MDC_IDC_SET_ZONE_TYPE_ATP_5 | type anti-tachycardia pacing pulse_5 ¦ zone ¦ setting ¦ idc ¦ mdc | 732101 | __MDC_IDC_ENUM_ATP_TYPE
MDC_IDC_SET_ZONE_TYPE_ATP_6 | type anti-tachycardia pacing pulse_6 ¦ zone ¦ setting ¦ idc ¦ mdc | 732102 | __MDC_IDC_ENUM_ATP_TYPE
MDC_IDC_SET_ZONE_TYPE_ATP_7 | type anti-tachycardia pacing pulse_7 ¦ zone ¦ setting ¦ idc ¦ mdc | 732103 | __MDC_IDC_ENUM_ATP_TYPE
MDC_IDC_SET_ZONE_TYPE_ATP_8 | type anti-tachycardia pacing pulse_8 ¦ zone ¦ setting ¦ idc ¦ mdc | 732104 | __MDC_IDC_ENUM_ATP_TYPE
MDC_IDC_SET_ZONE_TYPE_ATP_9 | type anti-tachycardia pacing pulse_9 ¦ zone ¦ setting ¦ idc ¦ mdc | 732105 | __MDC_IDC_ENUM_ATP_TYPE
MDC_IDC_SET_ZONE_TYPE_ATP_10 | type anti-tachycardia pacing pulse_10 ¦ zone ¦ setting ¦ idc ¦ mdc | 732106 | __MDC_IDC_ENUM_ATP_TYPE
MDC_IDC_SET_ZONE_NUM_ATP_SEQS | number of anti-tachycardia pacing pulse sequences ¦ zone ¦ setting ¦ idc ¦ mdc | 732160
MDC_IDC_SET_ZONE_NUM_ATP_SEQS_1 | number of anti-tachycardia pacing pulse sequences_1 ¦ zone ¦ setting ¦ idc ¦ mdc | 732161
MDC_IDC_SET_ZONE_NUM_ATP_SEQS_2 | number of anti-tachycardia pacing pulse sequences_2 ¦ zone ¦ setting ¦ idc ¦ mdc | 732162
MDC_IDC_SET_ZONE_NUM_ATP_SEQS_3 | number of anti-tachycardia pacing pulse sequences_3 ¦ zone ¦ setting ¦ idc ¦ mdc | 732163
MDC_IDC_SET_ZONE_NUM_ATP_SEQS_4 | number of anti-tachycardia pacing pulse sequences_4 ¦ zone ¦ setting ¦ idc ¦ mdc | 732164
MDC_IDC_SET_ZONE_NUM_ATP_SEQS_5 | number of anti-tachycardia pacing pulse sequences_5 ¦ zone ¦ setting ¦ idc ¦ mdc | 732165
MDC_IDC_SET_ZONE_NUM_ATP_SEQS_6 | number of anti-tachycardia pacing pulse sequences_6 ¦ zone ¦ setting ¦ idc ¦ mdc | 732166
MDC_IDC_SET_ZONE_NUM_ATP_SEQS_7 | number of anti-tachycardia pacing pulse sequences_7 ¦ zone ¦ setting ¦ idc ¦ mdc | 732167
MDC_IDC_SET_ZONE_NUM_ATP_SEQS_8 | number of anti-tachycardia pacing pulse sequences_8 ¦ zone ¦ setting ¦ idc ¦ mdc | 732168
MDC_IDC_SET_ZONE_NUM_ATP_SEQS_9 | number of anti-tachycardia pacing pulse sequences_9 ¦ zone ¦ setting ¦ idc ¦ mdc | 732169
MDC_IDC_SET_ZONE_NUM_ATP_SEQS_10 | number of anti-tachycardia pacing pulse sequences_10 ¦ zone ¦ setting ¦ idc ¦ mdc | 732170
MDC_IDC_SET_ZONE_SHOCK_ENERGY | shock energy ¦ zone ¦ setting ¦ idc ¦ mdc | 732224
MDC_IDC_SET_ZONE_SHOCK_ENERGY_1 | shock energy_1 ¦ zone ¦ setting ¦ idc ¦ mdc | 732225
MDC_IDC_SET_ZONE_SHOCK_ENERGY_2 | shock energy_2 ¦ zone ¦ setting ¦ idc ¦ mdc | 732226
MDC_IDC_SET_ZONE_SHOCK_ENERGY_3 | shock energy_3 ¦ zone ¦ setting ¦ idc ¦ mdc | 732227
MDC_IDC_SET_ZONE_SHOCK_ENERGY_4 | shock energy_4 ¦ zone ¦ setting ¦ idc ¦ mdc | 732228
MDC_IDC_SET_ZONE_SHOCK_ENERGY_5 | shock energy_5 ¦ zone ¦ setting ¦ idc ¦ mdc | 732229
MDC_IDC_SET_ZONE_SHOCK_ENERGY_6 | shock energy_6 ¦ zone ¦ setting ¦ idc ¦ mdc | 732230
MDC_IDC_SET_ZONE_SHOCK_ENERGY_7 | shock energy_7 ¦ zone ¦ setting ¦ idc ¦ mdc | 732231
MDC_IDC_SET_ZONE_SHOCK_ENERGY_8 | shock energy_8 ¦ zone ¦ setting ¦ idc ¦ mdc | 732232
MDC_IDC_SET_ZONE_SHOCK_ENERGY_9 | shock energy_9 ¦ zone ¦ setting ¦ idc ¦ mdc | 732233
MDC_IDC_SET_ZONE_SHOCK_ENERGY_10 | shock energy_10 ¦ zone ¦ setting ¦ idc ¦ mdc | 732234
MDC_IDC_SET_ZONE_NUM_SHOCKS | number of shocks ¦ zone ¦ setting ¦ idc ¦ mdc | 732288
MDC_IDC_SET_ZONE_NUM_SHOCKS_1 | number of shocks_1 ¦ zone ¦ setting ¦ idc ¦ mdc | 732289
MDC_IDC_SET_ZONE_NUM_SHOCKS_2 | number of shocks_2 ¦ zone ¦ setting ¦ idc ¦ mdc | 732290
MDC_IDC_SET_ZONE_NUM_SHOCKS_3 | number of shocks_3 ¦ zone ¦ setting ¦ idc ¦ mdc | 732291
MDC_IDC_SET_ZONE_NUM_SHOCKS_4 | number of shocks_4 ¦ zone ¦ setting ¦ idc ¦ mdc | 732292
MDC_IDC_SET_ZONE_NUM_SHOCKS_5 | number of shocks_5 ¦ zone ¦ setting ¦ idc ¦ mdc | 732293
MDC_IDC_SET_ZONE_NUM_SHOCKS_6 | number of shocks_6 ¦ zone ¦ setting ¦ idc ¦ mdc | 732294
MDC_IDC_SET_ZONE_NUM_SHOCKS_7 | number of shocks_7 ¦ zone ¦ setting ¦ idc ¦ mdc | 732295
MDC_IDC_SET_ZONE_NUM_SHOCKS_8 | number of shocks_8 ¦ zone ¦ setting ¦ idc ¦ mdc | 732296
MDC_IDC_SET_ZONE_NUM_SHOCKS_9 | number of shocks_9 ¦ zone ¦ setting ¦ idc ¦ mdc | 732297
MDC_IDC_SET_ZONE_NUM_SHOCKS_10 | number of shocks_10 ¦ zone ¦ setting ¦ idc ¦ mdc | 732298
MDC_IDC_STAT_DTM | date time ¦ statistic ¦ idc ¦ mdc | 737488
MDC_IDC_STAT_DTM_START | date time_start ¦ statistic ¦ idc ¦ mdc | 737489
MDC_IDC_STAT_DTM_END | date time_end ¦ statistic ¦ idc ¦ mdc | 737490
MDC_IDC_STAT_HEART_RATE_DTM | date time ¦ heart rate ¦ statistic ¦ idc ¦ mdc | 737616
MDC_IDC_STAT_HEART_RATE_DTM_START | date time_start ¦ heart rate ¦ statistic ¦ idc ¦ mdc | 737617
MDC_IDC_STAT_HEART_RATE_DTM_END | date time_end ¦ heart rate ¦ statistic ¦ idc ¦ mdc | 737618
MDC_IDC_STAT_HEART_RATE_ATRIAL | atrial ¦ heart rate ¦ statistic ¦ idc ¦ mdc | 737632
MDC_IDC_STAT_HEART_RATE_ATRIAL_MAX | atrial_max ¦ heart rate ¦ statistic ¦ idc ¦ mdc | 737633
MDC_IDC_STAT_HEART_RATE_ATRIAL_MIN | atrial_min ¦ heart rate ¦ statistic ¦ idc ¦ mdc | 737634
MDC_IDC_STAT_HEART_RATE_ATRIAL_MEAN | atrial_mean ¦ heart rate ¦ statistic ¦ idc ¦ mdc | 737635
MDC_IDC_STAT_HEART_RATE_VENTRICULAR | ventricular ¦ heart rate ¦ statistic ¦ idc ¦ mdc | 737648
MDC_IDC_STAT_HEART_RATE_VENTRICULAR_MAX | ventricular_max ¦ heart rate ¦ statistic ¦ idc ¦ mdc | 737649
MDC_IDC_STAT_HEART_RATE_VENTRICULAR_MIN | ventricular_min ¦ heart rate ¦ statistic ¦ idc ¦ mdc | 737650
MDC_IDC_STAT_HEART_RATE_VENTRICULAR_MEAN | ventricular_mean ¦ heart rate ¦ statistic ¦ idc ¦ mdc | 737651
MDC_IDC_STAT_BRADY_DTM | date time ¦ brady ¦ statistic ¦ idc ¦ mdc | 737504
MDC_IDC_STAT_BRADY_DTM_START | date time_start ¦ brady ¦ statistic ¦ idc ¦ mdc | 737505
MDC_IDC_STAT_BRADY_DTM_END | date time_end ¦ brady ¦ statistic ¦ idc ¦ mdc | 737506
MDC_IDC_STAT_BRADY_RA_PERCENT_PACED | right atrial percent paced ¦ brady ¦ statistic ¦ idc ¦ mdc | 737520
MDC_IDC_STAT_BRADY_RV_PERCENT_PACED | right ventricle percent paced ¦ brady ¦ statistic ¦ idc ¦ mdc | 737536
MDC_IDC_STAT_BRADY_AP_VP_PERCENT | ap-vp percent ¦ brady ¦ statistic ¦ idc ¦ mdc | 737552
MDC_IDC_STAT_BRADY_AS_VP_PERCENT | as-vp percent ¦ brady ¦ statistic ¦ idc ¦ mdc | 737568
MDC_IDC_STAT_BRADY_AP_VS_PERCENT | ap-vs percent ¦ brady ¦ statistic ¦ idc ¦ mdc | 737584
MDC_IDC_STAT_BRADY_AS_VS_PERCENT | as-vs percent ¦ brady ¦ statistic ¦ idc ¦ mdc | 737600
MDC_IDC_STAT_AT_DTM | date time ¦ atrial tachy ¦ statistic ¦ idc ¦ mdc | 737664
MDC_IDC_STAT_AT_DTM_START | date time_start ¦ atrial tachy ¦ statistic ¦ idc ¦ mdc | 737665
MDC_IDC_STAT_AT_DTM_END | date time_end ¦ atrial tachy ¦ statistic ¦ idc ¦ mdc | 737666
MDC_IDC_STAT_AT_MODE_SW_MAX_DURATION | mode switch maximum duration ¦ atrial tachy ¦ statistic ¦ idc ¦ mdc | 737680
MDC_IDC_STAT_AT_BURDEN_PERCENT | burden percent ¦ atrial tachy ¦ statistic ¦ idc ¦ mdc | 737696
MDC_IDC_STAT_AT_MODE_SW_PERCENT_TIME | mode switch percent of time ¦ atrial tachy ¦ statistic ¦ idc ¦ mdc | 737712
MDC_IDC_STAT_AT_MODE_SW_PERCENT_TIME_PER_DAY | mode switch percent of time per day ¦ atrial tachy ¦ statistic ¦ idc ¦ mdc | 737728
MDC_IDC_STAT_AT_MODE_SW_COUNT | mode switch count ¦ atrial tachy ¦ statistic ¦ idc ¦ mdc | 737744
MDC_IDC_STAT_AT_MODE_SW_COUNT_PER_DAY | mode switch count per day ¦ atrial tachy ¦ statistic ¦ idc ¦ mdc | 737760
MDC_IDC_STAT_CRT_DTM | date time ¦ cardiac resynchronization therapy ¦ statistic ¦ idc ¦ mdc | 737776
MDC_IDC_STAT_CRT_DTM_START | date time_start ¦ cardiac resynchronization therapy ¦ statistic ¦ idc ¦ mdc | 737777
MDC_IDC_STAT_CRT_DTM_END | date time_end ¦ cardiac resynchronization therapy ¦ statistic ¦ idc ¦ mdc | 737778
MDC_IDC_STAT_CRT_LV_PERCENT_PACED | left ventricle percent paced ¦ cardiac resynchronization therapy ¦ statistic ¦ idc ¦ mdc | 737792
MDC_IDC_STAT_CRT_PERCENT_PACED | percent paced ¦ cardiac resynchronization therapy ¦ statistic ¦ idc ¦ mdc | 737808
MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_DELIVERED_RECENT | recent shocks delivered ¦ tachy therapy ¦ statistic ¦ idc ¦ mdc | 737824
MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_DELIVERED_TOTAL | total shocks delivered ¦ tachy therapy ¦ statistic ¦ idc ¦ mdc | 737840
MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_ABORTED_RECENT | recent shocks aborted ¦ tachy therapy ¦ statistic ¦ idc ¦ mdc | 737856
MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_ABORTED_TOTAL | total shocks aborted ¦ tachy therapy ¦ statistic ¦ idc ¦ mdc | 737872
MDC_IDC_STAT_TACHYTHERAPY_ATP_DELIVERED_RECENT | recent anti-tachycardia pacing delivered ¦ tachy therapy ¦ statistic ¦ idc ¦ mdc | 737888
MDC_IDC_STAT_TACHYTHERAPY_ATP_DELIVERED_TOTAL | total anti-tachycardia pacing delivered ¦ tachy therapy ¦ statistic ¦ idc ¦ mdc | 737904
MDC_IDC_STAT_TACHYTHERAPY_TOTAL_DTM | date time ¦ total ¦ tachy therapy ¦ statistic ¦ idc ¦ mdc | 737920
MDC_IDC_STAT_TACHYTHERAPY_TOTAL_DTM_START | date time_start ¦ total ¦ tachy therapy ¦ statistic ¦ idc ¦ mdc | 737921
MDC_IDC_STAT_TACHYTHERAPY_TOTAL_DTM_END | date time_end ¦ total ¦ tachy therapy ¦ statistic ¦ idc ¦ mdc | 737922
MDC_IDC_STAT_TACHYTHERAPY_RECENT_DTM | date time ¦ recent ¦ tachy therapy ¦ statistic ¦ idc ¦ mdc | 737936
MDC_IDC_STAT_TACHYTHERAPY_RECENT_DTM_START | date time_start ¦ recent ¦ tachy therapy ¦ statistic ¦ idc ¦ mdc | 737937
MDC_IDC_STAT_TACHYTHERAPY_RECENT_DTM_END | date time_end ¦ recent ¦ tachy therapy ¦ statistic ¦ idc ¦ mdc | 737938
MDC_IDC_STAT_EPISODE_TYPE | type ¦ episode ¦ statistic ¦ idc ¦ mdc | 737952 | __MDC_IDC_ENUM_EPISODE_TYPE
MDC_IDC_STAT_EPISODE_TYPE_INDUCED | type induced ¦ episode ¦ statistic ¦ idc ¦ mdc | 737968 | __MDC_IDC_ENUM_EPISODE_TYPE_INDUCED
MDC_IDC_STAT_EPISODE_VENDOR_TYPE | vendor type ¦ episode ¦ statistic ¦ idc ¦ mdc | 737984 | __MDC_IDC_ENUM_EPISODE_VENDOR_TYPE
MDC_IDC_STAT_EPISODE_RECENT_COUNT | recent count ¦ episode ¦ statistic ¦ idc ¦ mdc | 738000
MDC_IDC_STAT_EPISODE_RECENT_COUNT_DTM | date time ¦ recent ¦ episode ¦ statistic ¦ idc ¦ mdc | 738016
MDC_IDC_STAT_EPISODE_RECENT_COUNT_DTM_START | date time_start ¦ recent ¦ episode ¦ statistic ¦ idc ¦ mdc | 738017
MDC_IDC_STAT_EPISODE_RECENT_COUNT_DTM_END | date time_end ¦ recent ¦ episode ¦ statistic ¦ idc ¦ mdc | 738018
MDC_IDC_STAT_EPISODE_TOTAL_COUNT | total count ¦ episode ¦ statistic ¦ idc ¦ mdc | 738032
MDC_IDC_STAT_EPISODE_TOTAL_COUNT_DTM | date time ¦ total ¦ episode ¦ statistic ¦ idc ¦ mdc | 738048
MDC_IDC_STAT_EPISODE_TOTAL_COUNT_DTM_START | date time_start ¦ total ¦ episode ¦ statistic ¦ idc ¦ mdc | 738049
MDC_IDC_STAT_EPISODE_TOTAL_COUNT_DTM_END | date time_end ¦ total ¦ episode ¦ statistic ¦ idc ¦ mdc | 738050
MDC_IDC_EPISODE_ID | identifier ¦ episode ¦ idc ¦ mdc | 739536
MDC_IDC_EPISODE_DTM | date time ¦ episode ¦ idc ¦ mdc | 739552
MDC_IDC_EPISODE_TYPE | type ¦ episode ¦ idc ¦ mdc | 739568 | __MDC_IDC_ENUM_EPISODE_TYPE
MDC_IDC_EPISODE_TYPE_INDUCED | type induced ¦ episode ¦ idc ¦ mdc | 739584 | __MDC_IDC_ENUM_EPISODE_TYPE_INDUCED
MDC_IDC_EPISODE_VENDOR_TYPE | vendor type ¦ episode ¦ idc ¦ mdc | 739600 | __MDC_IDC_ENUM_EPISODE_VENDOR_TYPE
MDC_IDC_EPISODE_ATRIAL_INTERVAL_AT_DETECTION | atrial interval at detection ¦ episode ¦ idc ¦ mdc | 739616
MDC_IDC_EPISODE_ATRIAL_INTERVAL_AT_TERMINATION | atrial interval at termination ¦ episode ¦ idc ¦ mdc | 739632
MDC_IDC_EPISODE_VENTRICULAR_INTERVAL_AT_DETECTION | ventricular interval at detection ¦ episode ¦ idc ¦ mdc | 739648
MDC_IDC_EPISODE_VENTRICULAR_INTERVAL_AT_TERMINATION | ventricular interval at termination ¦ episode ¦ idc ¦ mdc | 739664
MDC_IDC_EPISODE_DETECTION_THERAPY_DETAILS | detection therapy details ¦ episode ¦ idc ¦ mdc | 739680
MDC_IDC_EPISODE_THERAPY_RESULT | therapy result ¦ episode ¦ idc ¦ mdc | 739696 | __MDC_IDC_ENUM_EPISODE_THERAPY_RESULT
MDC_IDC_EPISODE_DURATION | duration ¦ episode ¦ idc ¦ mdc | 739712

## Table IEEEDoc
Reference ID Prepend MDC_IDC_ (Note: Spaces are only included for readability improvements) | Display Name | Definition | Type | For­ mat | Unit of Measure | Enumerator ID Prepend MDC_IDC_ (Note: Spaces are only included for readability improvements) | K | O | C | CC
MDC_IDC | Implantable Cardiac Device System | - | GROUP | - | - | - | - | - | - | -
MDC_IDC_DEV | Implantable Cardiac Device | The group of cardiac device attributes. | GROUP | - | - | - | - | - | 1..1 | - | Do not edit this table
MDC_IDC_DEV_TYPE | Implantable Cardiac Device Type | The type of cardiac device. | Enum | N/A | - | _MDC_IDC_ENUM_PG_TYPE | - | R | - | - | This is an instant automatic copy of all gray columns on the
MDC_IDC_DEV_MODEL | Implantable Cardiac Device Model | The model identifier of a cardiac device. | String | 50 | - | - | Y | R | - | - | IDC Nomenclature tab to be used in the IEEE 11073 document.
MDC_IDC_DEV_SERIAL | Implantable Cardiac Device Serial Number | The serial number of a cardiac device. | String | 40 | - | - | Y | R | - | - | After import into MS Word: Remove all 'MDC_IDC_' and 'mdc ¦ idc ¦ '
MDC_IDC_DEV_MFG | Implantable Cardiac Device Manufacturer | The manufacturer of the cardiac device. | Enum | N/A | - | _MDC_IDC_ENUM_MFG | Y | R | - | - | below the header, replace '_' by '<space>_' and replace
MDC_IDC_DEV_IMPLANT_DT | Implantable Cardiac Device Implant Date | The implant date of the cardiac device. | DateTime | N/A | - | - | - | - | - | - | gray shading of GROUPS by a lighter version for better readability
MDC_IDC_DEV_IMPLANTER | Implantable Cardiac Device Implanter | The name of the physician that implanted the cardiac device. | String | 50 | - | - | - | - | - | -
MDC_IDC_DEV_IMPLANTER_CONTACT_INFO | Implantable Cardiac Device Implanter Contact Information | The contact information of the physician that implanted the cardiac device. | String | 100 | - | - | - | - | - | - | Legend
MDC_IDC_DEV_IMPLANTING_FACILITY | Implantable Cardiac Device Implanting Facility | The facility (clinic / hospital) where the cardiac device was implanted. | String | 100 | - | - | - | - | - | - | K --- Key
MDC_IDC_LEAD | Implantable Lead Attributes | The group of lead attributes. | GROUP | - | - | - | - | - | 0..N | - | O --- Optionality
MDC_IDC_LEAD_MODEL | Implantable Lead Model | The model of the lead. | String | 50 | - | - | Y | RE | - | - | C --- Cardinality
MDC_IDC_LEAD_SERIAL | Implantable Lead Serial Number | The serial number of the lead. | String | 40 | - | - | Y | RE | - | - | CC --- Co Constraint Requirements
MDC_IDC_LEAD_MFG | Implantable Lead Manufacturer | The manufacturer of the lead. | Enum | N/A | - | _MDC_IDC_ENUM_MFG | Y | RE | - | - | Ver --- Version
MDC_IDC_LEAD_IMPLANT_DT | Implantable Lead Implant Date | The implant date of the lead. | DateTime | N/A | - | - | - | - | - | -
MDC_IDC_LEAD_POLARITY_TYPE | Implantable Lead Polarity Type | The number of electrodes on the lead. | Enum | N/A | - | _MDC_IDC_ENUM_LEAD_POLARITY_TYPE | - | - | - | -
MDC_IDC_LEAD_LOCATION | Implantable Lead Location | The fixation location of the lead, usually indicating a chamber of the heart. | Enum | N/A | - | _MDC_IDC_ENUM_LEAD_LOCATION_CHAMBER | - | - | - | -
MDC_IDC_LEAD_LOCATION_DETAIL_1 | Implantable Lead Location Detail 1 | The first word of a combination of words describing further details of the location. | Enum | N/A | - | _MDC_IDC_ENUM_LEAD_LOCATION_DETAIL | - | - | - | -
MDC_IDC_LEAD_LOCATION_DETAIL_2 | Implantable Lead Location Detail 2 | The second word of a combination of words describing further details of the location. | Enum | N/A | - | _MDC_IDC_ENUM_LEAD_LOCATION_DETAIL | - | - | - | -
MDC_IDC_LEAD_LOCATION_DETAIL_3 | Implantable Lead Location Detail 3 | The third word of a combination of words describing further details of the location. | Enum | N/A | - | _MDC_IDC_ENUM_LEAD_LOCATION_DETAIL | - | - | - | -
MDC_IDC_LEAD_CONNECTION_STATUS | Implantable Lead Connection Status | The physical connection status of the lead, either connected or abandoned. | Enum | N/A | - | _MDC_IDC_ENUM_LEAD_STATUS | - | - | - | -
MDC_IDC_LEAD_SPECIAL_FUNCTION | Implantable Lead Special Function | A description of any special attribute or function of the lead. | String | 20 | - | - | - | - | - | -
MDC_IDC_SESS | Interrogation Session | The group of session attributes. | GROUP | - | - | - | - | - | 1..1 | -
MDC_IDC_SESS_DTM | Date Time Interrogation Session | The date and time of the in-clinic or remote interrogation session. | DateTime | N/A | - | - | - | R | - | -
MDC_IDC_SESS_TYPE | Type Interrogation Session | The type of device interaction that generated the current data set. | Enum | N/A | - | _MDC_IDC_ENUM_SESS_TYPE | - | R | - | -
MDC_IDC_SESS_REPROGRAMMED | Re-programmed During Session | The indication of whether the device was re-programmed during the session. | Enum | N/A | - | _MDC_IDC_ENUM_SESS_REPROGRAMMED | - | - | - | -
MDC_IDC_SESS_DTM_PREVIOUS | Date Time Previous Interrogation Session | The date and time of a previous in-clinic or remote interrogation. | DateTime | N/A | - | - | - | - | - | -
MDC_IDC_SESS_TYPE_PREVIOUS | Type Previous Interrogation Session | The type of device interaction of the previous interrogation. | Enum | N/A | - | _MDC_IDC_ENUM_SESS_TYPE | - | - | - | -
MDC_IDC_SESS_REPROGRAMMED_PREVIOUS | Re-programmed During Previous Session | The indication of whether the device was re-programmed along with the previous interrogation. | Enum | N/A | - | _MDC_IDC_ENUM_SESS_REPROGRAMMED | - | - | - | -
MDC_IDC_SESS_CLINICIAN_NAME | Clinician Name | The name of the clinician that is responsible for the examination. | String | 50 | - | - | - | - | - | -
MDC_IDC_SESS_CLINICIAN_CONTACT_INFORMATION | Clinician Contact Information | The contact information for the responsible clinician. | String | 100 | - | - | - | - | - | -
MDC_IDC_SESS_CLINIC_NAME | Clinic Name | The name of the clinic where the examination took place. | String | 50 | - | - | - | - | - | -
MDC_IDC_MSMT | Measurements | The group of measurement attributes. | GROUP | - | - | - | - | - | 0..1 | -
MDC_IDC_MSMT_DTM_[STRTEND] | Measurement Date Time | The date and time of a discrete or a group of measurements. | DateTime | N/A | - | - | - | - | - | -
MDC_IDC_MSMT_BATTERY | Battery Measurements | The group of battery measurement attributes. | GROUP | - | - | - | - | - | 0..N | -
MDC_IDC_MSMT_BATTERY_DTM | Battery Date Time of Measurements | The date and time of the battery measurements. | DateTime | N/A | - | - | - | - | - | -
MDC_IDC_MSMT_BATTERY_STATUS | Battery Status | The different stages of battery depletion. | Enum | N/A | - | _MDC_IDC_ENUM_BATTERY_STATUS | - | - | - | -
MDC_IDC_MSMT_BATTERY_VOLTAGE | Battery Voltage | The measured battery voltage. | Numeric | F6.2 | V | - | - | - | - | -
MDC_IDC_MSMT_BATTERY_IMPEDANCE | Battery Impedance | The measured battery impedance. | Numeric | F6.0 | Ohm | - | - | - | - | -
MDC_IDC_MSMT_BATTERY_REMAINING_LONGEVITY | Battery Remaining Longevity | The estimated amount of battery capacity remaining in months. | Numeric | F6.0 | mo | - | - | - | - | -
MDC_IDC_MSMT_BATTERY_REMAINING_PERCENTAGE | Battery Remaining Percentage | The estimated amount of battery capacity remaining in percent. | Numeric | F3.1 | % | - | - | - | - | -
MDC_IDC_MSMT_BATTERY_RRT_TRIGGER | Battery RRT Trigger | The criteria for determining battery end of service condition. | String | 50 | - | - | - | - | - | -
MDC_IDC_MSMT_CAP | Capacitor Measurements | The group of capacitor measurement attributes. | GROUP | - | - | - | - | - | 0..N | -
MDC_IDC_MSMT_CAP_CHARGE_DTM | Capacitor Last Charge Date Time | The date and time of the capacitor charge. | DateTime | N/A | - | - | - | - | - | -
MDC_IDC_MSMT_CAP_CHARGE_TIME | Capacitor Charge Time | The duration in seconds of the capacitor charge. | Numeric | F6.1 | s | - | - | - | - | -
MDC_IDC_MSMT_CAP_CHARGE_ENERGY | Capacitor Charge Energy | The amount of energy stored after the capacitor charge. | Numeric | F4.0 | J | - | - | - | - | -
MDC_IDC_MSMT_CAP_CHARGE_TYPE | Capacitor Charge Type | The type of the capacitor charge. | Enum | N/A | - | _MDC_IDC_ENUM_CHARGE_TYPE | - | - | - | -
MDC_IDC_MSMT_LEADCHNL_[CHAMBER] | Lead Channel Measurements | The group of measurement attributes per a given chamber. | GROUP | - | - | - | - | - | 0..1 | -
MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_DTM_[STRTEND] | Lead Channel Measurements Date and Time | The date and time of the lead channel measurements. | DateTime | N/A | - | - | - | - | - | -
MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_LEAD_CHANNEL_STATUS | Lead Channel Status | The indication whether to check the lead. | Enum | N/A | - | _MDC_IDC_ENUM_CHANNEL_STATUS | - | - | - | -
MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_SENSING | Lead Channel Sensing Measurements | The group of sensing measurement attributes per a given chamber. | GROUP | - | - | - | - | - | 0..1 | -
MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_SENSING_INTR_AMPL_[MMM] | Lead Channel Sensing Intrinsic Amplitude | The measured amplitude of the intrinsic cardiac signal. | Numeric | F3.2 | mV | - | - | - | - | -
MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_SENSING_POLARITY | Lead Channel Sensing Polarity | The type of polarity used for the intrinsic amplitude measurement. | Enum | N/A | - | _MDC_IDC_ENUM_POLARITY | - | - | - | -
MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_PACING_THRESHOLD | Lead Channel Pacing Threshold Measurements | The group of pacing threshold measurement attributes per a given chamber. | GROUP | - | - | - | - | - | 0..1 | -
MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_PACING_THRESHOLD_AMPLITUDE | Lead Channel Pacing Threshold Amplitude | The minimum pulse amplitude needed for pacing capture for a given pulse width. | Numeric | F3.3 | V | - | - | - | - | -
MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_PACING_THRESHOLD_PULSEWIDTH | Lead Channel Pacing Threshold Pulse Width | The pulse width that was applied when determining the pacing threshold amplitude. | Numeric | F3.2 | ms | - | - | - | - | -
MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_PACING_THRESHOLD_MEASUREMENT_METHOD | Lead Channel Pacing Threshold Measurement Method | The method that was used to obtain the pacing threshold. | Enum | N/A | - | _MDC_IDC_ENUM_MEASUREMENT_METHOD | - | - | - | -
MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_PACING_THRESHOLD_POLARITY | Lead Channel Pacing Threshold Polarity | The type of lead polarity that was used for the pacing threshold measurement. | Enum | N/A | - | _MDC_IDC_ENUM_POLARITY | - | - | - | -
MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_IMPEDANCE | Lead Channel Impedance Measurements | The group of impedance measurement attributes per a given chamber. | GROUP | - | - | - | - | - | 0..1 | -
MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_IMPEDANCE_VALUE | Lead Channel Impedance Value | The sum of the impedances of the lead wires and the electrode-myocardial interface. | Numeric | F6.0 | Ohm | - | - | - | - | -
MDC_IDC_MSMT_LEADCHNL_[CHAMBER]_IMPEDANCE_POLARITY | Lead Channel Impedance Polarity | The type of lead polarity used for measuring the lead impedance. | Enum | N/A | - | _MDC_IDC_ENUM_POLARITY | - | - | - | -
MDC_IDC_MSMT_LEADHVCHNL | Lead High Voltage Channel Measurements | The group of high voltage channel measurement attributes. | GROUP | - | - | - | - | - | 0..N | -
MDC_IDC_MSMT_LEADHVCHNL_DTM_[STRTEND] | Lead High Voltage Channel Date Time | The date and time of the high voltage lead channel measurements. | DateTime | N/A | - | - | - | - | - | -
MDC_IDC_MSMT_LEADHVCHNL_IMPEDANCE | Lead High Voltage Channel Impedance | The sum of the impedances of the shock lead wires and the electrode-myocardial interface. | Numeric | F6.0 | Ohm | - | - | - | - | -
MDC_IDC_MSMT_LEADHVCHNL_MEASUREMENT_TYPE | Lead High Voltage Channel Measurement Type | The electric pulse type used for measuring the shock lead impedance. | Enum | N/A | - | _MDC_IDC_ENUM_HVCHNL_MEASUREMENT_TYPE | - | - | - | -
MDC_IDC_MSMT_LEADHVCHNL_STATUS | Lead High Voltage Channel Status | The indication whether to check the high voltage lead. | Enum | N/A | - | _MDC_IDC_ENUM_CHANNEL_STATUS | - | - | - | -
MDC_IDC_SET | Settings | The group of settings. | GROUP | - | - | - | - | - | 0..1 | -
MDC_IDC_SET_CRT | CRT Settings | The group of CRT settings. | GROUP | - | - | - | - | - | 0..1 | -
MDC_IDC_SET_CRT_LVRV_DELAY | CRT LV-RV Delay | The time between the LV and RV pacing pulses (positive values indicate LV is first). | Numeric | F4.0 | ms | - | - | - | - | -
MDC_IDC_SET_CRT_PACED_CHAMBERS | Ventricular chambers paced during CRT pacing. | The ventricular chambers paced during CRT pacing. | Enum | N/A | - | _MDC_IDC_ENUM_CRT_PACED_CHAMBERS | - | - | - | -
MDC_IDC_SET_MAGNET | Magnet Settings | The group of magnet settings. | GROUP | - | - | - | - | - | 0..1 | -
MDC_IDC_SET_MAGNET_RESP | Magnet Response | A description of how the device responds to a magnet. | String | 100 | - | - | - | - | - | -
MDC_IDC_SET_LEADCHNL_[CHAMBER] | Lead Channel Settings | The group of settings per a given chamber. | GROUP | - | - | - | - | - | 0..1 | -
MDC_IDC_SET_LEADCHNL_[CHAMBER]_SENSING | Lead Channel Settings Sensing | The group of sensing settings per a given chamber. | GROUP | - | - | - | - | - | 0..1 | -
MDC_IDC_SET_LEADCHNL_[CHAMBER]_SENSING_SENSITIVITY | Lead Channel Setting Sensing Sensitivity | The smallest electrical signal programmed to be detected by the device's sensing circuitry. | Numeric | F3.2 | mV | - | - | - | - | -
MDC_IDC_SET_LEADCHNL_[CHAMBER]_SENSING_POLARITY | Lead Channel Setting Sensing Polarity | The indication of unipolar or bipolar configuration for sensing. | Enum | N/A | - | _MDC_IDC_ENUM_POLARITY | - | - | - | -
MDC_IDC_SET_LEADCHNL_[CHAMBER]_SENSING_ANODE_LOCATION_[1..3] | Lead Channel Setting Sensing Anode Location | The location of the electrodes that define the sensing vector's anode, noting that sensing is polarity neutral. An anode may consist of more than one electrode. | Enum | N/A | - | _MDC_IDC_ENUM_ELECTRODE_LOCATION | - | - | - | CC1
MDC_IDC_SET_LEADCHNL_[CHAMBER]_SENSING_ANODE_ELECTRODE_[1..3] | Lead Channel Setting Sensing Anode Terminal | The type of the electrodes that define the sensing vector's anode, noting that sensing is polarity neutral. An anode may consist of more than one electrode. | Enum | N/A | - | _MDC_IDC_ENUM_ELECTRODE_NAME | - | - | - | CC1
MDC_IDC_SET_LEADCHNL_[CHAMBER]_SENSING_CATHODE_LOCATION_[1..3] | Lead Channel Setting Sensing Cathode Location | The location of the electrodes that define the sensing vector's cathode, noting that sensing is polarity neutral. A cathode may consist of more than one electrode. | Enum | N/A | - | _MDC_IDC_ENUM_ELECTRODE_LOCATION | - | - | - | CC1
MDC_IDC_SET_LEADCHNL_[CHAMBER]_SENSING_CATHODE_ELECTRODE_[1..3] | Lead Channel Setting Sensing Cathode Terminal | The type of the electrodes that define the sensing vector's cathode, noting that sensing is polarity neutral. A cathode may consist of more than one electrode. | Enum | N/A | - | _MDC_IDC_ENUM_ELECTRODE_NAME | - | - | - | CC1
MDC_IDC_SET_LEADCHNL_[CHAMBER]_SENSING_ADAPTATION_MODE | Lead Channel Setting Sensing Adaptation Mode | Specifies whether the sensitivity is fixed or adapting. | Enum | N/A | - | _MDC_IDC_ENUM_SENSING_ADAPTATION_MODE | - | - | - | -
MDC_IDC_SET_LEADCHNL_[CHAMBER]_PACING | Lead Channel Settings Pacing | The group of pacing settings per a given chamber. | GROUP | - | - | - | - | - | 0..1 | -
MDC_IDC_SET_LEADCHNL_[CHAMBER]_PACING_AMPLITUDE | Lead Channel Setting Pacing Amplitude | The pacing output amplitude. | Numeric | F3.3 | V | - | - | - | - | -
MDC_IDC_SET_LEADCHNL_[CHAMBER]_PACING_PULSEWIDTH | Lead Channel Setting Pacing Pulse Width | The pacing output pulse width. | Numeric | F3.2 | ms | - | - | - | - | -
MDC_IDC_SET_LEADCHNL_[CHAMBER]_PACING_POLARITY | Lead Channel Setting Pacing Polarity | The indication of unipolar or bipolar configuration for pacing. | Enum | N/A | - | _MDC_IDC_ENUM_POLARITY | - | - | - | -
MDC_IDC_SET_LEADCHNL_[CHAMBER]_PACING_ANODE_LOCATION_[1..3] | Lead Channel Setting Pacing Anode Location | The location of the electrodes that define the pacing vector's anode. An anode may consist of more than one electrode. | Enum | N/A | - | _MDC_IDC_ENUM_ELECTRODE_LOCATION | - | - | - | CC1
MDC_IDC_SET_LEADCHNL_[CHAMBER]_PACING_ANODE_ELECTRODE_[1..3] | Lead Channel Setting Pacing Anode Terminal | The type of the electrodes that define the pacing vector's anode. An anode may consist of more than one electrode. | Enum | N/A | - | _MDC_IDC_ENUM_ELECTRODE_NAME | - | - | - | CC1
MDC_IDC_SET_LEADCHNL_[CHAMBER]_PACING_CATHODE_LOCATION_[1..3] | Lead Channel Setting Pacing Cathode Location | The location of the electrodes that define the pacing vector's cathode. A cathode may consist of more than one electrode. | Enum | N/A | - | _MDC_IDC_ENUM_ELECTRODE_LOCATION | - | - | - | CC1
MDC_IDC_SET_LEADCHNL_[CHAMBER]_PACING_CATHODE_ELECTRODE_[1..3] | Lead Channel Setting Pacing Cathode Terminal | The type of the electrodes that define the pacing vector's cathode. A cathode may consist of more than one electrode. | Enum | N/A | - | _MDC_IDC_ENUM_ELECTRODE_NAME | - | - | - | CC1
MDC_IDC_SET_LEADCHNL_[CHAMBER]_PACING_CAPTURE_MODE | Lead Channel Setting Pacing Capture Mode | The method the device uses for managing pacing capture. | Enum | N/A | - | _MDC_IDC_ENUM_PACING_CAPTURE_MODE | - | - | - | -
MDC_IDC_SET_LEADHVCHNL | Lead High Voltage Channel Setting | The group of high voltage settings. | GROUP | - | - | - | - | - | 0..1 | -
MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR | Lead High Voltage Channel Setting Shock Vector | The group of high voltage shock vector settings. | GROUP | - | - | - | - | - | 0..1 | -
MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_LOCATION_[1..3] | Lead High Voltage Channel Setting Shock Vector Anode Location | The location of the anode electrode that defines the shocking vector for the lead. An anode may consist of more than one electrode. | Enum | N/A | - | _MDC_IDC_ENUM_ELECTRODE_LOCATION | - | - | - | CC1
MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_ELECTRODE_[1..3] | Lead High Voltage Channel Setting Shock Vector Anode Terminal | The anode electrode terminal that defines the shocking vector for the lead. An anode may consist of more than one electrode. | Enum | N/A | - | _MDC_IDC_ENUM_ELECTRODE_NAME | - | - | - | CC1
MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_CATHODE_LOCATION_[1..3] | Lead High Voltage Channel Setting Shock Vector Cathode Location | The location of the cathode electrode that defines the shocking vector for the lead. A cathode may consist of more than one electrode. | Enum | N/A | - | _MDC_IDC_ENUM_ELECTRODE_LOCATION | - | - | - | CC1
MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_CATHODE_ELECTRODE_[1..3] | Lead High Voltage Channel Setting Shock Vector Cathode Terminal | The cathode electrode terminal that defines the shocking vector for the lead. A cathode may consist of more than one electrode. | Enum | N/A | - | _MDC_IDC_ENUM_ELECTRODE_NAME | - | - | - | CC1
MDC_IDC_SET_BRADY | Brady Settings | The group of brady therapy settings. | GROUP | - | - | - | - | - | 0..1 | -
MDC_IDC_SET_BRADY_MODE | Brady Setting Mode (NBG Code) | The brady pacing mode according to the NBG standard. | Enum | N/A | - | _MDC_IDC_ENUM_BRADY_MODE | - | - | - | -
MDC_IDC_SET_BRADY_VENDOR_MODE | Brady Setting Mode (Vendor Specific) | The brady pacing mode as defined by vendor specific codes. | String | 20 | - | - | - | - | - | -
MDC_IDC_SET_BRADY_LOWRATE | Brady Setting Lower Rate Limit | The rate at which the device paces in the absence of a cardiac rhythm faster than the lower rate and without influence from features that can affect the pacing rate. | Numeric | F3.0 | {beats}/min | - | - | - | - | -
MDC_IDC_SET_BRADY_HYSTRATE | Brady Setting Hysteresis Rate | The lower rate for the hysteresis feature. | Numeric | F3.0 | {beats}/min | - | - | - | - | -
MDC_IDC_SET_BRADY_NIGHT_RATE | Brady Setting Night Rate | The lower rate for the night rate feature. | Numeric | F3.0 | {beats}/min | - | - | - | - | -
MDC_IDC_SET_BRADY_SENSOR_TYPE | Brady Setting Sensor Type | The type of sensor applied for rate adaptive pacing. | String | 20 | - | - | - | - | - | -
MDC_IDC_SET_BRADY_MAX_TRACKING_RATE | Brady Setting Maximum Tracking Rate | The fastest atrial rate at which ventricular pacing occurs with 1:1 synchrony. | Numeric | F3.0 | {beats}/min | - | - | - | - | -
MDC_IDC_SET_BRADY_MAX_SENSOR_RATE | Brady Setting Maximum Sensor Rate | The fastest sensor-driven pacing rate that can be achieved in a rate-adaptive pacing system. | Numeric | F3.0 | {beats}/min | - | - | - | - | -
MDC_IDC_SET_BRADY_SAV_DELAY_[HIGHLOW] | Brady Setting SAV Delay | The interval from an intrinsic (sensed) P wave to a paced ventricular event. | Numeric | F4.0 | ms | - | - | - | - | -
MDC_IDC_SET_BRADY_PAV_DELAY_[HIGHLOW] | Brady Setting PAV Delay | The interval from a paced atrial event to a paced ventricular event. | Numeric | F4.0 | ms | - | - | - | - | -
MDC_IDC_SET_BRADY_AT_MODE_SWITCH_MODE | Brady Setting AT Mode Switch Mode | The atrial tachycardia mode switch mode. | Enum | N/A | - | _MDC_IDC_ENUM_BRADY_MODE | - | - | - | -
MDC_IDC_SET_BRADY_AT_MODE_SWITCH_RATE | Brady Setting AT Mode Switch Rate | The atrial tachycardia mode switch rate. | Numeric | F3.0 | {beats}/min | - | - | - | - | -
MDC_IDC_SET_TACHYTHERAPY | Tachy Therapy Settings | The group of high level tachy therapy settings. | GROUP | - | - | - | - | - | 0..1 | -
MDC_IDC_SET_TACHYTHERAPY_VSTAT | Tachy Therapy Setting Ventricular Status | The high level programmed status for ventricular tachy therapies. | Enum | N/A | - | _MDC_IDC_ENUM_THERAPY_STATUS | - | - | - | -
MDC_IDC_SET_TACHYTHERAPY_ASTAT | Tachy Therapy Setting Atrial Status | The high level programmed status for atrial tachy therapies. | Enum | N/A | - | _MDC_IDC_ENUM_THERAPY_STATUS | - | - | - | -
MDC_IDC_SET_ZONE | Zone Settings | The group of zone settings. | GROUP | - | - | - | - | - | 0..N | -
MDC_IDC_SET_ZONE_TYPE | Zone Setting Type Category | The general type of the zone for tachy therapy detection. | Enum | N/A | - | _MDC_IDC_ENUM_ZONE_TYPE | Y | RE | - | -
MDC_IDC_SET_ZONE_VENDOR_TYPE | Zone Setting Vendor Type Category | The vendor specific type of the zone for tachy therapy detection. | Enum | N/A | - | _MDC_IDC_ENUM_ZONE_VENDOR_TYPE | Y | RE | - | -
MDC_IDC_SET_ZONE_STATUS | Zone Setting Status | The status of the zone, whether it is is active or not. | Enum | N/A | - | _MDC_IDC_ENUM_ZONE_STATUS | - | - | - | -
MDC_IDC_SET_ZONE_DETECTION_INTERVAL | Zone Setting Detection Interval | The upper interval limit of the zone for tachy therapy detection. | Numeric | F4.0 | ms | - | - | - | - | -
MDC_IDC_SET_ZONE_DETECTION_BEATS_NUMERATOR | Zone Setting Detection Beats Numerator | The numerator portion of the tachy detection ratio. | Numeric | F3.0 | {beats} | - | - | - | - | CC2
MDC_IDC_SET_ZONE_DETECTION_BEATS_DENOMINATOR | Zone Setting Detection Beats Denominator | The denominator portion of the tachy detection ratio. | Numeric | F3.0 | {beats} | - | - | - | - | CC2
MDC_IDC_SET_ZONE_DETECTION_DETAILS | Zone Setting Detection Details | A text describing arrhythmia detection details. | String | 50 | - | - | - | - | - | -
MDC_IDC_SET_ZONE_TYPE_ATP_[1..10] | Zone Setting ATP Type | The type of anti-tachycardia pacing (ATP) sequences programmed per ATP type. | Enum | N/A | - | _MDC_IDC_ENUM_ATP_TYPE | - | - | - | CC3
MDC_IDC_SET_ZONE_NUM_ATP_SEQS_[1..10] | Zone Setting Number of ATP Sequences | The number of anti-tachycardia pacing pulse (ATP) sequences programmed per ATP type. | Numeric | F3.0 | {seq} | - | - | - | - | CC3
MDC_IDC_SET_ZONE_SHOCK_ENERGY_[1..10] | Zone Setting Shock Energy | The shock energy delivered at a particular sequence of a zone therapy. | Numeric | F4.1 | J | - | - | - | - | CC4
MDC_IDC_SET_ZONE_NUM_SHOCKS_[1..10] | Zone Setting Number of Shocks | The maximum number of shocks set to be delivered at the given shock energy. | Numeric | F3.0 | {shocks} | - | - | - | - | CC4
MDC_IDC_STAT | Statistics | The group of statistics. | GROUP | - | - | - | - | - | 0..1 | -
MDC_IDC_STAT_DTM_[STRTEND] | Statistic Date Time | The date and time period for the aggregated statistics. | DateTime | N/A | - | - | - | RE | - | -
MDC_IDC_STAT_HEART_RATE | Statistic Heart Rate | The group of heart rate statistics. | GROUP | - | - | - | - | - | 0..1 | -
MDC_IDC_STAT_HEART_RATE_DTM_[STRTEND] | Statistic Heart Rate Date Time | The date and time period for the heart rate statistics. | DateTime | N/A | - | - | - | RE | - | CC5
MDC_IDC_STAT_HEART_RATE_ATRIAL_[MMM] | Statistic Atrial Heart Rate | The mean atrial heart rate over the specified time period. | Numeric | F3.0 | {beats}/min | - | - | - | - | -
MDC_IDC_STAT_HEART_RATE_VENTRICULAR_[MMM] | Statistic Ventricular Heart Rate | The mean ventricular heart rate over the specified time period. | Numeric | F3.0 | {beats}/min | - | - | - | - | -
MDC_IDC_STAT_BRADY | Brady Statistics | The group of brady statistics. | GROUP | - | - | - | - | - | 0..1 | -
MDC_IDC_STAT_BRADY_DTM_[STRTEND] | Brady Statistic Date Time | The date time period of the brady statistics. | DateTime | N/A | - | - | - | RE | - | CC5
MDC_IDC_STAT_BRADY_RA_PERCENT_PACED | Brady Statistic RA Percent Paced | The percentage of pacing events in the right atrial chamber over the specified time period. | Numeric | F3.1 | % | - | - | - | - | -
MDC_IDC_STAT_BRADY_RV_PERCENT_PACED | Brady Statistic RV Percent Paced | The percentage of pacing events in the right ventricular chamber over the specified time period. | Numeric | F3.1 | % | - | - | - | - | -
MDC_IDC_STAT_BRADY_AP_VP_PERCENT | Brady Statistic AP VP Percent | The percentage of atrial pace - ventricular pace sequences in relationship to all AV sequences over the specified time period. | Numeric | F3.1 | % | - | - | - | - | -
MDC_IDC_STAT_BRADY_AS_VP_PERCENT | Brady Statistic AS VP Percent | The percentage of atrial sense - ventricular pace sequences in relationship to all AV sequences over the specified time period. | Numeric | F3.1 | % | - | - | - | - | -
MDC_IDC_STAT_BRADY_AP_VS_PERCENT | Brady Statistic AP VS Percent | The percentage of atrial pace - ventricular sense sequences in relationship to all AV sequences over the specified time period. | Numeric | F3.1 | % | - | - | - | - | -
MDC_IDC_STAT_BRADY_AS_VS_PERCENT | Brady Statistic AS VS Percent | The percentage of atrial sense - ventricular sense sequences in relationship to all AV sequences over the specified time period. | Numeric | F3.1 | % | - | - | - | - | -
MDC_IDC_STAT_AT | Atrial Tachy Statistics | The group of atrial tachyarrhythmia statistics. | GROUP | - | - | - | - | - | 0..1 | -
MDC_IDC_STAT_AT_DTM_[STRTEND] | Atrial Tachy Statistic Date Time | The date and time period for the AT statistics. | DateTime | N/A | - | - | - | RE | - | CC5
MDC_IDC_STAT_AT_MODE_SW_MAX_DURATION | Atrial Tachy Statistic Maximum Mode Switch Duration | The maximum contiguous time the device is in AT mode switch over the specified time period. | Numeric | F8.0 | s | - | - | - | - | -
MDC_IDC_STAT_AT_BURDEN_PERCENT | Atrial Tachy Statistic AT/AF Burden Percent | The percent of time in AT AF per day over the specified time period. | Numeric | F3.0 | % | - | - | - | - | -
MDC_IDC_STAT_AT_MODE_SW_PERCENT_TIME | Atrial Tachy Statistic Percent Time In Mode Switch | The percentage of time the device is in AT mode switch over the specified time period. | Numeric | F3.0 | % | - | - | - | - | -
MDC_IDC_STAT_AT_MODE_SW_PERCENT_TIME_PER_DAY | Atrial Tachy Statistic Percent Time In Mode Switch Per Day | The percentage of time per day the device is in AT mode switch over the specified time period. | Numeric | F3.0 | % | - | - | - | - | -
MDC_IDC_STAT_AT_MODE_SW_COUNT | Atrial Tachy Statistic Number Of Mode Switches | The number of mode switches over the specified time period. | Numeric | F8.0 | {switches} | - | - | - | - | -
MDC_IDC_STAT_AT_MODE_SW_COUNT_PER_DAY | Atrial Tachy Statistic Number Of Mode Switches Per Day | The average number of mode switches per day over the specified time period. | Numeric | F8.0 | {switches} | - | - | - | - | -
MDC_IDC_STAT_CRT | CRT Statistics | The group of CRT statistics. | GROUP | - | - | - | - | - | 0..1 | -
MDC_IDC_STAT_CRT_DTM_[STRTEND] | CRT Statistic Date Time | The date and time period for the CRT statistics. | DateTime | N/A | - | - | - | RE | - | CC5
MDC_IDC_STAT_CRT_LV_PERCENT_PACED | CRT Statistic LV Percent Paced | The percentage of LV pacing events in relation to all LV events over the specified time period. | Numeric | F3.1 | % | - | - | - | - | -
MDC_IDC_STAT_CRT_PERCENT_PACED | CRT Statistic CRT Percent Paced | The percentage of LV pacing events due to CRT mode in relation to all ventricular events over the specified time period. | Numeric | F3.1 | % | - | - | - | - | -
MDC_IDC_STAT_TACHYTHERAPY | Tachy Therapy Statistics | The group of tachy therapy statistics. | GROUP | - | - | - | - | - | 0..1 | -
MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_DELIVERED_RECENT | Tachy Therapy Statistic Recent Shocks Delivered | The number of shocks applied over the specified time period. | Numeric | F5.0 | {shocks} | - | - | - | - | -
MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_DELIVERED_TOTAL | Tachy Therapy Statistic Total Shocks Delivered | The total number of shocks applied since implantation or device reset. | Numeric | F5.0 | {shocks} | - | - | - | - | -
MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_ABORTED_RECENT | Tachy Therapy Statistic Recent Shocks Aborted | The number of shocks aborted over the specified time period. | Numeric | F5.0 | {shocks} | - | - | - | - | -
MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_ABORTED_TOTAL | Tachy Therapy Statistic Total Shocks Aborted | The total number of shocks aborted since implantation or device reset. | Numeric | F5.0 | {shocks} | - | - | - | - | -
MDC_IDC_STAT_TACHYTHERAPY_ATP_DELIVERED_RECENT | Tachy Therapy Statistic Recent ATP Delivered | The number of anti-tachycardia pacing (ATP) sequences delivered over the specified time period. | Numeric | F6.0 | {seq} | - | - | - | - | -
MDC_IDC_STAT_TACHYTHERAPY_ATP_DELIVERED_TOTAL | Tachy Therapy Statistic Total ATP Delivered | The total number of anti-tachycardia pacing (ATP) sequences delivered since implantation or device reset. | Numeric | F6.0 | {seq} | - | - | - | - | -
MDC_IDC_STAT_TACHYTHERAPY_TOTAL_DTM_[STRTEND] | Tachy Therapy Statistic Total Date Time | The date and time period for the total tachy therapy statistics. | DateTime | N/A | - | - | - | - | - | -
MDC_IDC_STAT_TACHYTHERAPY_RECENT_DTM_[STRTEND] | Tachy Therapy Statistic Recent Date Time | The date and time period for the recent tachy therapy statistics. | DateTime | N/A | - | - | - | RE | - | CC5
MDC_IDC_STAT_EPISODE | Episode Statistics | The group of episode statistics. | GROUP | - | - | - | - | - | 0..N | -
MDC_IDC_STAT_EPISODE_TYPE | Episode Statistic Type Category | The generic type of the episode being reported. | Enum | N/A | - | _MDC_IDC_ENUM_EPISODE_TYPE | Y | RE | - | -
MDC_IDC_STAT_EPISODE_TYPE_INDUCED | Episode Statistic Type Induced | The indication of whether the episode has been induced or not. | Enum | N/A | - | _MDC_IDC_ENUM_EPISODE_TYPE_INDUCED | - | - | - | -
MDC_IDC_STAT_EPISODE_VENDOR_TYPE | Episode Statistic Vendor Type Category | The vendor type of the episode being reported. | Enum | N/A | - | _MDC_IDC_ENUM_EPISODE_VENDOR_TYPE | Y | RE | - | -
MDC_IDC_STAT_EPISODE_RECENT_COUNT | Episode Statistic Recent Count | The number of episodes per episode type over the specified time period. | Numeric | F7.0 | - | - | - | - | - | -
MDC_IDC_STAT_EPISODE_RECENT_COUNT_DTM_[STRTEND] | Episode Statistic Recent Date Time | The date and time period for the recent episode counts. | DateTime | N/A | - | - | - | RE | - | -
MDC_IDC_STAT_EPISODE_TOTAL_COUNT | Episode Statistic Total Count | The number of episodes per episode type since implantation or device reset. | Numeric | F7.0 | - | - | - | - | - | -
MDC_IDC_STAT_EPISODE_TOTAL_COUNT_DTM_[STRTEND] | Episode Statistic Total Date Time | The date and time period for the total episode counts. | DateTime | N/A | - | - | - | - | - | -
MDC_IDC_EPISODE | Episode | The group of episodes. | GROUP | - | - | - | - | - | 0..N | -
MDC_IDC_EPISODE_ID | Episode Identifier | The unique identifier for the episode. | String | 100 | - | - | Y | RE | - | -
MDC_IDC_EPISODE_DTM | Episode Date Time | The detection date and time of the episode. | DateTime | N/A | - | - | - | - | - | -
MDC_IDC_EPISODE_TYPE | Episode Type Category | The generic type of the episode being reported. | Enum | N/A | - | _MDC_IDC_ENUM_EPISODE_TYPE | - | RE | - | -
MDC_IDC_EPISODE_TYPE_INDUCED | Episode Type Induced Flag | The indication of whether the episode has been induced or not. | Enum | N/A | - | _MDC_IDC_ENUM_EPISODE_TYPE_INDUCED | - | - | - | -
MDC_IDC_EPISODE_VENDOR_TYPE | Episode Vendor Type Category | The vendor type of the episode being reported. | Enum | N/A | - | _MDC_IDC_ENUM_EPISODE_VENDOR_TYPE | - | RE | - | -
MDC_IDC_EPISODE_ATRIAL_INTERVAL_AT_DETECTION | Episode Detection Interval Atrial | The atrial interval at the time of detection. | Numeric | F4.0 | ms | - | - | - | - | -
MDC_IDC_EPISODE_ATRIAL_INTERVAL_AT_TERMINATION | Episode Termination Interval Atrial | The atrial interval at the time of termination. | Numeric | F4.0 | ms | - | - | - | - | -
MDC_IDC_EPISODE_VENTRICULAR_INTERVAL_AT_DETECTION | Episode Detection Interval Ventricular | The ventricular interval at the time of detection. | Numeric | F4.0 | ms | - | - | - | - | -
MDC_IDC_EPISODE_VENTRICULAR_INTERVAL_AT_TERMINATION | Episode Termination Interval Ventricular | The ventricular interval at the time of termination. | Numeric | F4.0 | ms | - | - | - | - | -
MDC_IDC_EPISODE_DETECTION_THERAPY_DETAILS | Episode Detection And Therapy Details | A text describing details about the detection and therapy delivered. | String | 100 | - | - | - | - | - | -
MDC_IDC_EPISODE_THERAPY_RESULT | Episode Therapy Result | The indicator of therapy success. | Enum | N/A | - | _MDC_IDC_ENUM_EPISODE_THERAPY_RESULT | - | - | - | -
MDC_IDC_EPISODE_DURATION | Episode Duration | The duration of the episode in seconds. | Numeric | F8.0 | s | - | - | - | - | -

## Co-Constr.
Co-Constraint-ID | Co-Constraint Requirements
CC1 | Both location and terminal must be provided when ANODE or CATHODE are defined.
CC2 | Detection numerator and denominator must be provided as a pair.
CC3 | Type of ATPs and number of ATP sequences must be provided as a pair.
CC4 | Shock energy and number of shocks must be provided as a pair.
CC5 | A date time range must be provided for a statistical measurements

## Discriminators
ID | Discriminator | Bits | # Identifier | Values | Description | Offset
MDC_IDC_DISC_CHAMBER | CHAMBER | 3 | 1 | RA | Right Atrial | 0
|  |  | 1 | RV | Right Ventricle | 1
|  |  | 1 | LA | Left Atrial | 2
|  |  | 1 | LV | Left Ventricle | 3
|  |  | 1 | MIXED | Multiple Chambers | 4
MDC_IDC_DISC_MMM | MMM | 2 | 2 | MIN | Minimum | 0
|  |  | 2 | MAX | Maximum | 1
|  |  | 2 | MEAN | Mean | 2
MDC_IDC_DISC_HIGHLOW | HIGHLOW | 2 | 3 | HIGH | High | 0
|  |  | 3 | LOW | Low | 1
MDC_IDC_DISC_1_10 | 1..10 | 4 | 4 | 1 | First Instance of Ten | 0
|  |  | 4 | 2 | Second Instance of Ten | 1
|  |  | 4 | 3 | Third Instance of Ten | 2
|  |  | 4 | 4 | Fourth Instance of Ten | 3
|  |  | 4 | 5 | Fifth Instance of Ten | 4
|  |  | 4 | 6 | Sixth Instance of Ten | 5
|  |  | 4 | 7 | Seventh Instance of Ten | 6
|  |  | 4 | 8 | Eighth Instance of Ten | 7
|  |  | 4 | 9 | Ninth Instance of Ten | 8
|  |  | 4 | 10 | Tenth Instance of Ten | 9
MDC_IDC_DISC_1_3 | 1..3 | 3 | 5 | 1 | First Instance of Three | 0
|  |  | 5 | 2 | Second Instance of Three | 1
|  |  | 5 | 3 | Third Instance of Three | 2
MDC_IDC_DISC_STRTEND | STRTEND | 4 | 6 | START | Start Date | 0
|  |  | 6 | END | End Date | 1

## Normative Enumerations
Enumerator Identifier | Enumerator RefID | EnumerationCode [mnemonic] | Display Name | Definition | Version | Code | Enum Sub Table ID | ID in Sub Table | IEEE Enum code [dec]
MDC_IDC_ENUM_DEV_TYPE | MDC_IDC_ENUM_DEV_TYPE_IPG | IPG | Pacemaker | A battery powered implantable device that uses low voltage pulses to stimulate the heart when it is beating too slow. | 1.0.0 | 753665 | 1 | 1 | 32769
| MDC_IDC_ENUM_DEV_TYPE_ICD | ICD | Defibrillator | A battery powered implantable device that uses high voltage energy to shock the heart when it is beating too fast and mostly includes pacemaker functionality. | 1.0.0 | 753666 | 1 | 2 | 32770
| MDC_IDC_ENUM_DEV_TYPE_CRT_D | CRT_D | Cardiac Resynchronization Therapy - Defibrillator | A battery powered implantable device that uses low voltage pulses to stimulate the ventricles of the heart in multiple locations to increase blood pumping efficiency and also includes pacemaker and defibrillator functionality. | 1.0.0 | 753667 | 1 | 3 | 32771
| MDC_IDC_ENUM_DEV_TYPE_CRT_P | CRT_P | Cardiac Resynchronization Therapy - Pacemaker | A battery powered implantable device that uses low voltage pulses to stimulate the ventricles of the heart in multiple locations to increase blood pumping efficiency and also includes pacemaker functionality. | 1.0.0 | 753668 | 1 | 4 | 32772
| MDC_IDC_ENUM_DEV_TYPE_Monitor | Monitor | Implantable Diagnostic Monitor | A battery powered implantable device that monitors and records cardiac activity and does not deliver any therapy. | 1.0.0 | 753669 | 1 | 5 | 32773
| MDC_IDC_ENUM_DEV_TYPE_Other | Other | Other | A device type other than those included in the nomenclature. | 1.0.0 | 753670 | 1 | 6 | 32774
MDC_IDC_ENUM_MFG | MDC_IDC_ENUM_MFG_ANGN | ANGN | Angeion Corp. | A manufacturer name | 1.0.0 | 753729 | 2 | 1 | 32833
| MDC_IDC_ENUM_MFG_APC | APC | American Pacemaker Corp. | A manufacturer name | 1.0.0 | 753730 | 2 | 2 | 32834
| MDC_IDC_ENUM_MFG_BIO | BIO | Biotronik | A manufacturer name | 1.0.0 | 753731 | 2 | 3 | 32835
| MDC_IDC_ENUM_MFG_BSX | BSX | Boston Scientific | A manufacturer name | 1.0.0 | 753732 | 2 | 4 | 32836
| MDC_IDC_ENUM_MFG_CCS | CCS | Cardiac Control Systems | A manufacturer name | 1.0.0 | 753733 | 2 | 5 | 32837
| MDC_IDC_ENUM_MFG_CIM | CIM | Cardiac Impulse | A manufacturer name | 1.0.0 | 753734 | 2 | 6 | 32838
| MDC_IDC_ENUM_MFG_CPM | CPM | Cardio-Pace Medical | A manufacturer name | 1.0.0 | 753735 | 2 | 7 | 32839
| MDC_IDC_ENUM_MFG_CKP | CKP | Cook Pacemaker Corp. | A manufacturer name | 1.0.0 | 753736 | 2 | 8 | 32840
| MDC_IDC_ENUM_MFG_CATM | CATM | Coratomic Inc. | A manufacturer name | 1.0.0 | 753737 | 2 | 9 | 32841
| MDC_IDC_ENUM_MFG_COR | COR | Cordis | A manufacturer name | 1.0.0 | 753738 | 2 | 10 | 32842
| MDC_IDC_ENUM_MFG_ELA | ELA | ELA Medical | A manufacturer name | 1.0.0 | 753739 | 2 | 11 | 32843
| MDC_IDC_ENUM_MFG_GDT | GDT | Guidant | A manufacturer name | 1.0.0 | 753740 | 2 | 12 | 32844
| MDC_IDC_ENUM_MFG_IMC | IMC | Intermedics | A manufacturer name | 1.0.0 | 753741 | 2 | 13 | 32845
| MDC_IDC_ENUM_MFG_IMT | IMT | Implantronik | A manufacturer name | 1.0.0 | 753742 | 2 | 14 | 32846
| MDC_IDC_ENUM_MFG_MCO | MCO | Medico | A manufacturer name | 1.0.0 | 753743 | 2 | 15 | 32847
| MDC_IDC_ENUM_MFG_MDT | MDT | Medtronic | A manufacturer name | 1.0.0 | 753744 | 2 | 16 | 32848
| MDC_IDC_ENUM_MFG_OSC | OSC | Oscor | A manufacturer name | 1.0.0 | 753745 | 2 | 17 | 32849
| MDC_IDC_ENUM_MFG_OSY | OSY | Osypka | A manufacturer name | 1.0.0 | 753746 | 2 | 18 | 32850
| MDC_IDC_ENUM_MFG_PCS | PCS | Pacesetter | A manufacturer name | 1.0.0 | 753747 | 2 | 19 | 32851
| MDC_IDC_ENUM_MFG_SIE | SIE | Siemens | A manufacturer name | 1.0.0 | 753748 | 2 | 20 | 32852
| MDC_IDC_ENUM_MFG_SOM | SOM | Somedics | A manufacturer name | 1.0.0 | 753749 | 2 | 21 | 32853
| MDC_IDC_ENUM_MFG_SOR | SOR | Sorin | A manufacturer name | 1.0.0 | 753750 | 2 | 22 | 32854
| MDC_IDC_ENUM_MFG_STJ | STJ | St.Jude Medical | A manufacturer name | 1.0.0 | 753751 | 2 | 23 | 32855
| MDC_IDC_ENUM_MFG_STO | STO | Stoeckert | A manufacturer name | 1.0.0 | 753752 | 2 | 24 | 32856
| MDC_IDC_ENUM_MFG_TEL | TEL | Teletronics | A manufacturer name | 1.0.0 | 753753 | 2 | 25 | 32857
| MDC_IDC_ENUM_MFG_VEN | VEN | Ventritex | A manufacturer name | 1.0.0 | 753754 | 2 | 26 | 32858
| MDC_IDC_ENUM_MFG_VIT | VIT | Vitatron | A manufacturer name | 1.0.0 | 753755 | 2 | 27 | 32859
| MDC_IDC_ENUM_MFG_Other | Other | Other | A manufacturer other than that contained in the current nomenclature. | 1.0.0 | 753756 | 2 | 28 | 32860
MDC_IDC_ENUM_LEAD_POLARITY_TYPE | MDC_IDC_ENUM_LEAD_POLARITY_TYPE_UNI | UNI | Unipolar Lead | A lead with a single electrode. | 1.0.0 | 753793 | 3 | 1 | 32897
| MDC_IDC_ENUM_LEAD_POLARITY_TYPE_BI | BI | Bipolar Lead | A lead with a two electrodes. | 1.0.0 | 753794 | 3 | 2 | 32898
| MDC_IDC_ENUM_LEAD_POLARITY_TYPE_TRI | TRI | Tripolar Lead | A lead with a three electrodes. | 1.0.0 | 753795 | 3 | 3 | 32899
| MDC_IDC_ENUM_LEAD_POLARITY_TYPE_QUAD | QUAD | Quadripolar Lead | A lead with a four electrodes. | 1.0.0 | 753796 | 3 | 4 | 32900
| MDC_IDC_ENUM_LEAD_POLARITY_TYPE_MULTI | MULTI | Multipolar Lead | A lead with more than four electrodes. | 1.0.0 | 753797 | 3 | 5 | 32901
| MDC_IDC_ENUM_LEAD_POLARITY_TYPE_Unknown | Unknown | Unknown | An indicator showing that polarity type is unknown for the lead. | 1.0.0 | 753798 | 3 | 6 | 32902
MDC_IDC_ENUM_LEAD_LOCATION_CHAMBER | MDC_IDC_ENUM_LEAD_LOCATION_CHAMBER_LA | LA | Left Atrium | The left atrium of the heart. | 1.0.0 | 753857 | 4 | 1 | 32961
| MDC_IDC_ENUM_LEAD_LOCATION_CHAMBER_LV | LV | Left Ventricle | The left ventricle of the heart. | 1.0.0 | 753858 | 4 | 2 | 32962
| MDC_IDC_ENUM_LEAD_LOCATION_CHAMBER_RA | RA | Right Atrium | The right atrium of the heart. | 1.0.0 | 753859 | 4 | 3 | 32963
| MDC_IDC_ENUM_LEAD_LOCATION_CHAMBER_RV | RV | Right Ventricle | The right ventricle of the heart. | 1.0.0 | 753860 | 4 | 4 | 32964
| MDC_IDC_ENUM_LEAD_LOCATION_CHAMBER_OTHER | OTHER | Other | A location other than the heart chambers. | 1.0.0 | 753861 | 4 | 5 | 32965
| MDC_IDC_ENUM_LEAD_LOCATION_CHAMBER_Unknown | Unknown | Unknown | An indicator showing that the lead location is unknown for the lead. | 1.0.0 | 753862 | 4 | 6 | 32966
MDC_IDC_ENUM_LEAD_LOCATION_DETAIL | MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_Anterior | Anterior | Anterior | The front of the heart or further descriptor of the cardiac veins. | 1.0.0 | 753921 | 5 | 1 | 33025
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_Apex | Apex | Apex | The highest or lowest physical location in a heart chamber. | 1.0.0 | 753922 | 5 | 2 | 33026
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_Appendage | Appendage | Appendage | A muscular pouch in the left atrium of the heart. | 1.0.0 | 753923 | 5 | 3 | 33027
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_CardiacVein | CardiacVein | Cardiac Vein | The cardiac vein of the heart. | 1.0.0 | 753924 | 5 | 4 | 33028
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_VenaCava | VenaCava | Vena Cava | The vena cava of the heart. | 1.0.0 | 753925 | 5 | 5 | 33029
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_CoronarySinus | CoronarySinus | Coronary Sinus | A specific cardiac vein of the heart. | 1.0.0 | 753926 | 5 | 6 | 33030
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_Distal | Distal | Distal | A location furthest from a specified reference. | 1.0.0 | 753927 | 5 | 7 | 33031
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_Epicardial | Epicardial | Epicardial | The outside surface of the heart. | 1.0.0 | 753928 | 5 | 8 | 33032
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_FreeWall | FreeWall | Free Wall | A further detail for a location within the heart. | 1.0.0 | 753929 | 5 | 9 | 33033
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_Great | Great | Great | A further descriptor of the cardiac veins. | 1.0.0 | 753930 | 5 | 10 | 33034
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_High | High | High | A further descriptor for a location within the heart. | 1.0.0 | 753931 | 5 | 11 | 33035
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_Lateral | Lateral | Lateral | A further descriptor for a location within the heart. | 1.0.0 | 753932 | 5 | 12 | 33036
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_Left | Left | Left | A further descriptor for a location within the heart. | 1.0.0 | 753933 | 5 | 13 | 33037
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_Low | Low | Low | A further descriptor for a location within the heart. | 1.0.0 | 753934 | 5 | 14 | 33038
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_Marginal | Marginal | Marginal | A further descriptor for a location within the heart. | 1.0.0 | 753935 | 5 | 15 | 33039
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_Middle | Middle | Middle | A further descriptor for a location within the heart. | 1.0.0 | 753936 | 5 | 16 | 33040
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_Myocardial | Myocardial | Myocardial | The muscle tissue of the heart. | 1.0.0 | 753937 | 5 | 17 | 33041
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_OutflowTract | OutflowTract | Outflow Tract | The part of the left ventricle where the blood exits through the pulmonary artery. | 1.0.0 | 753938 | 5 | 18 | 33042
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_Posterior | Posterior | Posterior | The back of the heart or further descriptor of the cardiac veins. | 1.0.0 | 753939 | 5 | 19 | 33043
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_Proximal | Proximal | Proximal | A location nearest to a specified reference. | 1.0.0 | 753940 | 5 | 20 | 33044
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_Right | Right | Right | A further descriptor for a location within the heart. | 1.0.0 | 753941 | 5 | 21 | 33045
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_Septum | Septum | Septum | The walls that separate the chambers of the heart. | 1.0.0 | 753942 | 5 | 22 | 33046
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_Subclavian | Subclavian | Subclavian | A major artery connected to the heart. | 1.0.0 | 753943 | 5 | 23 | 33047
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_Subcutaneous | Subcutaneous | Subcutaneous | A location under the skin. | 1.0.0 | 753944 | 5 | 24 | 33048
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_Superior | Superior | Superior | A specific cardiac vein of the heart. | 1.0.0 | 753945 | 5 | 25 | 33049
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_HISbundle | HISbundle | HIS bundle | The electrical fibers in the heart, inside the septum between the left and right ventricles. | 1.0.0 | 753946 | 5 | 26 | 33050
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_Other | Other | Other | A location other than that which is contained in current nomenclature. | 1.0.0 | 753947 | 5 | 27 | 33051
| MDC_IDC_ENUM_LEAD_LOCATION_DETAIL_Unknown | Unknown | Unknown | An indicator showing that the lead location detail is unknown for the lead. | 1.0.0 | 753948 | 5 | 28 | 33052
MDC_IDC_ENUM_LEAD_STATUS | MDC_IDC_ENUM_LEAD_STATUS_Connected | Connected | Connected | An indicator that a cardiac lead is functional and in use. | 1.0.0 | 753985 | 6 | 1 | 33089
| MDC_IDC_ENUM_LEAD_STATUS_Abandoned | Abandoned | Abandoned | An indicator that a cardiac lead is non-functional and not in use. | 1.0.0 | 753986 | 6 | 2 | 33090
MDC_IDC_ENUM_SESS_TYPE | MDC_IDC_ENUM_SESS_TYPE_Implant | Implant | Implant | An indicator that the data was acquired at the time of the surgical implant. | 1.0.0 | 754049 | 7 | 1 | 33153
| MDC_IDC_ENUM_SESS_TYPE_InClinic | InClinic | In Clinic | An indicator that the data was acquired when the patient was in a hospital or clinic. | 1.0.0 | 754050 | 7 | 2 | 33154
| MDC_IDC_ENUM_SESS_TYPE_Remote | Remote | Remote | An indicator that the data was acquired while the patient was outside a hospital or clinic. | 1.0.0 | 754051 | 7 | 3 | 33155
| MDC_IDC_ENUM_SESS_TYPE_RemoteDeviceInitiated | RemoteDeviceInitiated | Remote Device Initiated | An indicator that the data was acquired while the patient was outside a hospital or clinic using transmission equipment and because of a condition the device detected. | 1.0.0 | 754052 | 7 | 4 | 33156
| MDC_IDC_ENUM_SESS_TYPE_RemoteScheduled | RemoteScheduled | Remote Scheduled | An indicator that the data was acquired while the patient was outside a hospital or clinic using transmission equipment and at a time that was planned in advance. | 1.0.0 | 754053 | 7 | 5 | 33157
| MDC_IDC_ENUM_SESS_TYPE_RemotePatientInitiated | RemotePatientInitiated | Remote Patient Initiated | An indicator that the data was acquired while the patient was outside a hospital or clinic using transmission equipment and because of an action by the patient at a time that was not planned in advance. | 1.0.0 | 754054 | 7 | 6 | 33158
| MDC_IDC_ENUM_SESS_TYPE_Other | Other | Other | An interrogation session type not currently defined within the nomenclature. | 1.0.0 | 754055 | 7 | 7 | 33159
| MDC_IDC_ENUM_SESS_TYPE_Unknown | Unknown | Unknown | An indicator showing that the session type is unknown for the interrogation session. | 1.0.0 | 754056 | 7 | 8 | 33160
MDC_IDC_ENUM_BATTERY_STATUS | MDC_IDC_ENUM_BATTERY_STATUS_BOS | BOS | Beginning of Service | The device battery is at the beginning of its service. | 1.0.0 | 754113 | 8 | 1 | 33217
| MDC_IDC_ENUM_BATTERY_STATUS_EOS | EOS | End of Service | The device battery is depleted. | 1.0.0 | 754114 | 8 | 2 | 33218
| MDC_IDC_ENUM_BATTERY_STATUS_RRT | RRT | Recommended Replacement Time | An indicator that the device should be replaced. | 1.0.0 | 754115 | 8 | 3 | 33219
| MDC_IDC_ENUM_BATTERY_STATUS_MOS | MOS | Middle of Service | The device battery is in use but not depleted or nearing depletion. | 1.0.0 | 754116 | 8 | 4 | 33220
| MDC_IDC_ENUM_BATTERY_STATUS_Unknown | Unknown | Unknown | An indicator showing that the battery status is unknown for the battery. | 1.0.0 | 754117 | 8 | 5 | 33221
MDC_IDC_ENUM_CHARGE_TYPE | MDC_IDC_ENUM_CHARGE_TYPE_Shock | Shock | Shock | The measurement was made when the capacitor was being charged for the purpose of delivering a high voltage shock. | 1.0.0 | 754177 | 9 | 1 | 33281
| MDC_IDC_ENUM_CHARGE_TYPE_Reformation | Reformation | Reformation | The measurement was made when the capacitor was being charged for routine maintenance not involving shock delivery. | 1.0.0 | 754178 | 9 | 2 | 33282
| MDC_IDC_ENUM_CHARGE_TYPE_Unknown | Unknown | Unknown | An indicator showing that the charge type is unknown for the charge. | 1.0.0 | 754179 | 9 | 3 | 33283
MDC_IDC_ENUM_CHANNEL_STATUS | MDC_IDC_ENUM_CHANNEL_STATUS_CheckLead | CheckLead | Check Lead | A notification that the lead appears to be not functioning properly per the device vendor specific criterion. | 1.0.0 | 754241 | 10 | 1 | 33345
| MDC_IDC_ENUM_CHANNEL_STATUS_Null | Null | Null | The absence of a lead status notification. | 1.0.0 | 754242 | 10 | 2 | 33346
MDC_IDC_ENUM_POLARITY | MDC_IDC_ENUM_POLARITY_UNI | UNI | Unipolar | The lead measurement or setting is between one lead electrode and the device can. | 1.0.0 | 754305 | 11 | 1 | 33409
| MDC_IDC_ENUM_POLARITY_BI | BI | Bipolar | The lead measurement or setting is between two lead electrodes. | 1.0.0 | 754306 | 11 | 2 | 33410
| MDC_IDC_ENUM_POLARITY_Unknown | Unknown | Unknown | An indicator showing that the polarity is unknown for the electrode. | 1.0.0 | 754307 | 11 | 3 | 33411
MDC_IDC_ENUM_MEASUREMENT_METHOD | MDC_IDC_ENUM_MEASUREMENT_METHOD_ProgrammerManual | ProgrammerManual | Programmer Manual | A manual measurement made using a programmer. The measurement was made by a clinician programming different device values and noting when changes were observed. | 1.0.0 | 754369 | 12 | 1 | 33473
| MDC_IDC_ENUM_MEASUREMENT_METHOD_ProgrammerAutomatic | ProgrammerAutomatic | Programmer Automatic | The measurement was made by a clinician using an automatic routine within the programmer. | 1.0.0 | 754370 | 12 | 2 | 33474
| MDC_IDC_ENUM_MEASUREMENT_METHOD_DeviceAutomatic | DeviceAutomatic | Device Automatic | The measurement was made by an automatic routine within the device. | 1.0.0 | 754371 | 12 | 3 | 33475
| MDC_IDC_ENUM_MEASUREMENT_METHOD_Unknown | Unknown | Unknown | An indicator showing that the measurement method is unknown. | 1.0.0 | 754372 | 12 | 4 | 33476
MDC_IDC_ENUM_HVCHNL_MEASUREMENT_TYPE | MDC_IDC_ENUM_HVCHNL_MEASUREMENT_TYPE_LowVoltage | LowVoltage | Low Voltage Pulse | The measurement was made using voltage levels associated with a cardiac pacing pulses or lower. | 1.0.0 | 754433 | 13 | 1 | 33537
| MDC_IDC_ENUM_HVCHNL_MEASUREMENT_TYPE_Shock | Shock | Shock | The measurement was made using high voltage energy during delivery of a shock. | 1.0.0 | 754434 | 13 | 2 | 33538
| MDC_IDC_ENUM_HVCHNL_MEASUREMENT_TYPE_Unknown | Unknown | Unknown | An indicator showing that the measurement type is unknown for the high voltage channel measurement. | 1.0.0 | 754435 | 13 | 3 | 33539
MDC_IDC_ENUM_ELECTRODE_LOCATION | MDC_IDC_ENUM_ELECTRODE_LOCATION_RA | RA | Right Atrium | The right atrium of the heart. | 1.0.0 | 754497 | 14 | 1 | 33601
| MDC_IDC_ENUM_ELECTRODE_LOCATION_RV | RV | Right Ventricle | The right ventricle of the heart. | 1.0.0 | 754498 | 14 | 2 | 33602
| MDC_IDC_ENUM_ELECTRODE_LOCATION_LA | LA | Left Atrium | The left atrium of the heart. | 1.0.0 | 754499 | 14 | 3 | 33603
| MDC_IDC_ENUM_ELECTRODE_LOCATION_LV | LV | Left Ventricle | The left ventricle of the heart. | 1.0.0 | 754500 | 14 | 4 | 33604
| MDC_IDC_ENUM_ELECTRODE_LOCATION_SVC | SVC | Superior Vena Cava | The superior vena cava. | 1.0.0 | 754501 | 14 | 5 | 33605
| MDC_IDC_ENUM_ELECTRODE_LOCATION_Other | Other | Other | Other location. | 1.0.0 | 754502 | 14 | 6 | 33606
| MDC_IDC_ENUM_ELECTRODE_LOCATION_Unknown | Unknown | Unknown | An indicator showing that the location is unknown for the electrode. | 1.0.0 | 754503 | 14 | 7 | 33607
MDC_IDC_ENUM_ELECTRODE_NAME | MDC_IDC_ENUM_ELECTRODE_NAME_Tip | Tip | Tip | A tip electrode of the lead. | 1.0.0 | 754561 | 15 | 1 | 33665
| MDC_IDC_ENUM_ELECTRODE_NAME_Ring | Ring | Ring | A ring electrode of the lead. | 1.0.0 | 754562 | 15 | 2 | 33666
| MDC_IDC_ENUM_ELECTRODE_NAME_Ring1 | Ring1 | Ring1 | The first ring electrode of a multi-electrode lead relative to the distal end of the lead. | 1.0.0 | 754563 | 15 | 3 | 33667
| MDC_IDC_ENUM_ELECTRODE_NAME_Ring2 | Ring2 | Ring2 | The second ring electrode of a multi-electrode lead relative to the distal end of the lead. | 1.0.0 | 754564 | 15 | 4 | 33668
| MDC_IDC_ENUM_ELECTRODE_NAME_Ring3 | Ring3 | Ring3 | The third ring electrode of a multi-electrode lead relative to the distal end of the lead. | 1.0.0 | 754565 | 15 | 5 | 33669
| MDC_IDC_ENUM_ELECTRODE_NAME_Ring4 | Ring4 | Ring4 | The fourth ring electrode of a multi-electrode lead relative to the distal end of the lead. | 1.0.0 | 754566 | 15 | 6 | 33670
| MDC_IDC_ENUM_ELECTRODE_NAME_Coil | Coil | Coil | A coil electrode of the lead. | 1.0.0 | 754567 | 15 | 7 | 33671
| MDC_IDC_ENUM_ELECTRODE_NAME_SubQPatch | SubQPatch | Subcutaneous Patch | A subcutaneous patch electrode. | 1.0.0 | 754568 | 15 | 8 | 33672
| MDC_IDC_ENUM_ELECTRODE_NAME_Can | Can | Can | The housing electrode of the device. | 1.0.0 | 754569 | 15 | 9 | 33673
| MDC_IDC_ENUM_ELECTRODE_NAME_Other1 | Other1 | Other1 | Other electrode 1. | 1.0.0 | 754570 | 15 | 10 | 33674
| MDC_IDC_ENUM_ELECTRODE_NAME_Other2 | Other2 | Other2 | Other electrode 2. | 1.0.0 | 754571 | 15 | 11 | 33675
| MDC_IDC_ENUM_ELECTRODE_NAME_Other3 | Other3 | Other3 | Other electrode 3. | 1.0.0 | 754572 | 15 | 12 | 33676
| MDC_IDC_ENUM_ELECTRODE_NAME_Other4 | Other4 | Other4 | Other electrode 4. | 1.0.0 | 754573 | 15 | 13 | 33677
| MDC_IDC_ENUM_ELECTRODE_NAME_Other5 | Other5 | Other5 | Other electrode 5. | 1.0.0 | 754574 | 15 | 14 | 33678
| MDC_IDC_ENUM_ELECTRODE_NAME_Other6 | Other6 | Other6 | Other electrode 6. | 1.0.0 | 754575 | 15 | 15 | 33679
| MDC_IDC_ENUM_ELECTRODE_NAME_Other7 | Other7 | Other7 | Other electrode 7. | 1.0.0 | 754576 | 15 | 16 | 33680
| MDC_IDC_ENUM_ELECTRODE_NAME_Other8 | Other8 | Other8 | Other electrode 8. | 1.0.0 | 754577 | 15 | 17 | 33681
| MDC_IDC_ENUM_ELECTRODE_NAME_Other9 | Other9 | Other9 | Other electrode 9. | 1.0.0 | 754578 | 15 | 18 | 33682
| MDC_IDC_ENUM_ELECTRODE_NAME_Other10 | Other10 | Other10 | Other electrode 10. | 1.0.0 | 754579 | 15 | 19 | 33683
| MDC_IDC_ENUM_ELECTRODE_NAME_Unknown | Unknown | Unknown | An indicator showing that the type is unknown for the electrode. | 1.0.0 | 754580 | 15 | 20 | 33684
MDC_IDC_ENUM_SENSING_ADAPTATION_MODE | MDC_IDC_ENUM_SENSING_ADAPTATION_MODE_AdaptiveSensing | AdaptiveSensing | Adaptive | The sensing threshold changes automatically. | 1.0.0 | 754625 | 16 | 1 | 33729
| MDC_IDC_ENUM_SENSING_ADAPTATION_MODE_FixedSensing | FixedSensing | Fixed | The sensing threshold is fixed. | 1.0.0 | 754626 | 16 | 2 | 33730
MDC_IDC_ENUM_PACING_CAPTURE_MODE | MDC_IDC_ENUM_PACING_CAPTURE_MODE_AdaptiveCapture | AdaptiveCapture | Adaptive | The pacing stimulus changes automatically to maintain capture. | 1.0.0 | 754689 | 17 | 1 | 33793
| MDC_IDC_ENUM_PACING_CAPTURE_MODE_FixedPacing | FixedPacing | Fixed Pacing | The pacing stimulus is fixed. | 1.0.0 | 754690 | 17 | 2 | 33794
| MDC_IDC_ENUM_PACING_CAPTURE_MODE_MonitorCapture | MonitorCapture | Monitor | The device measures the stimulus threshold for cardiac pacing but does not automatically change the stimulus. | 1.0.0 | 754691 | 17 | 3 | 33795
| MDC_IDC_ENUM_PACING_CAPTURE_MODE_Other | Other | Other | A capture mode other than those included in the nomenclature. | 1.0.0 | 754692 | 17 | 4 | 33796
MDC_IDC_ENUM_BRADY_MODE | MDC_IDC_ENUM_BRADY_MODE_AAI | AAI | AAI | The pacing mode per the NBG standard. | 1.0.0 | 754753 | 18 | 1 | 33857
| MDC_IDC_ENUM_BRADY_MODE_AAIR | AAIR | AAIR | The pacing mode per the NBG standard. | 1.0.0 | 754754 | 18 | 2 | 33858
| MDC_IDC_ENUM_BRADY_MODE_AAT | AAT | AAT | The pacing mode per the NBG standard. | 1.0.0 | 754755 | 18 | 3 | 33859
| MDC_IDC_ENUM_BRADY_MODE_AATR | AATR | AATR | The pacing mode per the NBG standard. | 1.0.0 | 754756 | 18 | 4 | 33860
| MDC_IDC_ENUM_BRADY_MODE_AOO | AOO | AOO | The pacing mode per the NBG standard. | 1.0.0 | 754757 | 18 | 5 | 33861
| MDC_IDC_ENUM_BRADY_MODE_AOOR | AOOR | AOOR | The pacing mode per the NBG standard. | 1.0.0 | 754758 | 18 | 6 | 33862
| MDC_IDC_ENUM_BRADY_MODE_DAD | DAD | DAD | The pacing mode per the NBG standard. | 1.0.0 | 754759 | 18 | 7 | 33863
| MDC_IDC_ENUM_BRADY_MODE_DDD | DDD | DDD | The pacing mode per the NBG standard. | 1.0.0 | 754760 | 18 | 8 | 33864
| MDC_IDC_ENUM_BRADY_MODE_DDDR | DDDR | DDDR | The pacing mode per the NBG standard. | 1.0.0 | 754761 | 18 | 9 | 33865
| MDC_IDC_ENUM_BRADY_MODE_DDI | DDI | DDI | The pacing mode per the NBG standard. | 1.0.0 | 754762 | 18 | 10 | 33866
| MDC_IDC_ENUM_BRADY_MODE_DDIR | DDIR | DDIR | The pacing mode per the NBG standard. | 1.0.0 | 754763 | 18 | 11 | 33867
| MDC_IDC_ENUM_BRADY_MODE_DOO | DOO | DOO | The pacing mode per the NBG standard. | 1.0.0 | 754764 | 18 | 12 | 33868
| MDC_IDC_ENUM_BRADY_MODE_DOOR | DOOR | DOOR | The pacing mode per the NBG standard. | 1.0.0 | 754765 | 18 | 13 | 33869
| MDC_IDC_ENUM_BRADY_MODE_DVI | DVI | DVI | The pacing mode per the NBG standard. | 1.0.0 | 754766 | 18 | 14 | 33870
| MDC_IDC_ENUM_BRADY_MODE_DVIR | DVIR | DVIR | The pacing mode per the NBG standard. | 1.0.0 | 754767 | 18 | 15 | 33871
| MDC_IDC_ENUM_BRADY_MODE_VAT | VAT | VAT | The pacing mode per the NBG standard. | 1.0.0 | 754768 | 18 | 16 | 33872
| MDC_IDC_ENUM_BRADY_MODE_VDD | VDD | VDD | The pacing mode per the NBG standard. | 1.0.0 | 754769 | 18 | 17 | 33873
| MDC_IDC_ENUM_BRADY_MODE_VDDR | VDDR | VDDR | The pacing mode per the NBG standard. | 1.0.0 | 754770 | 18 | 18 | 33874
| MDC_IDC_ENUM_BRADY_MODE_VOO | VOO | VOO | The pacing mode per the NBG standard. | 1.0.0 | 754771 | 18 | 19 | 33875
| MDC_IDC_ENUM_BRADY_MODE_VOOR | VOOR | VOOR | The pacing mode per the NBG standard. | 1.0.0 | 754772 | 18 | 20 | 33876
| MDC_IDC_ENUM_BRADY_MODE_VVI | VVI | VVI | The pacing mode per the NBG standard. | 1.0.0 | 754773 | 18 | 21 | 33877
| MDC_IDC_ENUM_BRADY_MODE_VVIR | VVIR | VVIR | The pacing mode per the NBG standard. | 1.0.0 | 754774 | 18 | 22 | 33878
| MDC_IDC_ENUM_BRADY_MODE_VVT | VVT | VVT | The pacing mode per the NBG standard. | 1.0.0 | 754775 | 18 | 23 | 33879
| MDC_IDC_ENUM_BRADY_MODE_OOO | OOO | OOO | The pacing mode per the NBG standard. | 1.0.0 | 754776 | 18 | 24 | 33880
| MDC_IDC_ENUM_BRADY_MODE_OOOR | OOOR | OOOR | The pacing mode per the NBG standard. | 1.0.0 | 754777 | 18 | 25 | 33881
| MDC_IDC_ENUM_BRADY_MODE_VDI | VDI | VDI | The pacing mode per the NBG standard. | 1.0.0 | 754778 | 18 | 26 | 33882
| MDC_IDC_ENUM_BRADY_MODE_VDIR | VDIR | VDIR | The pacing mode per the NBG standard. | 1.0.0 | 754779 | 18 | 27 | 33883
| MDC_IDC_ENUM_BRADY_MODE_ODO | ODO | ODO | The pacing mode per the NBG standard. | 1.0.0 | 754780 | 18 | 28 | 33884
| MDC_IDC_ENUM_BRADY_MODE_VVTR | VVTR | VVTR | The pacing mode per the NBG standard. | 1.0.0 | 754781 | 18 | 29 | 33885
| MDC_IDC_ENUM_BRADY_MODE_DDTR | DDTR | DDTR | The pacing mode per the NBG standard. | 1.0.0 | 754782 | 18 | 30 | 33886
| MDC_IDC_ENUM_BRADY_MODE_OVO | OVO | OVO | The pacing mode per the NBG standard. | 1.0.0 | 754783 | 18 | 31 | 33887
| MDC_IDC_ENUM_BRADY_MODE_DVTR | DVTR | DVTR | The pacing mode per the NBG standard. | 1.0.0 | 754784 | 18 | 32 | 33888
| MDC_IDC_ENUM_BRADY_MODE_DDT | DDT | DDT | The pacing mode per the NBG standard. | 1.0.0 | 754785 | 18 | 33 | 33889
| MDC_IDC_ENUM_BRADY_MODE_VDTR | VDTR | VDTR | The pacing mode per the NBG standard. | 1.0.0 | 754786 | 18 | 34 | 33890
| MDC_IDC_ENUM_BRADY_MODE_VDT | VDT | VDT | The pacing mode per the NBG standard. | 1.0.0 | 754787 | 18 | 35 | 33891
| MDC_IDC_ENUM_BRADY_MODE_ADIR | ADIR | ADIR | The pacing mode per the NBG standard. | 1.0.0 | 754788 | 18 | 36 | 33892
| MDC_IDC_ENUM_BRADY_MODE_ADI | ADI | ADI | The pacing mode per the NBG standard. | 1.0.0 | 754789 | 18 | 37 | 33893
| MDC_IDC_ENUM_BRADY_MODE_Unknown | Unknown | Unknown | An indicator showing that the brady mode is unknown for the brady device. | 1.0.0 | 754790 | 18 | 38 | 33894
MDC_IDC_ENUM_THERAPY_STATUS | MDC_IDC_ENUM_THERAPY_STATUS_On | On | On | An indicator showing that tachy therapies are on. | 1.0.0 | 754817 | 19 | 1 | 33921
| MDC_IDC_ENUM_THERAPY_STATUS_Off | Off | Off | An indicator showing that tachy therapies are off. | 1.0.0 | 754818 | 19 | 2 | 33922
| MDC_IDC_ENUM_THERAPY_STATUS_Unknown | Unknown | Unknown | An indicator showing that the tachy therapy status is unknown for the tachy device. | 1.0.0 | 754819 | 19 | 3 | 33923
MDC_IDC_ENUM_EPISODE_TYPE | MDC_IDC_ENUM_EPISODE_TYPE_Epis_VF | Epis_VF | VF | An episode type of ventricular fibrillation. | 1.0.0 | 754881 | 20 | 1 | 33985
| MDC_IDC_ENUM_EPISODE_TYPE_Epis_VT | Epis_VT | VT | An episode type of ventricular tachycardia. | 1.0.0 | 754882 | 20 | 2 | 33986
| MDC_IDC_ENUM_EPISODE_TYPE_Epis_ATAF | Epis_ATAF | AT/AF | An episode type of atrial tachyarrhythmia without being specific about it being fibrillation, flutter, or tachycardia. | 1.0.0 | 754883 | 20 | 3 | 33987
| MDC_IDC_ENUM_EPISODE_TYPE_Epis_SVT | Epis_SVT | SVT | An episode type of supraventricular tachycardia. | 1.0.0 | 754884 | 20 | 4 | 33988
| MDC_IDC_ENUM_EPISODE_TYPE_Epis_Monitor | Epis_Monitor | Monitor | An episode type that is only being observed. | 1.0.0 | 754885 | 20 | 5 | 33989
| MDC_IDC_ENUM_EPISODE_TYPE_Epis_PeriodicEGM | Epis_PeriodicEGM | Periodic EGM | An episode type that is being observed periodically. | 1.0.0 | 754886 | 20 | 6 | 33990
| MDC_IDC_ENUM_EPISODE_TYPE_Epis_PatientActivated | Epis_PatientActivated | Patient Activated | An episode type that is triggered by the patient. | 1.0.0 | 754887 | 20 | 7 | 33991
| MDC_IDC_ENUM_EPISODE_TYPE_Epis_Other | Epis_Other | Other | An episode that is not included in the enumerations. | 1.0.0 | 754888 | 20 | 8 | 33992
MDC_IDC_ENUM_ZONE_TYPE | MDC_IDC_ENUM_ZONE_TYPE_Zone_VF | Zone_VF | VF | An zone type of ventricular fibrillation. | 1.0.0 | 754945 | 21 | 1 | 34049
| MDC_IDC_ENUM_ZONE_TYPE_Zone_VT | Zone_VT | VT | An zone type of ventricular tachycardia. | 1.0.0 | 754946 | 21 | 2 | 34050
| MDC_IDC_ENUM_ZONE_TYPE_Zone_ATAF | Zone_ATAF | AT/AF | An zone type of atrial tachyarrhythmia without being specific about it being fibrillation, flutter, or tachycardia. | 1.0.0 | 754947 | 21 | 3 | 34051
| MDC_IDC_ENUM_ZONE_TYPE_Zone_Other | Zone_Other | Other | Indicates a selection other than any in the list of selections. | 1.0.0 | 754948 | 21 | 4 | 34052
MDC_IDC_ENUM_ZONE_STATUS | MDC_IDC_ENUM_ZONE_STATUS_Active | Active | Active | An indicator showing that tachyarrhythmia detection and therapy is on for the designated chamber. | 1.0.0 | 755009 | 22 | 1 | 34113
| MDC_IDC_ENUM_ZONE_STATUS_Inactive | Inactive | Inactive | An Indicator showing that tachyarrhythmia detection and therapy is off for the designated chamber. | 1.0.0 | 755010 | 22 | 2 | 34114
| MDC_IDC_ENUM_ZONE_STATUS_Monitor | Monitor | Monitor | An indicator showing that tachyarrhythmia detection is on but therapy is off for the designated chamber. | 1.0.0 | 755011 | 22 | 3 | 34115
| MDC_IDC_ENUM_ZONE_STATUS_Unknown | Unknown | Unknown | An indicator showing that tachyarrhythmia detection an therapy is unknown for the designated chamber. | 1.0.0 | 755012 | 22 | 4 | 34116
MDC_IDC_ENUM_ATP_TYPE | MDC_IDC_ENUM_ATP_TYPE_Burst | Burst | Burst | A tachyarrhythmia therapy that uses pacing stimulus at a fixed pacing rate for a specific number of pulses. | 1.0.0 | 755073 | 23 | 1 | 34177
| MDC_IDC_ENUM_ATP_TYPE_Ramp | Ramp | Ramp | A tachyarrhythmia therapy that uses pacing stimulus at a pacing rate that increases or decreases. | 1.0.0 | 755074 | 23 | 2 | 34178
| MDC_IDC_ENUM_ATP_TYPE_BurstScan | BurstScan | Burst+Scan | A tachyarrhythmia therapy that uses pacing stimulus at a fixed pacing rate with the time between ATP pulses being reduced with each subsequent burst. | 1.0.0 | 755075 | 23 | 3 | 34179
| MDC_IDC_ENUM_ATP_TYPE_RampScan | RampScan | Ramp+Scan | A tachyarrhythmia therapy that uses pacing stimulus at a pacing rate with the time between ATP pulses being reduced with each subsequent burst. | 1.0.0 | 755076 | 23 | 4 | 34180
| MDC_IDC_ENUM_ATP_TYPE_Other | Other | Other | Indicates a selection other than any in the list of selections. | 1.0.0 | 755077 | 23 | 5 | 34181
| MDC_IDC_ENUM_ATP_TYPE_Unknown | Unknown | Unknown | An indicator showing that ATP type is unknown for the ATP pulses. | 1.0.0 | 755078 | 23 | 6 | 34182
MDC_IDC_ENUM_EPISODE_THERAPY_RESULT | MDC_IDC_ENUM_EPISODE_THERAPY_RESULT_Successful | Successful | Successful | An indicator that shows the specific therapy was successful in terminating the tachyarrhythmia. | 1.0.0 | 755137 | 24 | 1 | 34241
| MDC_IDC_ENUM_EPISODE_THERAPY_RESULT_Unsuccessful | Unsuccessful | Unsuccessful | An indicator that shows the specific therapy was not successful in terminating the tachyarrhythmia. | 1.0.0 | 755138 | 24 | 2 | 34242
| MDC_IDC_ENUM_EPISODE_THERAPY_RESULT_Unknown | Unknown | Unknown | An indicator that shows the results of the specific therapy was unkown. | 1.0.0 | 755139 | 24 | 3 | 34243
MDC_IDC_ENUM_SESS_REPROGRAMMED | MDC_IDC_ENUM_SESS_REPROGRAMMED_YES | YES | YES | An indicator that shows the device was re-programmed during the last session. | 1.0.0 | 755201 | 25 | 1 | 34305
| MDC_IDC_ENUM_SESS_REPROGRAMMED_NO | NO | NO | An indicator that shows the device was not re-programmed during the last session. | 1.0.0 | 755202 | 25 | 2 | 34306
| MDC_IDC_ENUM_SESS_REPROGRAMMED_Unknown | Unknown | Unknown | An indicator that shows that device re-programming is unknown. | 1.0.0 | 755203 | 25 | 3 | 34307
MDC_IDC_ENUM_CRT_PACED_CHAMBERS | MDC_IDC_ENUM_CRT_PACED_CHAMBERS_RV_Only | RV_Only | RVOnly | An indicator that shows that only the RV chamber is being paced for resynchronization. | 1.0.0 | 755265 | 26 | 1 | 34369
| MDC_IDC_ENUM_CRT_PACED_CHAMBERS_LV_Only | LV_Only | LVOnly | An indicator that shows that only the LV chamber is being paced for resynchronization. | 1.0.0 | 755266 | 26 | 2 | 34370
| MDC_IDC_ENUM_CRT_PACED_CHAMBERS_BiV | BiV | BiV | An indicator that shows that both the RV and LV chamber are being paced for resynchronization. | 1.0.0 | 755267 | 26 | 3 | 34371
| MDC_IDC_ENUM_CRT_PACED_CHAMBERS_Unknown | Unknown | Unknown | An indicator showing that CRT paced chambers are unknown for the CRT device. | 1.0.0 | 755268 | 26 | 4 | 34372
MDC_IDC_ENUM_EPISODE_TYPE_INDUCED | MDC_IDC_ENUM_EPISODE_TYPE_INDUCED_YES | YES | YES | An indicator that shows that the episode was induced. | 1.0.0 | 755329 | 27 | 1 | 34433
| MDC_IDC_ENUM_EPISODE_TYPE_INDUCED_NO | NO | NO | An indicator that shows that the episode was not induced. | 1.0.0 | 755330 | 27 | 2 | 34434
| MDC_IDC_ENUM_EPISODE_TYPE_INDUCED_Unknown | Unknown | Unknown | An indicator showing that whether the episode was induced or not unknown for the recorded episode. | 1.0.0 | 755331 | 27 | 3 | 34435

## BIO Enumerations
Enumerator Root | Enumerator RefID | EnumerationCode [mnemonic] | Display Name | Definition | Version | Code | Enum Vendor Section | Enum Vendor Offset Lookup | ID in Sub Table | IEEE Enum code [dec]
MDC_IDC_ENUM_EPISODE_VENDOR_TYPE | MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BIO-Epis_VF | BIO-Epis_VF | VF | An episode of ventricular fibrillation (BIOTRONIK). | 1.0.0 | 770049 | 1 | 1 | 1 | 49153
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BIO-Epis_VT | BIO-Epis_VT | VT | An episode of ventricular tachycardia (BIOTRONIK). | 1.0.0 | 770050 | 1 | 1 | 2 | 49154
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BIO-Epis_VT1 | BIO-Epis_VT1 | VT1 | An episode VT1 of ventricular tachycardia (BIOTRONIK). | 1.0.0 | 770051 | 1 | 1 | 3 | 49155
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BIO-Epis_VT2 | BIO-Epis_VT2 | VT2 | An episode VT2 of ventricular tachycardia (BIOTRONIK). | 1.0.0 | 770052 | 1 | 1 | 4 | 49156
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BIO-Epis_NSVT | BIO-Epis_NSVT | NSVT | An episode of a non-sustained ventricular tachycardia (BIOTRONIK). | 1.0.0 | 770053 | 1 | 1 | 5 | 49157
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BIO-Epis_SVT | BIO-Epis_SVT | SVT | An episode of supraventricular tachycardia (BIOTRONIK). | 1.0.0 | 770054 | 1 | 1 | 6 | 49158
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BIO-Epis_ATAF | BIO-Epis_ATAF | AT/AF | An episode of atrial tachyarrhythmia without being specific about it being fibrillation, flutter, or tachycardia (BIOTRONIK). | 1.0.0 | 770055 | 1 | 1 | 7 | 49159
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BIO-Epis_AFlut | BIO-Epis_AFlut | AFlut | An episode of atrial flutter (BIOTRONIK). | 1.0.0 | 770056 | 1 | 1 | 8 | 49160
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BIO-Epis_AFib | BIO-Epis_AFib | AFib | An episode of atrial fibrillation (BIOTRONIK). | 1.0.0 | 770057 | 1 | 1 | 9 | 49161
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BIO-Epis_AT | BIO-Epis_AT | AT | An episode of atrial tachycardia (BIOTRONIK). | 1.0.0 | 770058 | 1 | 1 | 10 | 49162
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BIO-Epis_ModeSw | BIO-Epis_ModeSw | ModeSw | An episode of mode switching (BIOTRONIK). | 1.0.0 | 770059 | 1 | 1 | 11 | 49163
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BIO-Epis_AMon | BIO-Epis_AMon | Atr. Mon | An atrial monitoring episode (BIOTRONIK). | 1.0.0 | 770060 | 1 | 1 | 12 | 49164
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BIO-Epis_VMon | BIO-Epis_VMon | Ven. Mon | A ventricular monitoring episode (BIOTRONIK). | 1.0.0 | 770061 | 1 | 1 | 13 | 49165
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BIO-Epis_HVR | BIO-Epis_HVR | High Ven. Rate | An episode of high ventricular rate (BIOTRONIK). | 1.0.0 | 770062 | 1 | 1 | 14 | 49166
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BIO-Epis_HAR | BIO-Epis_HAR | High Atr. Rate | An episode of high atrial rate (BIOTRONIK). | 1.0.0 | 770063 | 1 | 1 | 15 | 49167
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BIO-Epis_HR | BIO-Epis_HR | High Rate | An episode of high rate (BIOTRONIK). | 1.0.0 | 770064 | 1 | 1 | 16 | 49168
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BIO-Epis_PMT | BIO-Epis_PMT | PMT | An episode of pacemaker mediated tachycardia (BIOTRONIK). | 1.0.0 | 770065 | 1 | 1 | 17 | 49169
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BIO-Epis_PIEGM | BIO-Epis_PIEGM | Periodic IEGM | A periodic IEGM (BIOTRONIK). | 1.0.0 | 770066 | 1 | 1 | 18 | 49170
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BIO-Epis_PAT | BIO-Epis_PAT | Pat. Triggered | A patient triggered IEGM (BIOTRONIK). | 1.0.0 | 770067 | 1 | 1 | 19 | 49171
MDC_IDC_ENUM_ZONE_VENDOR_TYPE | MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BIO-Zone_VF | BIO-Zone_VF | VF | The ventricular fibrillation zone (BIOTRONIK). | 1.0.0 | 770113 | 2 | 1 | 1 | 49217
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BIO-Zone_VT | BIO-Zone_VT | VT | The ventricular tachycardia zone (BIOTRONIK). | 1.0.0 | 770114 | 2 | 1 | 2 | 49218
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BIO-Zone_VT1 | BIO-Zone_VT1 | VT1 | The ventricular tachycardia zone 1 (BIOTRONIK). | 1.0.0 | 770115 | 2 | 1 | 3 | 49219
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BIO-Zone_VT2 | BIO-Zone_VT2 | VT2 | The ventricular tachycardia zone 2 (BIOTRONIK). | 1.0.0 | 770116 | 2 | 1 | 4 | 49220
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BIO-Zone_SVT | BIO-Zone_SVT | SVT | The supraventricular tachycardia zone (BIOTRONIK). | 1.0.0 | 770117 | 2 | 1 | 5 | 49221
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BIO-Zone_ATAF | BIO-Zone_ATAF | AT/AF | The atrial tachycardia / atrial fibrillation zone (BIOTRONIK). | 1.0.0 | 770118 | 2 | 1 | 6 | 49222
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BIO-Zone_AFlut | BIO-Zone_AFlut | AFlut | The atrial flutter zone (BIOTRONIK). | 1.0.0 | 770119 | 2 | 1 | 7 | 49223
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BIO-Zone_AFib | BIO-Zone_AFib | AFib | The atrial fibrillation zone (BIOTRONIK). | 1.0.0 | 770120 | 2 | 1 | 8 | 49224
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BIO-Zone_AT | BIO-Zone_AT | AT | The atrial tachycardia zone (BIOTRONIK). | 1.0.0 | 770121 | 2 | 1 | 9 | 49225
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BIO-Zone_ModeSw | BIO-Zone_ModeSw | ModeSw | The mode switching zone (BIOTRONIK). | 1.0.0 | 770122 | 2 | 1 | 10 | 49226
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BIO-Zone_AMon | BIO-Zone_AMon | Atr. Mon | The atrial monitoring zone (BIOTRONIK). | 1.0.0 | 770123 | 2 | 1 | 11 | 49227
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BIO-Zone_VMon | BIO-Zone_VMon | Ven. Mon | The ventricular monitoring zone (BIOTRONIK). | 1.0.0 | 770124 | 2 | 1 | 12 | 49228
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BIO-Zone_HVR | BIO-Zone_HVR | High Ven. Rate | The high ventricular rate zone (BIOTRONIK). | 1.0.0 | 770125 | 2 | 1 | 13 | 49229
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BIO-Zone_HAR | BIO-Zone_HAR | High Atr. Rate | The high atrial rate zone (BIOTRONIK). | 1.0.0 | 770126 | 2 | 1 | 14 | 49230
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BIO-Zone_HR | BIO-Zone_HR | High Rate | The high rate zone (BIOTRONIK). | 1.0.0 | 770127 | 2 | 1 | 15 | 49231

## BSX Enumerations
Enumerator Root | Enumerator RefID | EnumerationCode [mnemonic] | Display Name | Definition | Version | Code | Enum Vendor Section | Enum Vendor Offset Lookup | ID in Sub Table | IEEE Enum code [dec]
MDC_IDC_ENUM_EPISODE_VENDOR_TYPE | MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BSX-Epis_VF | BSX-Epis_VF | VF | A Boston Scientific episode sub-classification of ventricular fibrillation. | 1.0.0 | 771073 | 1 | 2 | 1 | 50177
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BSX-Epis_VT | BSX-Epis_VT | VT | A Boston Scientific episode sub-classification of ventricular tachycardia. | 1.0.0 | 771074 | 1 | 2 | 2 | 50178
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BSX-Epis_VT-1 | BSX-Epis_VT-1 | VT-1 | A Boston Scientific episode sub-classification of fast ventricular tachycardia. | 1.0.0 | 771075 | 1 | 2 | 3 | 50179
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BSX-Epis_SVT | BSX-Epis_SVT | SVT | A Boston Scientific episode sub-classification of supraventricular tachycardia. | 1.0.0 | 771076 | 1 | 2 | 4 | 50180
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BSX-Epis_NSVT | BSX-Epis_NSVT | NSVT | A Boston Scientific episode sub-classification of non-sustained ventricular tachycardia. | 1.0.0 | 771077 | 1 | 2 | 5 | 50181
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BSX-Epis_ATR | BSX-Epis_ATR | ATR | A Boston Scientific episode sub-classification of atrial tachycardia response. | 1.0.0 | 771078 | 1 | 2 | 6 | 50182
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BSX-Epis_PMT | BSX-Epis_PMT | PMT | A Boston Scientific episode sub-classification of pacemaker mediated tachycardia. | 1.0.0 | 771079 | 1 | 2 | 7 | 50183
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BSX-Epis_PTM | BSX-Epis_PTM | PTM | A Boston Scientific episode sub-classification of patient triggered monitor. | 1.0.0 | 771080 | 1 | 2 | 8 | 50184
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BSX-Epis_Afib | BSX-Epis_Afib | Afib | A Boston Scientific episode sub-classification of atrial fibrillation. | 1.0.0 | 771081 | 1 | 2 | 9 | 50185
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BSX-Epis_AF | BSX-Epis_AF | AF | A Boston Scientific episode sub-classification of atrial fibrillation. | 1.0.0 | 771082 | 1 | 2 | 10 | 50186
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BSX-Epis_AT | BSX-Epis_AT | AT | A Boston Scientific episode sub-classification of atrial tachycardia. | 1.0.0 | 771083 | 1 | 2 | 11 | 50187
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BSX-Epis_RMS | BSX-Epis_RMS | RMS | A Boston Scientific episode sub-classification of reverse mode switch episode. | 1.0.0 | 771084 | 1 | 2 | 12 | 50188
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_BSX-Epis_APMRT | BSX-Epis_APMRT | APMRT | A Boston Scientific episode sub-classification of | 1.0.0 | 771085 | 1 | 2 | 13 | 50189
MDC_IDC_ENUM_ZONE_VENDOR_TYPE | MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BSX-Zone_VT | BSX-Zone_VT | VT | A Boston Scientific zone sub-classification of ventricular tachycardia. | 1.0.0 | 771137 | 2 | 2 | 1 | 50241
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BSX-Zone_VT-1 | BSX-Zone_VT-1 | VT-1 | A Boston Scientific zone sub-classification of fast ventricular tachycardia. | 1.0.0 | 771138 | 2 | 2 | 2 | 50242
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BSX-Zone_VF | BSX-Zone_VF | VF | A Boston Scientific zone sub-classification of fast ventricular fibrillation. | 1.0.0 | 771139 | 2 | 2 | 3 | 50243
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BSX-Zone_SVT | BSX-Zone_SVT | SVT | A Boston Scientific zone sub-classification of supraventricular tachycardia. | 1.0.0 | 771140 | 2 | 2 | 4 | 50244
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BSX-Zone_Afib | BSX-Zone_Afib | Afib | A Boston Scientific zone sub-classification of atrial fibrillation. | 1.0.0 | 771141 | 2 | 2 | 5 | 50245
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BSX-Zone_AF | BSX-Zone_AF | AF | A Boston Scientific zone sub-classification of atrial fibrillation. | 1.0.0 | 771142 | 2 | 2 | 6 | 50246
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_BSX-Zone_AT | BSX-Zone_AT | AT | A Boston Scientific zone sub-classification of atrial tachycardia. | 1.0.0 | 771143 | 2 | 2 | 7 | 50247

## MDT Enumerations
Enumerator Root | Enumerator RefID | Enumeration Code [mnemonic] | Display Name | Definition | Version | Code | Enum Vendor Section | Enum Vendor Offset Lookup | ID in Sub Table | IEEE Enum code [dec]
MDC_IDC_ENUM_EPISODE_VENDOR_TYPE | MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_MDT-Eps_VF | MDT-Eps_VF | VF | A Medtronic episode sub-classification of ventricular fibrillation. | 1.0.0 | 772097 | 1 | 3 | 1 | 51201
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_MDT-Epis_VT | MDT-Epis_VT | VT | A Medtronic episode sub-classification of ventricular tachyarrhythmia. | 1.0.0 | 772098 | 1 | 3 | 2 | 51202
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_MDT-Epis_Fast_VT | MDT-Epis_Fast_VT | Fast VT | A Medtronic episode sub-classification of fast ventricular tachyarrhythmia. | 1.0.0 | 772099 | 1 | 3 | 3 | 51203
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_MDT-Epis_AT_AF | MDT-Epis_AT_AF | AT AF | A Medtronic episode sub-classification of atrial tachyarrhythmia without being specific about it being fibrillation, flutter, or tachycardia. | 1.0.0 | 772100 | 1 | 3 | 4 | 51204
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_MDT-Epis_Fast_AT_AF | MDT-Epis_Fast_AT_AF | Fast AT AF | A Medtronic episode sub-classification of fast atrial tachyarrhythmia without being specific about it being fibrillation, flutter, or tachycardia. | 1.0.0 | 772101 | 1 | 3 | 5 | 51205
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_MDT-Epis_SVT | MDT-Epis_SVT | SVT | A Medtronic episode sub-classification of supraventricular tachycardia. | 1.0.0 | 772102 | 1 | 3 | 6 | 51206
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_MDT-Epis_NS_VT | MDT-Epis_NS_VT | NSVT | A Medtronic episode sub-classification of non-sustained ventricular tachycardia. | 1.0.0 | 772103 | 1 | 3 | 7 | 51207
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_MDT-Epis_Periodic_EGM | MDT-Epis_Periodic_EGM | Periodic EGM | A Medtronic episode sub-classification of periodic EGM recording. | 1.0.0 | 772104 | 1 | 3 | 8 | 51208
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_MDT-Epis_Rate_Drop | MDT-Epis_Rate_Drop | Rate Drop | A Medtronic episode sub-classification of sudden drop in heart rate. | 1.0.0 | 772105 | 1 | 3 | 9 | 51209
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_MDT-Epis_V_Sensing | MDT-Epis_V_Sensing | V Sensing | A Medtronic episode sub-classification of ventricular sensing resulting in inhibition of ventricular pacing therapy. | 1.0.0 | 772106 | 1 | 3 | 10 | 51210
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_MDT-Epis_HR | MDT-Epis_HR | High Rate | A Medtronic episode sub-classification of high heart rate without designation of being atrial or ventricular. | 1.0.0 | 772107 | 1 | 3 | 11 | 51211
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_MDT-Epis_VHR | MDT-Epis_VHR | V High Rate | A Medtronic episode sub-classification of high ventricular rate. | 1.0.0 | 772108 | 1 | 3 | 12 | 51212
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_MDT-Epis_AHR | MDT-Epis_AHR | A High Rate | A Medtronic episode sub-classification of high atrial rate. | 1.0.0 | 772109 | 1 | 3 | 13 | 51213
MDC_IDC_ENUM_ZONE_VENDOR_TYPE | MDC_IDC_ENUM_ZONE_VENDOR_TYPE_MDT-Zone_VF | MDT-Zone_VF | VF | A Medtronic zone sub-classification of ventricular fibrillation. | 1.0.0 | 772161 | 2 | 3 | 1 | 51265
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_MDT-Zone_VT | MDT-Zone_VT | VT | A Medtronic zone sub-classification of ventricular tachyarrhythmia. | 1.0.0 | 772162 | 2 | 3 | 2 | 51266
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_MDT-Zone_FastVT | MDT-Zone_FastVT | FastVT | A Medtronic zone sub-classification of fast ventricular tachyarrhythmia. | 1.0.0 | 772163 | 2 | 3 | 3 | 51267
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_MDT-Zone_FastATAF | MDT-Zone_FastATAF | FastATAF | A Medtronic zone sub-classification of fast atrial tachyarrhythmia without being specific about it being fibrillation, flutter, or tachycardia. | 1.0.0 | 772164 | 2 | 3 | 4 | 51268
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_MDT-Zone_MonVT | MDT-Zone_MonVT | MonVT | A Medtronic zone sub-classification of ventricular tachycardia that is only being monitored with no therapy being delivered. | 1.0.0 | 772165 | 2 | 3 | 5 | 51269
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_MDT-Zone_MonATAF | MDT-Zone_MonATAF | MonATAF | A Medtronic zone sub-classification of atrial tachyarrhythmia without being specific about it being fibrillation, flutter, or tachycardia, that is being monitored with no therapy being delivered. | 1.0.0 | 772166 | 2 | 3 | 6 | 51270
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_MDT-Zone_HR | MDT-Zone_HR | High Rate | A Medtronic zone sub-classification of high heart rate without designation of being atrial or ventricular. | 1.0.0 | 772167 | 2 | 3 | 7 | 51271
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_MDT-Zone_VHR | MDT-Zone_VHR | V High Rate | A Medtronic zone sub-classification of high ventricular rate. | 1.0.0 | 772168 | 2 | 3 | 8 | 51272
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_MDT-Zone_AHR | MDT-Zone_AHR | A High Rate | A Medtronic zone sub-classification of high atrial rate. | 1.0.0 | 772169 | 2 | 3 | 9 | 51273

## STJ Enumerations
Enumerator Root | Enumerator RefID | Enumeration Code [mnemonic] | Display Name | Definition | Version | Code | Enum Vendor Section | Enum Vendor Offset Lookup | ID in Sub Table | IEEE Enum code [dec]
MDC_IDC_ENUM_EPISODE_VENDOR_TYPE | MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_VF | STJ-Epis_VF | VF | A St. Jude Medical episode sub-classification of ventricular fibrillation. | 1.0.0 | 773121 | 1 | 4 | 1 | 52225
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_VT | STJ-Epis_VT | VT | A St. Jude Medical episode sub-classification of ventricular tachyarrhythmia. | 1.0.0 | 773122 | 1 | 4 | 2 | 52226
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_SVT | STJ-Epis_SVT | SVT | A St. Jude Medical episode sub-classification of supraventricular tachycardia. | 1.0.0 | 773123 | 1 | 4 | 3 | 52227
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_VT1 | STJ-Epis_VT1 | VT1 | A St. Jude Medical episode sub-classification of the slowest tachycardia. | 1.0.0 | 773124 | 1 | 4 | 4 | 52228
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_VT2 | STJ-Epis_VT2 | VT2 | A St. Jude Medical episode sub-classification of the faster tachycardia. | 1.0.0 | 773125 | 1 | 4 | 5 | 52229
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_AT_AF | STJ-Epis_AT_AF | AT/AF | A St. Jude Medical episode sub-classification of atrial tachyarrhythmia without being specific about it being fibrillation, flutter, or tachycardia. | 1.0.0 | 773126 | 1 | 4 | 6 | 52230
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_AMS | STJ-Epis_AMS | AMS | A St. Jude Medical episode sub-classification of Auto Mode Switch. | 1.0.0 | 773127 | 1 | 4 | 7 | 52231
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_NoiseReversion | STJ-Epis_NoiseReversion | Noise Reversion | A St. Jude Medical episode sub-classification of Noise Revision. | 1.0.0 | 773128 | 1 | 4 | 8 | 52232
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_PMTTermination | STJ-Epis_PMTTermination | PMT Termination | A St. Jude Medical episode sub-classification of Pacemaker Mediated Tachycardia Termination. | 1.0.0 | 773129 | 1 | 4 | 9 | 52233
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_MagnetResp. | STJ-Epis_MagnetResp. | Magnet Response | A St. Jude Medical episode sub-clasification of Magnet Response. | 1.0.0 | 773130 | 1 | 4 | 10 | 52234
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_EmergencyShock | STJ-Epis_EmergencyShock | Emergency Shock | A St. Jude Medical episode sub-clasification of Emergency Shock. | 1.0.0 | 773131 | 1 | 4 | 11 | 52235
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_MorphTemplate | STJ-Epis_MorphTemplate | Morphology Template | A St. Jude Medical episode sub-clasification of Morphology Template update. | 1.0.0 | 773132 | 1 | 4 | 12 | 52236
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_ReturnToSinus | STJ-Epis_ReturnToSinus | Return to Sinus | A St. Jude Medical episode sub-clasification of Return to Sinus condition. | 1.0.0 | 773133 | 1 | 4 | 13 | 52237
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_HighVRate | STJ-Epis_HighVRate | High V-Rate | A St. Jude Medical episode sub-clasification of High Ventricular Rate. | 1.0.0 | 773134 | 1 | 4 | 14 | 52238
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_HighARate | STJ-Epis_HighARate | High A-Rate | A St. Jude Medical episode sub-clasification of High Atrial Rate. | 1.0.0 | 773135 | 1 | 4 | 15 | 52239
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_AMSEntry | STJ-Epis_AMSEntry | AMS Entry | A St. Jude Medical episode sub-clasification of Entry into Auto-Mode Switch. | 1.0.0 | 773136 | 1 | 4 | 16 | 52240
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_AMSExit | STJ-Epis_AMSExit | AMS Exit | A St. Jude Medical episode sub-clasification of Exit from Auto-Mode Switch. | 1.0.0 | 773137 | 1 | 4 | 17 | 52241
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_AdvHysteresis | STJ-Epis_AdvHysteresis | Advanced Hyseresis | A St. Jude Medical episode sub-clasification of Morphology Template update. | 1.0.0 | 773138 | 1 | 4 | 18 | 52242
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_ConsecutivePVCs | STJ-Epis_ConsecutivePVCs | Consecutive PVCs | A St. Jude Medical episode sub-clasification of Consecutive Premature Ventricular Contractions. | 1.0.0 | 773139 | 1 | 4 | 19 | 52243
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_LossOfCapture | STJ-Epis_LossOfCapture | Loss of Capture | A St. Jude Medical episode sub-clasification of Loss of Capture. | 1.0.0 | 773140 | 1 | 4 | 20 | 52244
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_CannotDetermineSummary | STJ-Epis_CannotDetermineSummary | Cannot Determine Summary | A St. Jude-Medical episode sub-classification with no episode summary. | 1.0.0 | 773141 | 1 | 4 | 21 | 52245
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_VTTimeout | STJ-Epis_VTTimeout | VT Timeout | A St. Jude-Medical episode sub-classification with ventricular tachycardia timeout. | 1.0.0 | 773142 | 1 | 4 | 22 | 52246
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_SVT_VTTimeout | STJ-Epis_SVT_VTTimeout | SVT/VT Timeout | A St. Jude-Medical episode sub-classification with either supraventricular tachycardia or ventricular tachycardia timeout. | 1.0.0 | 773143 | 1 | 4 | 23 | 52247
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_SVTTimeout | STJ-Epis_SVTTimeout | SVT Timeout | A St. Jude-Medical episode sub-classification of induced episode for device-based testing. | 1.0.0 | 773144 | 1 | 4 | 24 | 52248
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_DBT | STJ-Epis_DBT | DBT | A St. Jude Medical episode sub-clasification of DBT.. | 1.0.0 | 773145 | 1 | 4 | 25 | 52249
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_Non-sustained | STJ-Epis_Non-sustained | Non-sustained VT | A St. Jude Medical episode sub-classification of non-sustained VT | 1.0.0 | 773146 | 1 | 4 | 26 | 52250
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_VF_RVLeadNoise | STJ-Epis_VF_RVLeadNoise | RV Lead noise - VF | A St. Jude Medical episode sub-classification of Lead Noise seen as VF | 1.0.0 | 773147 | 1 | 4 | 27 | 52251
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_VT_RVLeadNoise | STJ-Epis_VT_RVLeadNoise | RV Lead noise - VT | A St. Jude Medical episode sub-classification of Lead Noise seen as VT | 1.0.0 | 773148 | 1 | 4 | 28 | 52252
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_VT1_RVLeadNoise | STJ-Epis_VT1_RVLeadNoise | RV Lead noise - VT1 | A St. Jude Medical episode sub-classification of Lead Noise seen as VT1 | 1.0.0 | 773149 | 1 | 4 | 29 | 52253
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_VT2_RVLeadNoise | STJ-Epis_VT2_RVLeadNoise | RV Lead noise - VT2 | A St. Jude Medical episode sub-classification of Lead Noise seen as VT2 | 1.0.0 | 773150 | 1 | 4 | 30 | 52254
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_VF_SecureSenseTimeout | STJ-Epis_VF_SecureSenseTimeout | Secure Sense - VF | A St. Jude Medical episode sub-classification of Secure Sense seen as VF | 1.0.0 | 773151 | 1 | 4 | 31 | 52255
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_VT_SecureSenseTimeout | STJ-Epis_VT_SecureSenseTimeout | Secure Sense Timeout - VT | A St. Jude Medical episode sub-classification of Secure Sense seen as VT | 1.0.0 | 773152 | 1 | 4 | 32 | 52256
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_VT1_SecureSenseTimeout | STJ-Epis_VT1_SecureSenseTimeout | Secure Sense Timeout - VT1 | A St. Jude Medical episode sub-classification of Secure Sense seen as VT1 | 1.0.0 | 773153 | 1 | 4 | 33 | 52257
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_VT2_SecureSenseTimeout | STJ-Epis_VT2_SecureSenseTimeout | Secure Sense Timeout - VT2 | A St. Jude Medical episode sub-classification of Secure Sense seen as VT2 | 1.0.0 | 773154 | 1 | 4 | 34 | 52258
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_SVT_RVLeadNoise | STJ-Epis_SVT_RVLeadNoise | Lead Noise - SVT | A St. Jude Medical episode sub-classification of Lead Noise seen as SVT | 1.0.0 | 773155 | 1 | 4 | 35 | 52259
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_SVTTimeout_RVLeadNoise | STJ-Epis_SVTTimeout_RVLeadNoise | Lead Noise - SVT Timeout | A St. Jude Medical episode sub-classification of Lead Noise seen as SVT timeout | 1.0.0 | 773156 | 1 | 4 | 36 | 52260
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_Non-sustained_RVLeadNoise | STJ-Epis_Non-sustained_RVLeadNoise | Lead Noise - Non-Sustained | A St. Jude Medical episode sub-classification of Lead Noise seen as Non-Sustained VT | 1.0.0 | 773157 | 1 | 4 | 37 | 52261
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_RVLeadNoise | STJ-Epis_RVLeadNoise | RV Lead Noise | A St. Jude Medical episode sub-classification of RV Lead Noise | 1.0.0 | 773158 | 1 | 4 | 38 | 52262
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_STJ-Epis_SecureSenseTimeout | STJ-Epis_SecureSenseTimeout | Secure Sense Timeout | A St. Jude Medical episode sub-classification of Secure Sense Timeout | 1.0.0 | 773159 | 1 | 4 | 39 | 52263
MDC_IDC_ENUM_ZONE_VENDOR_TYPE | MDC_IDC_ENUM_ZONE_VENDOR_TYPE_STJ-Zone_VF | STJ-Zone_VF | VF | A St. Jude Medical zone sub-classification of ventricular fibrillation. | 1.0.0 | 773185 | 2 | 4 | 1 | 52289
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_STJ-Zone_VT | STJ-Zone_VT | VT | A St. Jude Medical zone sub-classification of ventricular tachyarrhythmia. | 1.0.0 | 773186 | 2 | 4 | 2 | 52290
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_STJ-Zone_SVT | STJ-Zone_SVT | SVT | A St. Jude Medical zone sub-classification of supraventricular tachycardia. | 1.0.0 | 773187 | 2 | 4 | 3 | 52291
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_STJ-Zone_VT1 | STJ-Zone_VT1 | VT1 | A St. Jude Medical zone sub-classification of the slowest tachycardia. | 1.0.0 | 773188 | 2 | 4 | 4 | 52292
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_STJ-Zone_VT2 | STJ-Zone_VT2 | VT2 | A St. Jude Medical zone sub-classification of the faster tachycardia. | 1.0.0 | 773189 | 2 | 4 | 5 | 52293
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_STJ-Zone_AT_AF | STJ-Zone_AT_AF | AT/AF | A St. Jude Medical zone sub-classification of atrial tachyarrhythmia without being specific about it being fibrillation, flutter, or tachycardia. | 1.0.0 | 773190 | 2 | 4 | 6 | 52294

## SOR Enumerations
Enumerator Root | Enumerator RefID | EnumerationCode [mnemonic] | Display Name | Definition | Version | Code | Enum Vendor Section | Enum Vendor Offset Lookup | ID in Sub Table | IEEE Enum code [dec]
MDC_IDC_ENUM_EPISODE_VENDOR_TYPE | MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_SOR-Epis_VF | SOR-Epis_VF | VF | Episode of ventricular fibrillation (Sorin). | 1.0.0 | 774145 | 1 | 5 | 1 | 53249
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_SOR-Epis_FastVT | SOR-Epis_FastVT | FastVT | Episode of fast ventricular tachycardia (Sorin). | 1.0.0 | 774146 | 1 | 5 | 2 | 53250
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_SOR-Epis_VT | SOR-Epis_VT | VT | Episode of ventricular tachycardia (Sorin). | 1.0.0 | 774147 | 1 | 5 | 3 | 53251
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_SOR-Epis_SlowVT | SOR-Epis_SlowVT | SlowVT | Episode of slow ventricular tachycardia (Sorin). | 1.0.0 | 774148 | 1 | 5 | 4 | 53252
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_SOR-Epis_SVT | SOR-Epis_SVT | SVT | Episode of supraventricular tachycardia (Sorin). | 1.0.0 | 774149 | 1 | 5 | 5 | 53253
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_SOR-Epis_SVTVT | SOR-Epis_SVTVT | SVT-VT | Episode of treated supraventricular tachycardia (Sorin). | 1.0.0 | 774150 | 1 | 5 | 6 | 53254
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_SOR-Epis_SVTST | SOR-Epis_SVTST | SVT-ST | Episode of non treated supraventricular tachycardia (Sorin). | 1.0.0 | 774151 | 1 | 5 | 7 | 53255
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_SOR-Epis_ModeSwitch | SOR-Epis_ModeSwitch | Mode Switch | Episode of mode switching (Sorin). | 1.0.0 | 774152 | 1 | 5 | 8 | 53256
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_SOR-Epis_NST | SOR-Epis_NST | Non Sustained | Episode of a non-sustained ventricular tachycardia (Sorin). | 1.0.0 | 774153 | 1 | 5 | 9 | 53257
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_SOR-Epis_ABurst | SOR-Epis_ABurst | A Burst | Episode of an atrial burst (Sorin). | 1.0.0 | 774154 | 1 | 5 | 10 | 53258
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_SOR-Epis_VBurst | SOR-Epis_VBurst | V Burst | Episode of a ventricular burst (Sorin). | 1.0.0 | 774155 | 1 | 5 | 11 | 53259
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_SOR-Epis_AVBI | SOR-Epis_AVBI | AVB I | Episode of atrio-ventricular block I in SafeR mode (Sorin). | 1.0.0 | 774156 | 1 | 5 | 12 | 53260
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_SOR-Epis_AVBII | SOR-Epis_AVBII | AVB II | Episode of atrio-ventricular block II in SafeR mode (Sorin). | 1.0.0 | 774157 | 1 | 5 | 13 | 53261
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_SOR-Epis_AVBIII | SOR-Epis_AVBIII | AVB III | Episode of atrio-ventricular block III in SafeR mode (Sorin). | 1.0.0 | 774158 | 1 | 5 | 14 | 53262
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_SOR-Epis_Pause | SOR-Epis_Pause | Pause | Episode of pause in SafeR mode (Sorin). | 1.0.0 | 774159 | 1 | 5 | 15 | 53263
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_SOR-Epis_Security | SOR-Epis_Security | Security | Episode of AVB in security windows in SafeR mode (Sorin). | 1.0.0 | 774160 | 1 | 5 | 16 | 53264
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_SOR-Epis_PMT | SOR-Epis_PMT | PMT | Episode of PMT detection (Sorin). | 1.0.0 | 774161 | 1 | 5 | 17 | 53265
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_SOR-Epis_Impedance | SOR-Epis_Impedance | Impedance | Episode of impedance variation between two impedance tests (Sorin). | 1.0.0 | 774162 | 1 | 5 | 18 | 53266
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_SOR-Epis_VCapture | SOR-Epis_VCapture | V Capture | Episode of threshold variation between two automatic threshold tests (Sorin). | 1.0.0 | 774163 | 1 | 5 | 19 | 53267
| MDC_IDC_ENUM_EPISODE_VENDOR_TYPE_SOR-Epis_FU | SOR-Epis_FU | Follow-up | Episode during of follow-up (Sorin). | 1.0.0 | 774164 | 1 | 5 | 20 | 53268
MDC_IDC_ENUM_ZONE_VENDOR_TYPE | MDC_IDC_ENUM_ZONE_VENDOR_TYPE_SOR-Zone_VF | SOR-Zone_VF | VF | The ventricular fibrillation zone (Sorin). | 1.0.0 | 774209 | 2 | 5 | 1 | 53313
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_SOR-Zone_FastVT | SOR-Zone_FastVT | FastVT | The fast ventricular tachycardia zone (Sorin). | 1.0.0 | 774210 | 2 | 5 | 2 | 53314
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_SOR-Zone_VT | SOR-Zone_VT | VT | The ventricular tachycardia zone (Sorin). | 1.0.0 | 774211 | 2 | 5 | 3 | 53315
| MDC_IDC_ENUM_ZONE_VENDOR_TYPE_SOR-Zone_SlowVT | SOR-Zone_SlowVT | SlowVT | The slow ventricular tachycardia zone (Sorin). | 1.0.0 | 774212 | 2 | 5 | 4 | 53316

## Numeric Code Offsets
Channels | Term Bits | Discriminator Bits | Increment | Channel Size | Channel Offset
MDC_IDC_PG | 6 | 0 | 1 | 64 | 0
MDC_IDC_LEAD | 6 | 0 | 1 | 64 | 64
MDC_IDC_SESS | 6 | 0 | 1 | 64 | 128
MDC_IDC_MSMT | 7 | 6 | 64 | 128 | 192
MDC_IDC_SET | 7 | 6 | 64 | 128 | 8384
MDC_IDC_STAT | 7 | 4 | 16 | 128 | 16576
MDC_IDC_EPISODE | 7 | 4 | 16 | 128 | 18624
Partition Number | 11 |  | 720896
Partition Size | 65536
Total Offset for Nomenclature | 720896
End of Partition | 786432
Normative Enumerations Partition Offset | 32768
Normative Enumerations Partition Size | 16384
Normative Enumerations Table Size | 64
Vendor Enumerations Partition Offset | 49153
Vendor Enumerations Partition Size | 16384
Each Vendor Partition Size | 1024
Each Vendor Table Size | 64 | 16
Vendor | Number | Start | End
Biotronik | 1 | 49153 | 50176
Boston Scientific | 2 | 50177 | 51200
Medtronic | 3 | 51201 | 52224
St. Jude | 4 | 52225 | 53248
Sorin | 5 | 53249 | 54272
Vendor 6 TBD | 6 | 54273 | 55296
Vendor 7 TBD | 7 | 55297 | 56320
Vendor 8 TBD | 8 | 56321 | 57344
Vendor 9 TBD | 9 | 57345 | 58368
Vendor 10 TBD | 10 | 58369 | 59392
Vendor 11 TBD | 11 | 59393 | 60416
Vendor 12 TBD | 12 | 60417 | 61440
Vendor 13 TBD | 13 | 61441 | 62464
Vendor 14 TBD | 14 | 62465 | 63488
Vendor 15 TBD | 15 | 63489 | 64512
Vendor 16 TBD | 16 | 64513 | 65536

## Units of Measure
Unit_of_Measure | Symbol
percent | %
volts | V
millivolts | mV
ohms | Ohm
joules | J
minutes | min
milliseconds | ms
seconds | s
hours | h
months | mo
sequences | {seq}
shocks | {shocks}
mode switches | {switches}
heartbeats | {beats}
heartbeats per minute | {beats}/min

## Codes
Class Type | Metric Type | Units of Measure
Channel | Complex | % | percent
MDS | DateTime | mV | milivolts
Metric | Enum | V | volts
VMD | Numeric | Ohm | ohms
| Structured Numeric | J | joules
| String | ms | miliseconds
|  | s | seconds
|  | min | minutes
|  | h | hours
|  | d | days
|  | mo | months
|  | {seq} | sequences
|  | {shocks} | shocks
|  | {switches} | mode switches
|  | {beats} | heartbeats
|  | {beats}/min | heartbeats per minute

## HL7 Msg Generator
OBX-1 | OBX-2 | OBX-3.1 | OBX-3.2 | OBX-3.3 | OBX-4 | OBX-5.1 | OBX-5.2 | OBX-6 | OBX-8 | OBX-11 | OBX-12 | OBX-14 | OBX-17 | OBX-18 | Message
MSH¦^~\&¦APPNAME¦VENDOR¦CLINIC APP¦CLINIC ID¦20090422152341¦¦ORU^R01^ORU_R01¦12345¦P¦2.6
PID¦¦¦MODEL:XXX/SERIAL:YYY^^^BSC^U~123-12-1234^^^BSC^SS¦¦DOE^JOHN¦¦20070422153118¦M¦¦¦^^^^12345-1234
PV1¦1¦R¦¦¦¦¦MWELBY^Welby^Marcus^A^III^DR¦¦¦¦¦¦¦¦¦¦¦¦12345
OBR¦1¦¦123456¦REMOTE FOLLOW-UP¦¦¦20111025100000¦20111025100000¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦F
NTE¦1¦P¦Remote follow-up observation (a comment can go here concerning the OBR/interrogation section).
1 | CWE | 720897 | MDC_IDC_DEV_TYPE | MDC |  | CRT_D | ^Cardiac Resynchronization Therapy - Defibrillator |  |  | F |  |  |  |  | OBX¦1¦CWE¦720897^MDC_IDC_DEV_TYPE^MDC¦¦CRT_D^Cardiac Resynchronization Therapy - Defibrillator¦¦¦¦¦¦F¦¦¦¦¦¦¦
2 | ST | 720898 | MDC_IDC_DEV_MODEL | MDC |  | Device Model Name |  |  |  | F |  |  |  |  | OBX¦2¦ST¦720898^MDC_IDC_DEV_MODEL^MDC¦¦Device Model Name¦¦¦¦¦¦F¦¦¦¦¦¦¦
3 | ST | 720899 | MDC_IDC_DEV_SERIAL | MDC |  | 5867463524 |  |  |  | F |  |  |  |  | OBX¦3¦ST¦720899^MDC_IDC_DEV_SERIAL^MDC¦¦5867463524¦¦¦¦¦¦F¦¦¦¦¦¦¦
4 | CWE | 720900 | MDC_IDC_DEV_MFG | MDC |  | BIO | ^Biotronik |  |  | F |  |  |  |  | OBX¦4¦CWE¦720900^MDC_IDC_DEV_MFG^MDC¦¦BIO^Biotronik¦¦¦¦¦¦F¦¦¦¦¦¦¦
5 | DTM | 720901 | MDC_IDC_DEV_IMPLANT_DT | MDC |  | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦5¦DTM¦720901^MDC_IDC_DEV_IMPLANT_DT^MDC¦¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
6 | ST | 720902 | MDC_IDC_DEV_IMPLANTER | MDC |  | Dr. Miller |  |  |  | F |  |  |  |  | OBX¦6¦ST¦720902^MDC_IDC_DEV_IMPLANTER^MDC¦¦Dr. Miller¦¦¦¦¦¦F¦¦¦¦¦¦¦
7 | ST | 720903 | MDC_IDC_DEV_IMPLANTER_CONTACT_INFO | MDC |  | Phone: +1 12 345 6789 |  |  |  | F |  |  |  |  | OBX¦7¦ST¦720903^MDC_IDC_DEV_IMPLANTER_CONTACT_INFO^MDC¦¦Phone: +1 12 345 6789¦¦¦¦¦¦F¦¦¦¦¦¦¦
8 | ST | 720904 | MDC_IDC_DEV_IMPLANTING_FACILITY | MDC |  | Main Heart Center New York |  |  |  | F |  |  |  |  | OBX¦8¦ST¦720904^MDC_IDC_DEV_IMPLANTING_FACILITY^MDC¦¦Main Heart Center New York¦¦¦¦¦¦F¦¦¦¦¦¦¦
9 | ST | 720961 | MDC_IDC_LEAD_MODEL | MDC | 1 | Supersense |  |  |  | F |  |  |  |  | OBX¦9¦ST¦720961^MDC_IDC_LEAD_MODEL^MDC¦1¦Supersense¦¦¦¦¦¦F¦¦¦¦¦¦¦
10 | ST | 720962 | MDC_IDC_LEAD_SERIAL | MDC | 1 | 1234567812 |  |  |  | F |  |  |  |  | OBX¦10¦ST¦720962^MDC_IDC_LEAD_SERIAL^MDC¦1¦1234567812¦¦¦¦¦¦F¦¦¦¦¦¦¦
11 | CWE | 720963 | MDC_IDC_LEAD_MFG | MDC | 1 | STJ | ^St.Jude Medical |  |  | F |  |  |  |  | OBX¦11¦CWE¦720963^MDC_IDC_LEAD_MFG^MDC¦1¦STJ^St.Jude Medical¦¦¦¦¦¦F¦¦¦¦¦¦¦
12 | DTM | 720964 | MDC_IDC_LEAD_IMPLANT_DT | MDC | 1 | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦12¦DTM¦720964^MDC_IDC_LEAD_IMPLANT_DT^MDC¦1¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
13 | CWE | 720965 | MDC_IDC_LEAD_POLARITY_TYPE | MDC | 1 | UNI | ^Unipolar Lead |  |  | F |  |  |  |  | OBX¦13¦CWE¦720965^MDC_IDC_LEAD_POLARITY_TYPE^MDC¦1¦UNI^Unipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
14 | CWE | 720966 | MDC_IDC_LEAD_LOCATION | MDC | 1 | RA | ^Right Atrium |  |  | F |  |  |  |  | OBX¦14¦CWE¦720966^MDC_IDC_LEAD_LOCATION^MDC¦1¦RA^Right Atrium¦¦¦¦¦¦F¦¦¦¦¦¦¦
15 | CWE | 720967 | MDC_IDC_LEAD_LOCATION_DETAIL_1 | MDC | 1 | Appendage | ^Appendage |  |  | F |  |  |  |  | OBX¦15¦CWE¦720967^MDC_IDC_LEAD_LOCATION_DETAIL_1^MDC¦1¦Appendage^Appendage¦¦¦¦¦¦F¦¦¦¦¦¦¦
16 | CWE | 720970 | MDC_IDC_LEAD_CONNECTION_STATUS | MDC | 1 | Connected | ^Connected |  |  | F |  |  |  |  | OBX¦16¦CWE¦720970^MDC_IDC_LEAD_CONNECTION_STATUS^MDC¦1¦Connected^Connected¦¦¦¦¦¦F¦¦¦¦¦¦¦
17 | ST | 720971 | MDC_IDC_LEAD_SPECIAL_FUNCTION | MDC | 1 | Pressure Sensor |  |  |  | F |  |  |  |  | OBX¦17¦ST¦720971^MDC_IDC_LEAD_SPECIAL_FUNCTION^MDC¦1¦Pressure Sensor¦¦¦¦¦¦F¦¦¦¦¦¦¦
18 | ST | 720961 | MDC_IDC_LEAD_MODEL | MDC | 2 | Supersense |  |  |  | F |  |  |  |  | OBX¦18¦ST¦720961^MDC_IDC_LEAD_MODEL^MDC¦2¦Supersense¦¦¦¦¦¦F¦¦¦¦¦¦¦
19 | ST | 720962 | MDC_IDC_LEAD_SERIAL | MDC | 2 | 1234567813 |  |  |  | F |  |  |  |  | OBX¦19¦ST¦720962^MDC_IDC_LEAD_SERIAL^MDC¦2¦1234567813¦¦¦¦¦¦F¦¦¦¦¦¦¦
20 | CWE | 720963 | MDC_IDC_LEAD_MFG | MDC | 2 | BIO | ^Biotronik |  |  | F |  |  |  |  | OBX¦20¦CWE¦720963^MDC_IDC_LEAD_MFG^MDC¦2¦BIO^Biotronik¦¦¦¦¦¦F¦¦¦¦¦¦¦
21 | DTM | 720964 | MDC_IDC_LEAD_IMPLANT_DT | MDC | 2 | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦21¦DTM¦720964^MDC_IDC_LEAD_IMPLANT_DT^MDC¦2¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
22 | CWE | 720965 | MDC_IDC_LEAD_POLARITY_TYPE | MDC | 2 | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦22¦CWE¦720965^MDC_IDC_LEAD_POLARITY_TYPE^MDC¦2¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
23 | CWE | 720966 | MDC_IDC_LEAD_LOCATION | MDC | 2 | RV | ^Right Ventricle |  |  | F |  |  |  |  | OBX¦23¦CWE¦720966^MDC_IDC_LEAD_LOCATION^MDC¦2¦RV^Right Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
24 | CWE | 720967 | MDC_IDC_LEAD_LOCATION_DETAIL_1 | MDC | 2 | Apex | ^Apex |  |  | F |  |  |  |  | OBX¦24¦CWE¦720967^MDC_IDC_LEAD_LOCATION_DETAIL_1^MDC¦2¦Apex^Apex¦¦¦¦¦¦F¦¦¦¦¦¦¦
25 | CWE | 720970 | MDC_IDC_LEAD_CONNECTION_STATUS | MDC | 2 | Connected | ^Connected |  |  | F |  |  |  |  | OBX¦25¦CWE¦720970^MDC_IDC_LEAD_CONNECTION_STATUS^MDC¦2¦Connected^Connected¦¦¦¦¦¦F¦¦¦¦¦¦¦
26 | ST | 720971 | MDC_IDC_LEAD_SPECIAL_FUNCTION | MDC | 2 | N/A |  |  |  | F |  |  |  |  | OBX¦26¦ST¦720971^MDC_IDC_LEAD_SPECIAL_FUNCTION^MDC¦2¦N/A¦¦¦¦¦¦F¦¦¦¦¦¦¦
27 | ST | 720961 | MDC_IDC_LEAD_MODEL | MDC | 3 | Supersense |  |  |  | F |  |  |  |  | OBX¦27¦ST¦720961^MDC_IDC_LEAD_MODEL^MDC¦3¦Supersense¦¦¦¦¦¦F¦¦¦¦¦¦¦
28 | ST | 720962 | MDC_IDC_LEAD_SERIAL | MDC | 3 | 1234567814 |  |  |  | F |  |  |  |  | OBX¦28¦ST¦720962^MDC_IDC_LEAD_SERIAL^MDC¦3¦1234567814¦¦¦¦¦¦F¦¦¦¦¦¦¦
29 | CWE | 720963 | MDC_IDC_LEAD_MFG | MDC | 3 | BSX | ^Boston Scientific |  |  | F |  |  |  |  | OBX¦29¦CWE¦720963^MDC_IDC_LEAD_MFG^MDC¦3¦BSX^Boston Scientific¦¦¦¦¦¦F¦¦¦¦¦¦¦
30 | DTM | 720964 | MDC_IDC_LEAD_IMPLANT_DT | MDC | 3 | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦30¦DTM¦720964^MDC_IDC_LEAD_IMPLANT_DT^MDC¦3¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
31 | CWE | 720965 | MDC_IDC_LEAD_POLARITY_TYPE | MDC | 3 | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦31¦CWE¦720965^MDC_IDC_LEAD_POLARITY_TYPE^MDC¦3¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
32 | CWE | 720966 | MDC_IDC_LEAD_LOCATION | MDC | 3 | LV | ^Left Ventricle |  |  | F |  |  |  |  | OBX¦32¦CWE¦720966^MDC_IDC_LEAD_LOCATION^MDC¦3¦LV^Left Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
33 | CWE | 720967 | MDC_IDC_LEAD_LOCATION_DETAIL_1 | MDC | 3 | FreeWall | ^Free Wall |  |  | F |  |  |  |  | OBX¦33¦CWE¦720967^MDC_IDC_LEAD_LOCATION_DETAIL_1^MDC¦3¦FreeWall^Free Wall¦¦¦¦¦¦F¦¦¦¦¦¦¦
34 | CWE | 720970 | MDC_IDC_LEAD_CONNECTION_STATUS | MDC | 3 | Connected | ^Connected |  |  | F |  |  |  |  | OBX¦34¦CWE¦720970^MDC_IDC_LEAD_CONNECTION_STATUS^MDC¦3¦Connected^Connected¦¦¦¦¦¦F¦¦¦¦¦¦¦
35 | ST | 720971 | MDC_IDC_LEAD_SPECIAL_FUNCTION | MDC | 3 | N/A |  |  |  | F |  |  |  |  | OBX¦35¦ST¦720971^MDC_IDC_LEAD_SPECIAL_FUNCTION^MDC¦3¦N/A¦¦¦¦¦¦F¦¦¦¦¦¦¦
36 | DTM | 721025 | MDC_IDC_SESS_DTM | MDC |  | 20111025100000 |  |  |  | F |  |  |  |  | OBX¦36¦DTM¦721025^MDC_IDC_SESS_DTM^MDC¦¦20111025100000¦¦¦¦¦¦F¦¦¦¦¦¦¦
37 | CWE | 721026 | MDC_IDC_SESS_TYPE | MDC |  | Remote | ^Remote |  |  | F |  |  |  |  | OBX¦37¦CWE¦721026^MDC_IDC_SESS_TYPE^MDC¦¦Remote^Remote ¦¦¦¦¦¦F¦¦¦¦¦¦¦
38 | CWE | 721027 | MDC_IDC_SESS_REPROGRAMMED | MDC |  | NO | ^NO |  |  | F |  |  |  |  | OBX¦38¦CWE¦721027^MDC_IDC_SESS_REPROGRAMMED^MDC¦¦NO^NO¦¦¦¦¦¦F¦¦¦¦¦¦¦
39 | DTM | 721028 | MDC_IDC_SESS_DTM_PREVIOUS | MDC |  | 20090925100000 |  |  |  | F |  |  |  |  | OBX¦39¦DTM¦721028^MDC_IDC_SESS_DTM_PREVIOUS^MDC¦¦20090925100000¦¦¦¦¦¦F¦¦¦¦¦¦¦
40 | CWE | 721029 | MDC_IDC_SESS_TYPE_PREVIOUS | MDC |  | InClinic | ^In Clinic |  |  | F |  |  |  |  | OBX¦40¦CWE¦721029^MDC_IDC_SESS_TYPE_PREVIOUS^MDC¦¦InClinic^In Clinic¦¦¦¦¦¦F¦¦¦¦¦¦¦
41 | CWE | 721030 | MDC_IDC_SESS_REPROGRAMMED_PREVIOUS | MDC |  | YES | ^YES |  |  | F |  |  |  |  | OBX¦41¦CWE¦721030^MDC_IDC_SESS_REPROGRAMMED_PREVIOUS^MDC¦¦YES^YES¦¦¦¦¦¦F¦¦¦¦¦¦¦
42 | ST | 721031 | MDC_IDC_SESS_CLINICIAN_NAME | MDC |  | Dr. Anderson |  |  |  | F |  |  |  |  | OBX¦42¦ST¦721031^MDC_IDC_SESS_CLINICIAN_NAME^MDC¦¦Dr. Anderson¦¦¦¦¦¦F¦¦¦¦¦¦¦
43 | ST | 721032 | MDC_IDC_SESS_CLINICIAN_CONTACT_INFORMATION | MDC |  | Phone: +1 12 345 6789 |  |  |  | F |  |  |  |  | OBX¦43¦ST¦721032^MDC_IDC_SESS_CLINICIAN_CONTACT_INFORMATION^MDC¦¦Phone: +1 12 345 6789¦¦¦¦¦¦F¦¦¦¦¦¦¦
44 | ST | 721033 | MDC_IDC_SESS_CLINIC_NAME | MDC |  | Main Heart Center New Jersey |  |  |  | F |  |  |  |  | OBX¦44¦ST¦721033^MDC_IDC_SESS_CLINIC_NAME^MDC¦¦Main Heart Center New Jersey¦¦¦¦¦¦F¦¦¦¦¦¦¦
45 | DTM | 721216 | MDC_IDC_MSMT_BATTERY_DTM | MDC | 1 | 20090825100000 |  |  |  | F |  |  |  |  | OBX¦45¦DTM¦721216^MDC_IDC_MSMT_BATTERY_DTM^MDC¦1¦20090825100000¦¦¦¦¦¦F¦¦¦¦¦¦¦
46 | CWE | 721280 | MDC_IDC_MSMT_BATTERY_STATUS | MDC | 1 | MOS | ^Middle of Service |  |  | F |  |  |  |  | OBX¦46¦CWE¦721280^MDC_IDC_MSMT_BATTERY_STATUS^MDC¦1¦MOS^Middle of Service¦¦¦¦¦¦F¦¦¦¦¦¦¦
47 | NM | 721344 | MDC_IDC_MSMT_BATTERY_VOLTAGE | MDC | 1 | 6.3 |  | V |  | F |  |  |  |  | OBX¦47¦NM¦721344^MDC_IDC_MSMT_BATTERY_VOLTAGE^MDC¦1¦6.3¦V¦¦¦¦¦F¦¦¦¦¦¦¦
48 | NM | 721408 | MDC_IDC_MSMT_BATTERY_IMPEDANCE | MDC | 1 | 2500 |  | Ohm |  | F |  |  |  |  | OBX¦48¦NM¦721408^MDC_IDC_MSMT_BATTERY_IMPEDANCE^MDC¦1¦2500¦Ohm¦¦¦¦¦F¦¦¦¦¦¦¦
49 | NM | 721472 | MDC_IDC_MSMT_BATTERY_REMAINING_LONGEVITY | MDC | 1 | 39 |  | mo |  | F |  |  |  |  | OBX¦49¦NM¦721472^MDC_IDC_MSMT_BATTERY_REMAINING_LONGEVITY^MDC¦1¦39¦mo¦¦¦¦¦F¦¦¦¦¦¦¦
50 | NM | 721536 | MDC_IDC_MSMT_BATTERY_REMAINING_PERCENTAGE | MDC | 1 | 75 |  | % |  | F |  |  |  |  | OBX¦50¦NM¦721536^MDC_IDC_MSMT_BATTERY_REMAINING_PERCENTAGE^MDC¦1¦75¦%¦¦¦¦¦F¦¦¦¦¦¦¦
51 | ST | 721600 | MDC_IDC_MSMT_BATTERY_RRT_TRIGGER | MDC | 1 | Battery voltage < 5.7 V / Cap. Charge time > 12 s |  |  |  | F |  |  |  |  | OBX¦51¦ST¦721600^MDC_IDC_MSMT_BATTERY_RRT_TRIGGER^MDC¦1¦Battery voltage < 5.7 V / Cap. Charge time > 12 s¦¦¦¦¦¦F¦¦¦¦¦¦¦
52 | DTM | 721664 | MDC_IDC_MSMT_CAP_CHARGE_DTM | MDC | 1 | 20100601000000 |  |  |  | F |  |  |  |  | OBX¦52¦DTM¦721664^MDC_IDC_MSMT_CAP_CHARGE_DTM^MDC¦1¦20100601000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
53 | NM | 721728 | MDC_IDC_MSMT_CAP_CHARGE_TIME | MDC | 1 | 8.4 |  | s |  | F |  |  |  |  | OBX¦53¦NM¦721728^MDC_IDC_MSMT_CAP_CHARGE_TIME^MDC¦1¦8.4¦s¦¦¦¦¦F¦¦¦¦¦¦¦
54 | NM | 721792 | MDC_IDC_MSMT_CAP_CHARGE_ENERGY | MDC | 1 | 36 |  | J |  | F |  |  |  |  | OBX¦54¦NM¦721792^MDC_IDC_MSMT_CAP_CHARGE_ENERGY^MDC¦1¦36¦J¦¦¦¦¦F¦¦¦¦¦¦¦
55 | CWE | 721856 | MDC_IDC_MSMT_CAP_CHARGE_TYPE | MDC | 1 | Reformation | ^Reformation |  |  | F |  |  |  |  | OBX¦55¦CWE¦721856^MDC_IDC_MSMT_CAP_CHARGE_TYPE^MDC¦1¦Reformation^Reformation¦¦¦¦¦¦F¦¦¦¦¦¦¦
56 | DTM | 721920 | MDC_IDC_MSMT_LEADCHNL_RA_DTM | MDC |  | 20111025100000 |  |  |  | F |  |  |  |  | OBX¦56¦DTM¦721920^MDC_IDC_MSMT_LEADCHNL_RA_DTM^MDC¦¦20111025100000¦¦¦¦¦¦F¦¦¦¦¦¦¦
57 | DTM | 721921 | MDC_IDC_MSMT_LEADCHNL_RA_DTM_START | MDC |  | 20110925000000 |  |  |  | F |  |  |  |  | OBX¦57¦DTM¦721921^MDC_IDC_MSMT_LEADCHNL_RA_DTM_START^MDC¦¦20110925000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
58 | DTM | 721922 | MDC_IDC_MSMT_LEADCHNL_RA_DTM_END | MDC |  | 20111025100000 |  |  |  | F |  |  |  |  | OBX¦58¦DTM¦721922^MDC_IDC_MSMT_LEADCHNL_RA_DTM_END^MDC¦¦20111025100000¦¦¦¦¦¦F¦¦¦¦¦¦¦
59 | CWE | 721984 | MDC_IDC_MSMT_LEADCHNL_RA_LEAD_CHANNEL_STATUS | MDC |  | NULL | ^Null |  |  | F |  |  |  |  | OBX¦59¦CWE¦721984^MDC_IDC_MSMT_LEADCHNL_RA_LEAD_CHANNEL_STATUS^MDC¦¦NULL^Null¦¦¦¦¦¦F¦¦¦¦¦¦¦
60 | NM | 722048 | MDC_IDC_MSMT_LEADCHNL_RA_SENSING_INTR_AMPL | MDC |  | 2.5 |  | mV |  | F |  |  |  |  | OBX¦60¦NM¦722048^MDC_IDC_MSMT_LEADCHNL_RA_SENSING_INTR_AMPL^MDC¦¦2.5¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
61 | NM | 722049 | MDC_IDC_MSMT_LEADCHNL_RA_SENSING_INTR_AMPL_MAX | MDC |  | 2.6 |  | mV |  | F |  |  |  |  | OBX¦61¦NM¦722049^MDC_IDC_MSMT_LEADCHNL_RA_SENSING_INTR_AMPL_MAX^MDC¦¦2.6¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
62 | NM | 722050 | MDC_IDC_MSMT_LEADCHNL_RA_SENSING_INTR_AMPL_MIN | MDC |  | 2.2 |  | mV |  | F |  |  |  |  | OBX¦62¦NM¦722050^MDC_IDC_MSMT_LEADCHNL_RA_SENSING_INTR_AMPL_MIN^MDC¦¦2.2¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
63 | NM | 722051 | MDC_IDC_MSMT_LEADCHNL_RA_SENSING_INTR_AMPL_MEAN | MDC |  | 2.5 |  | mV |  | F |  |  |  |  | OBX¦63¦NM¦722051^MDC_IDC_MSMT_LEADCHNL_RA_SENSING_INTR_AMPL_MEAN^MDC¦¦2.5¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
64 | CWE | 722112 | MDC_IDC_MSMT_LEADCHNL_RA_SENSING_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦64¦CWE¦722112^MDC_IDC_MSMT_LEADCHNL_RA_SENSING_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
65 | NM | 722176 | MDC_IDC_MSMT_LEADCHNL_RA_PACING_THRESHOLD_AMPLITUDE | MDC |  | 0.7 |  | V |  | F |  |  |  |  | OBX¦65¦NM¦722176^MDC_IDC_MSMT_LEADCHNL_RA_PACING_THRESHOLD_AMPLITUDE^MDC¦¦0.7¦V¦¦¦¦¦F¦¦¦¦¦¦¦
66 | NM | 722240 | MDC_IDC_MSMT_LEADCHNL_RA_PACING_THRESHOLD_PULSEWIDTH | MDC |  | 0.5 |  | ms |  | F |  |  |  |  | OBX¦66¦NM¦722240^MDC_IDC_MSMT_LEADCHNL_RA_PACING_THRESHOLD_PULSEWIDTH^MDC¦¦0.5¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
67 | CWE | 722304 | MDC_IDC_MSMT_LEADCHNL_RA_PACING_THRESHOLD_MEASUREMENT_METHOD | MDC |  | ProgrammerManual | ^Programmer Manual |  |  | F |  |  |  |  | OBX¦67¦CWE¦722304^MDC_IDC_MSMT_LEADCHNL_RA_PACING_THRESHOLD_MEASUREMENT_METHOD^MDC¦¦ProgrammerManual^Programmer Manual¦¦¦¦¦¦F¦¦¦¦¦¦¦
68 | CWE | 722368 | MDC_IDC_MSMT_LEADCHNL_RA_PACING_THRESHOLD_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦68¦CWE¦722368^MDC_IDC_MSMT_LEADCHNL_RA_PACING_THRESHOLD_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
69 | NM | 722432 | MDC_IDC_MSMT_LEADCHNL_RA_IMPEDANCE_VALUE | MDC |  | 500 |  | Ohm |  | F |  |  |  |  | OBX¦69¦NM¦722432^MDC_IDC_MSMT_LEADCHNL_RA_IMPEDANCE_VALUE^MDC¦¦500¦Ohm¦¦¦¦¦F¦¦¦¦¦¦¦
70 | CWE | 722496 | MDC_IDC_MSMT_LEADCHNL_RA_IMPEDANCE_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦70¦CWE¦722496^MDC_IDC_MSMT_LEADCHNL_RA_IMPEDANCE_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
71 | DTM | 721924 | MDC_IDC_MSMT_LEADCHNL_RV_DTM | MDC |  | 20111025100000 |  |  |  | F |  |  |  |  | OBX¦71¦DTM¦721924^MDC_IDC_MSMT_LEADCHNL_RV_DTM^MDC¦¦20111025100000¦¦¦¦¦¦F¦¦¦¦¦¦¦
72 | DTM | 721925 | MDC_IDC_MSMT_LEADCHNL_RV_DTM_START | MDC |  | 20110925000000 |  |  |  | F |  |  |  |  | OBX¦72¦DTM¦721925^MDC_IDC_MSMT_LEADCHNL_RV_DTM_START^MDC¦¦20110925000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
73 | DTM | 721926 | MDC_IDC_MSMT_LEADCHNL_RV_DTM_END | MDC |  | 20111025100000 |  |  |  | F |  |  |  |  | OBX¦73¦DTM¦721926^MDC_IDC_MSMT_LEADCHNL_RV_DTM_END^MDC¦¦20111025100000¦¦¦¦¦¦F¦¦¦¦¦¦¦
74 | CWE | 721985 | MDC_IDC_MSMT_LEADCHNL_RV_LEAD_CHANNEL_STATUS | MDC |  | CheckLead | ^Check Lead |  |  | F |  |  |  |  | OBX¦74¦CWE¦721985^MDC_IDC_MSMT_LEADCHNL_RV_LEAD_CHANNEL_STATUS^MDC¦¦CheckLead^Check Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
75 | NM | 722052 | MDC_IDC_MSMT_LEADCHNL_RV_SENSING_INTR_AMPL | MDC |  | 4.7 |  | mV |  | F |  |  |  |  | OBX¦75¦NM¦722052^MDC_IDC_MSMT_LEADCHNL_RV_SENSING_INTR_AMPL^MDC¦¦4.7¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
76 | NM | 722053 | MDC_IDC_MSMT_LEADCHNL_RV_SENSING_INTR_AMPL_MAX | MDC |  | 4.9 |  | mV |  | F |  |  |  |  | OBX¦76¦NM¦722053^MDC_IDC_MSMT_LEADCHNL_RV_SENSING_INTR_AMPL_MAX^MDC¦¦4.9¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
77 | NM | 722054 | MDC_IDC_MSMT_LEADCHNL_RV_SENSING_INTR_AMPL_MIN | MDC |  | 4.0 |  | mV |  | F |  |  |  |  | OBX¦77¦NM¦722054^MDC_IDC_MSMT_LEADCHNL_RV_SENSING_INTR_AMPL_MIN^MDC¦¦4.0¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
78 | NM | 722055 | MDC_IDC_MSMT_LEADCHNL_RV_SENSING_INTR_AMPL_MEAN | MDC |  | 4.7 |  | mV |  | F |  |  |  |  | OBX¦78¦NM¦722055^MDC_IDC_MSMT_LEADCHNL_RV_SENSING_INTR_AMPL_MEAN^MDC¦¦4.7¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
79 | CWE | 722113 | MDC_IDC_MSMT_LEADCHNL_RV_SENSING_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦79¦CWE¦722113^MDC_IDC_MSMT_LEADCHNL_RV_SENSING_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
80 | NM | 722177 | MDC_IDC_MSMT_LEADCHNL_RV_PACING_THRESHOLD_AMPLITUDE | MDC |  | 0.6 |  | V |  | F |  |  |  |  | OBX¦80¦NM¦722177^MDC_IDC_MSMT_LEADCHNL_RV_PACING_THRESHOLD_AMPLITUDE^MDC¦¦0.6¦V¦¦¦¦¦F¦¦¦¦¦¦¦
81 | NM | 722241 | MDC_IDC_MSMT_LEADCHNL_RV_PACING_THRESHOLD_PULSEWIDTH | MDC |  | 0.5 |  | ms |  | F |  |  |  |  | OBX¦81¦NM¦722241^MDC_IDC_MSMT_LEADCHNL_RV_PACING_THRESHOLD_PULSEWIDTH^MDC¦¦0.5¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
82 | CWE | 722305 | MDC_IDC_MSMT_LEADCHNL_RV_PACING_THRESHOLD_MEASUREMENT_METHOD | MDC |  | DeviceAutomatic | ^Device Automatic |  |  | F |  |  |  |  | OBX¦82¦CWE¦722305^MDC_IDC_MSMT_LEADCHNL_RV_PACING_THRESHOLD_MEASUREMENT_METHOD^MDC¦¦DeviceAutomatic^Device Automatic¦¦¦¦¦¦F¦¦¦¦¦¦¦
83 | CWE | 722369 | MDC_IDC_MSMT_LEADCHNL_RV_PACING_THRESHOLD_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦83¦CWE¦722369^MDC_IDC_MSMT_LEADCHNL_RV_PACING_THRESHOLD_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
84 | NM | 722433 | MDC_IDC_MSMT_LEADCHNL_RV_IMPEDANCE_VALUE | MDC |  | 3000 |  | Ohm |  | F |  |  |  |  | OBX¦84¦NM¦722433^MDC_IDC_MSMT_LEADCHNL_RV_IMPEDANCE_VALUE^MDC¦¦3000¦Ohm¦¦¦¦¦F¦¦¦¦¦¦¦
85 | CWE | 722497 | MDC_IDC_MSMT_LEADCHNL_RV_IMPEDANCE_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦85¦CWE¦722497^MDC_IDC_MSMT_LEADCHNL_RV_IMPEDANCE_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
86 | DTM | 721932 | MDC_IDC_MSMT_LEADCHNL_LV_DTM | MDC |  | 20111025100000 |  |  |  | F |  |  |  |  | OBX¦86¦DTM¦721932^MDC_IDC_MSMT_LEADCHNL_LV_DTM^MDC¦¦20111025100000¦¦¦¦¦¦F¦¦¦¦¦¦¦
87 | DTM | 721933 | MDC_IDC_MSMT_LEADCHNL_LV_DTM_START | MDC |  | 20110925000000 |  |  |  | F |  |  |  |  | OBX¦87¦DTM¦721933^MDC_IDC_MSMT_LEADCHNL_LV_DTM_START^MDC¦¦20110925000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
88 | DTM | 721934 | MDC_IDC_MSMT_LEADCHNL_LV_DTM_END | MDC |  | 20111025100000 |  |  |  | F |  |  |  |  | OBX¦88¦DTM¦721934^MDC_IDC_MSMT_LEADCHNL_LV_DTM_END^MDC¦¦20111025100000¦¦¦¦¦¦F¦¦¦¦¦¦¦
89 | CWE | 721987 | MDC_IDC_MSMT_LEADCHNL_LV_LEAD_CHANNEL_STATUS | MDC |  | NULL | ^Null |  |  | F |  |  |  |  | OBX¦89¦CWE¦721987^MDC_IDC_MSMT_LEADCHNL_LV_LEAD_CHANNEL_STATUS^MDC¦¦NULL^Null¦¦¦¦¦¦F¦¦¦¦¦¦¦
90 | NM | 722060 | MDC_IDC_MSMT_LEADCHNL_LV_SENSING_INTR_AMPL | MDC |  | 3.5 |  | mV |  | F |  |  |  |  | OBX¦90¦NM¦722060^MDC_IDC_MSMT_LEADCHNL_LV_SENSING_INTR_AMPL^MDC¦¦3.5¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
91 | NM | 722061 | MDC_IDC_MSMT_LEADCHNL_LV_SENSING_INTR_AMPL_MAX | MDC |  | 3.8 |  | mV |  | F |  |  |  |  | OBX¦91¦NM¦722061^MDC_IDC_MSMT_LEADCHNL_LV_SENSING_INTR_AMPL_MAX^MDC¦¦3.8¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
92 | NM | 722062 | MDC_IDC_MSMT_LEADCHNL_LV_SENSING_INTR_AMPL_MIN | MDC |  | 2.2 |  | mV |  | F |  |  |  |  | OBX¦92¦NM¦722062^MDC_IDC_MSMT_LEADCHNL_LV_SENSING_INTR_AMPL_MIN^MDC¦¦2.2¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
93 | NM | 722063 | MDC_IDC_MSMT_LEADCHNL_LV_SENSING_INTR_AMPL_MEAN | MDC |  | 3.5 |  | mV |  | F |  |  |  |  | OBX¦93¦NM¦722063^MDC_IDC_MSMT_LEADCHNL_LV_SENSING_INTR_AMPL_MEAN^MDC¦¦3.5¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
94 | CWE | 722115 | MDC_IDC_MSMT_LEADCHNL_LV_SENSING_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦94¦CWE¦722115^MDC_IDC_MSMT_LEADCHNL_LV_SENSING_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
95 | NM | 722179 | MDC_IDC_MSMT_LEADCHNL_LV_PACING_THRESHOLD_AMPLITUDE | MDC |  | 0.6 |  | V |  | F |  |  |  |  | OBX¦95¦NM¦722179^MDC_IDC_MSMT_LEADCHNL_LV_PACING_THRESHOLD_AMPLITUDE^MDC¦¦0.6¦V¦¦¦¦¦F¦¦¦¦¦¦¦
96 | NM | 722243 | MDC_IDC_MSMT_LEADCHNL_LV_PACING_THRESHOLD_PULSEWIDTH | MDC |  | 0.5 |  | ms |  | F |  |  |  |  | OBX¦96¦NM¦722243^MDC_IDC_MSMT_LEADCHNL_LV_PACING_THRESHOLD_PULSEWIDTH^MDC¦¦0.5¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
97 | CWE | 722307 | MDC_IDC_MSMT_LEADCHNL_LV_PACING_THRESHOLD_MEASUREMENT_METHOD | MDC |  | ProgrammerAutomatic | ^Programmer Automatic |  |  | F |  |  |  |  | OBX¦97¦CWE¦722307^MDC_IDC_MSMT_LEADCHNL_LV_PACING_THRESHOLD_MEASUREMENT_METHOD^MDC¦¦ProgrammerAutomatic^Programmer Automatic¦¦¦¦¦¦F¦¦¦¦¦¦¦
98 | CWE | 722371 | MDC_IDC_MSMT_LEADCHNL_LV_PACING_THRESHOLD_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦98¦CWE¦722371^MDC_IDC_MSMT_LEADCHNL_LV_PACING_THRESHOLD_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
99 | NM | 722435 | MDC_IDC_MSMT_LEADCHNL_LV_IMPEDANCE_VALUE | MDC |  | 500 |  | Ohm |  | F |  |  |  |  | OBX¦99¦NM¦722435^MDC_IDC_MSMT_LEADCHNL_LV_IMPEDANCE_VALUE^MDC¦¦500¦Ohm¦¦¦¦¦F¦¦¦¦¦¦¦
100 | CWE | 722499 | MDC_IDC_MSMT_LEADCHNL_LV_IMPEDANCE_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦100¦CWE¦722499^MDC_IDC_MSMT_LEADCHNL_LV_IMPEDANCE_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
101 | DTM | 722560 | MDC_IDC_MSMT_LEADHVCHNL_DTM | MDC | 1 | 20111025100000 |  |  |  | F |  |  |  |  | OBX¦101¦DTM¦722560^MDC_IDC_MSMT_LEADHVCHNL_DTM^MDC¦1¦20111025100000¦¦¦¦¦¦F¦¦¦¦¦¦¦
102 | DTM | 722561 | MDC_IDC_MSMT_LEADHVCHNL_DTM_START | MDC | 1 | 20110925000000 |  |  |  | F |  |  |  |  | OBX¦102¦DTM¦722561^MDC_IDC_MSMT_LEADHVCHNL_DTM_START^MDC¦1¦20110925000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
103 | DTM | 722562 | MDC_IDC_MSMT_LEADHVCHNL_DTM_END | MDC | 1 | 20111025100000 |  |  |  | F |  |  |  |  | OBX¦103¦DTM¦722562^MDC_IDC_MSMT_LEADHVCHNL_DTM_END^MDC¦1¦20111025100000¦¦¦¦¦¦F¦¦¦¦¦¦¦
104 | NM | 722624 | MDC_IDC_MSMT_LEADHVCHNL_IMPEDANCE | MDC | 1 | 330 |  | Ohm |  | F |  |  |  |  | OBX¦104¦NM¦722624^MDC_IDC_MSMT_LEADHVCHNL_IMPEDANCE^MDC¦1¦330¦Ohm¦¦¦¦¦F¦¦¦¦¦¦¦
105 | CWE | 722688 | MDC_IDC_MSMT_LEADHVCHNL_MEASUREMENT_TYPE | MDC | 1 | LowVoltage | ^Low Voltage Pulse |  |  | F |  |  |  |  | OBX¦105¦CWE¦722688^MDC_IDC_MSMT_LEADHVCHNL_MEASUREMENT_TYPE^MDC¦1¦LowVoltage^Low Voltage Pulse¦¦¦¦¦¦F¦¦¦¦¦¦¦
106 | CWE | 722752 | MDC_IDC_MSMT_LEADHVCHNL_STATUS | MDC | 1 | CheckLead | ^Check Lead |  |  | F |  |  |  |  | OBX¦106¦CWE¦722752^MDC_IDC_MSMT_LEADHVCHNL_STATUS^MDC¦1¦CheckLead^Check Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
107 | NM | 729344 | MDC_IDC_SET_CRT_LVRV_DELAY | MDC |  | -20 |  | ms |  | F |  |  |  |  | OBX¦107¦NM¦729344^MDC_IDC_SET_CRT_LVRV_DELAY^MDC¦¦-20¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
108 | CWE | 729408 | MDC_IDC_SET_CRT_PACED_CHAMBERS | MDC |  | BiV | ^BiV |  |  | F |  |  |  |  | OBX¦108¦CWE¦729408^MDC_IDC_SET_CRT_PACED_CHAMBERS^MDC¦¦BiV^BiV¦¦¦¦¦¦F¦¦¦¦¦¦¦
109 | ST | 729472 | MDC_IDC_SET_MAGNET_RESP | MDC |  | Therapies inactive, VVIR |  |  |  | F |  |  |  |  | OBX¦109¦ST¦729472^MDC_IDC_SET_MAGNET_RESP^MDC¦¦Therapies inactive, VVIR¦¦¦¦¦¦F¦¦¦¦¦¦¦
110 | NM | 729536 | MDC_IDC_SET_LEADCHNL_RA_SENSING_SENSITIVITY | MDC |  | 0.8 |  | mV |  | F |  |  |  |  | OBX¦110¦NM¦729536^MDC_IDC_SET_LEADCHNL_RA_SENSING_SENSITIVITY^MDC¦¦0.8¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
111 | CWE | 729600 | MDC_IDC_SET_LEADCHNL_RA_SENSING_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦111¦CWE¦729600^MDC_IDC_SET_LEADCHNL_RA_SENSING_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
112 | CWE | 729665 | MDC_IDC_SET_LEADCHNL_RA_SENSING_ANODE_LOCATION_1 | MDC |  | RA | ^Right Atrium |  |  | F |  |  |  |  | OBX¦112¦CWE¦729665^MDC_IDC_SET_LEADCHNL_RA_SENSING_ANODE_LOCATION_1^MDC¦¦RA^Right Atrium¦¦¦¦¦¦F¦¦¦¦¦¦¦
113 | CWE | 729729 | MDC_IDC_SET_LEADCHNL_RA_SENSING_ANODE_ELECTRODE_1 | MDC |  | Tip | ^Tip |  |  | F |  |  |  |  | OBX¦113¦CWE¦729729^MDC_IDC_SET_LEADCHNL_RA_SENSING_ANODE_ELECTRODE_1^MDC¦¦Tip^Tip¦¦¦¦¦¦F¦¦¦¦¦¦¦
114 | CWE | 729793 | MDC_IDC_SET_LEADCHNL_RA_SENSING_CATHODE_LOCATION_1 | MDC |  | Other | ^Other |  |  | F |  |  |  |  | OBX¦114¦CWE¦729793^MDC_IDC_SET_LEADCHNL_RA_SENSING_CATHODE_LOCATION_1^MDC¦¦Other^Other¦¦¦¦¦¦F¦¦¦¦¦¦¦
115 | CWE | 729857 | MDC_IDC_SET_LEADCHNL_RA_SENSING_CATHODE_ELECTRODE_1 | MDC |  | Can | ^Can |  |  | F |  |  |  |  | OBX¦115¦CWE¦729857^MDC_IDC_SET_LEADCHNL_RA_SENSING_CATHODE_ELECTRODE_1^MDC¦¦Can^Can¦¦¦¦¦¦F¦¦¦¦¦¦¦
116 | CWE | 729920 | MDC_IDC_SET_LEADCHNL_RA_SENSING_ADAPTATION_MODE | MDC |  | AdaptiveSensing | ^Adaptive |  |  | F |  |  |  |  | OBX¦116¦CWE¦729920^MDC_IDC_SET_LEADCHNL_RA_SENSING_ADAPTATION_MODE^MDC¦¦AdaptiveSensing^Adaptive¦¦¦¦¦¦F¦¦¦¦¦¦¦
117 | NM | 729984 | MDC_IDC_SET_LEADCHNL_RA_PACING_AMPLITUDE | MDC |  | 1.8 |  | V |  | F |  |  |  |  | OBX¦117¦NM¦729984^MDC_IDC_SET_LEADCHNL_RA_PACING_AMPLITUDE^MDC¦¦1.8¦V¦¦¦¦¦F¦¦¦¦¦¦¦
118 | NM | 730048 | MDC_IDC_SET_LEADCHNL_RA_PACING_PULSEWIDTH | MDC |  | 0.5 |  | ms |  | F |  |  |  |  | OBX¦118¦NM¦730048^MDC_IDC_SET_LEADCHNL_RA_PACING_PULSEWIDTH^MDC¦¦0.5¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
119 | CWE | 730112 | MDC_IDC_SET_LEADCHNL_RA_PACING_POLARITY | MDC |  | UNI | ^Unipolar Lead |  |  | F |  |  |  |  | OBX¦119¦CWE¦730112^MDC_IDC_SET_LEADCHNL_RA_PACING_POLARITY^MDC¦¦UNI^Unipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
120 | CWE | 730177 | MDC_IDC_SET_LEADCHNL_RA_PACING_ANODE_LOCATION_1 | MDC |  | RA | ^Right Atrium |  |  | F |  |  |  |  | OBX¦120¦CWE¦730177^MDC_IDC_SET_LEADCHNL_RA_PACING_ANODE_LOCATION_1^MDC¦¦RA^Right Atrium¦¦¦¦¦¦F¦¦¦¦¦¦¦
121 | CWE | 730241 | MDC_IDC_SET_LEADCHNL_RA_PACING_ANODE_ELECTRODE_1 | MDC |  | Ring | ^Ring |  |  | F |  |  |  |  | OBX¦121¦CWE¦730241^MDC_IDC_SET_LEADCHNL_RA_PACING_ANODE_ELECTRODE_1^MDC¦¦Ring^Ring¦¦¦¦¦¦F¦¦¦¦¦¦¦
122 | CWE | 730305 | MDC_IDC_SET_LEADCHNL_RA_PACING_CATHODE_LOCATION_1 | MDC |  | Other | ^Other |  |  | F |  |  |  |  | OBX¦122¦CWE¦730305^MDC_IDC_SET_LEADCHNL_RA_PACING_CATHODE_LOCATION_1^MDC¦¦Other^Other¦¦¦¦¦¦F¦¦¦¦¦¦¦
123 | CWE | 730369 | MDC_IDC_SET_LEADCHNL_RA_PACING_CATHODE_ELECTRODE_1 | MDC |  | Can | ^Can |  |  | F |  |  |  |  | OBX¦123¦CWE¦730369^MDC_IDC_SET_LEADCHNL_RA_PACING_CATHODE_ELECTRODE_1^MDC¦¦Can^Can¦¦¦¦¦¦F¦¦¦¦¦¦¦
124 | CWE | 730432 | MDC_IDC_SET_LEADCHNL_RA_PACING_CAPTURE_MODE | MDC |  | FixedPacing | ^Fixed Pacing |  |  | F |  |  |  |  | OBX¦124¦CWE¦730432^MDC_IDC_SET_LEADCHNL_RA_PACING_CAPTURE_MODE^MDC¦¦FixedPacing^Fixed Pacing¦¦¦¦¦¦F¦¦¦¦¦¦¦
125 | NM | 729537 | MDC_IDC_SET_LEADCHNL_RV_SENSING_SENSITIVITY | MDC |  | 1.3 |  | mV |  | F |  |  |  |  | OBX¦125¦NM¦729537^MDC_IDC_SET_LEADCHNL_RV_SENSING_SENSITIVITY^MDC¦¦1.3¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
126 | CWE | 729601 | MDC_IDC_SET_LEADCHNL_RV_SENSING_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦126¦CWE¦729601^MDC_IDC_SET_LEADCHNL_RV_SENSING_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
127 | CWE | 729669 | MDC_IDC_SET_LEADCHNL_RV_SENSING_ANODE_LOCATION_1 | MDC |  | RV | ^Right Ventricle |  |  | F |  |  |  |  | OBX¦127¦CWE¦729669^MDC_IDC_SET_LEADCHNL_RV_SENSING_ANODE_LOCATION_1^MDC¦¦RV^Right Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
128 | CWE | 729733 | MDC_IDC_SET_LEADCHNL_RV_SENSING_ANODE_ELECTRODE_1 | MDC |  | Tip | ^Tip |  |  | F |  |  |  |  | OBX¦128¦CWE¦729733^MDC_IDC_SET_LEADCHNL_RV_SENSING_ANODE_ELECTRODE_1^MDC¦¦Tip^Tip¦¦¦¦¦¦F¦¦¦¦¦¦¦
129 | CWE | 729797 | MDC_IDC_SET_LEADCHNL_RV_SENSING_CATHODE_LOCATION_1 | MDC |  | RV | ^Right Ventricle |  |  | F |  |  |  |  | OBX¦129¦CWE¦729797^MDC_IDC_SET_LEADCHNL_RV_SENSING_CATHODE_LOCATION_1^MDC¦¦RV^Right Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
130 | CWE | 729861 | MDC_IDC_SET_LEADCHNL_RV_SENSING_CATHODE_ELECTRODE_1 | MDC |  | Ring | ^Ring |  |  | F |  |  |  |  | OBX¦130¦CWE¦729861^MDC_IDC_SET_LEADCHNL_RV_SENSING_CATHODE_ELECTRODE_1^MDC¦¦Ring^Ring¦¦¦¦¦¦F¦¦¦¦¦¦¦
131 | CWE | 729921 | MDC_IDC_SET_LEADCHNL_RV_SENSING_ADAPTATION_MODE | MDC |  | AdaptiveSensing | ^Adaptive |  |  | F |  |  |  |  | OBX¦131¦CWE¦729921^MDC_IDC_SET_LEADCHNL_RV_SENSING_ADAPTATION_MODE^MDC¦¦AdaptiveSensing^Adaptive¦¦¦¦¦¦F¦¦¦¦¦¦¦
132 | NM | 729985 | MDC_IDC_SET_LEADCHNL_RV_PACING_AMPLITUDE | MDC |  | 2.0 |  | V |  | F |  |  |  |  | OBX¦132¦NM¦729985^MDC_IDC_SET_LEADCHNL_RV_PACING_AMPLITUDE^MDC¦¦2.0¦V¦¦¦¦¦F¦¦¦¦¦¦¦
133 | NM | 730049 | MDC_IDC_SET_LEADCHNL_RV_PACING_PULSEWIDTH | MDC |  | 0.5 |  | ms |  | F |  |  |  |  | OBX¦133¦NM¦730049^MDC_IDC_SET_LEADCHNL_RV_PACING_PULSEWIDTH^MDC¦¦0.5¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
134 | CWE | 730113 | MDC_IDC_SET_LEADCHNL_RV_PACING_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦134¦CWE¦730113^MDC_IDC_SET_LEADCHNL_RV_PACING_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
135 | CWE | 730181 | MDC_IDC_SET_LEADCHNL_RV_PACING_ANODE_LOCATION_1 | MDC |  | RV | ^Right Ventricle |  |  | F |  |  |  |  | OBX¦135¦CWE¦730181^MDC_IDC_SET_LEADCHNL_RV_PACING_ANODE_LOCATION_1^MDC¦¦RV^Right Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
136 | CWE | 730245 | MDC_IDC_SET_LEADCHNL_RV_PACING_ANODE_ELECTRODE_1 | MDC |  | Tip | ^Tip |  |  | F |  |  |  |  | OBX¦136¦CWE¦730245^MDC_IDC_SET_LEADCHNL_RV_PACING_ANODE_ELECTRODE_1^MDC¦¦Tip^Tip¦¦¦¦¦¦F¦¦¦¦¦¦¦
137 | CWE | 730309 | MDC_IDC_SET_LEADCHNL_RV_PACING_CATHODE_LOCATION_1 | MDC |  | RV | ^Right Ventricle |  |  | F |  |  |  |  | OBX¦137¦CWE¦730309^MDC_IDC_SET_LEADCHNL_RV_PACING_CATHODE_LOCATION_1^MDC¦¦RV^Right Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
138 | CWE | 730373 | MDC_IDC_SET_LEADCHNL_RV_PACING_CATHODE_ELECTRODE_1 | MDC |  | Ring | ^Ring |  |  | F |  |  |  |  | OBX¦138¦CWE¦730373^MDC_IDC_SET_LEADCHNL_RV_PACING_CATHODE_ELECTRODE_1^MDC¦¦Ring^Ring¦¦¦¦¦¦F¦¦¦¦¦¦¦
139 | CWE | 730433 | MDC_IDC_SET_LEADCHNL_RV_PACING_CAPTURE_MODE | MDC |  | AdaptiveCapture | ^Adaptive |  |  | F |  |  |  |  | OBX¦139¦CWE¦730433^MDC_IDC_SET_LEADCHNL_RV_PACING_CAPTURE_MODE^MDC¦¦AdaptiveCapture^Adaptive¦¦¦¦¦¦F¦¦¦¦¦¦¦
140 | NM | 729539 | MDC_IDC_SET_LEADCHNL_LV_SENSING_SENSITIVITY | MDC |  | 1.0 |  | mV |  | F |  |  |  |  | OBX¦140¦NM¦729539^MDC_IDC_SET_LEADCHNL_LV_SENSING_SENSITIVITY^MDC¦¦1.0¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
141 | CWE | 729603 | MDC_IDC_SET_LEADCHNL_LV_SENSING_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦141¦CWE¦729603^MDC_IDC_SET_LEADCHNL_LV_SENSING_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
142 | CWE | 729677 | MDC_IDC_SET_LEADCHNL_LV_SENSING_ANODE_LOCATION_1 | MDC |  | LV | ^Left Ventricle |  |  | F |  |  |  |  | OBX¦142¦CWE¦729677^MDC_IDC_SET_LEADCHNL_LV_SENSING_ANODE_LOCATION_1^MDC¦¦LV^Left Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
143 | CWE | 729741 | MDC_IDC_SET_LEADCHNL_LV_SENSING_ANODE_ELECTRODE_1 | MDC |  | Tip | ^Tip |  |  | F |  |  |  |  | OBX¦143¦CWE¦729741^MDC_IDC_SET_LEADCHNL_LV_SENSING_ANODE_ELECTRODE_1^MDC¦¦Tip^Tip¦¦¦¦¦¦F¦¦¦¦¦¦¦
144 | CWE | 729805 | MDC_IDC_SET_LEADCHNL_LV_SENSING_CATHODE_LOCATION_1 | MDC |  | LV | ^Left Ventricle |  |  | F |  |  |  |  | OBX¦144¦CWE¦729805^MDC_IDC_SET_LEADCHNL_LV_SENSING_CATHODE_LOCATION_1^MDC¦¦LV^Left Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
145 | CWE | 729869 | MDC_IDC_SET_LEADCHNL_LV_SENSING_CATHODE_ELECTRODE_1 | MDC |  | Ring | ^Ring |  |  | F |  |  |  |  | OBX¦145¦CWE¦729869^MDC_IDC_SET_LEADCHNL_LV_SENSING_CATHODE_ELECTRODE_1^MDC¦¦Ring^Ring¦¦¦¦¦¦F¦¦¦¦¦¦¦
146 | CWE | 729923 | MDC_IDC_SET_LEADCHNL_LV_SENSING_ADAPTATION_MODE | MDC |  | FixedSensing | ^Fixed |  |  | F |  |  |  |  | OBX¦146¦CWE¦729923^MDC_IDC_SET_LEADCHNL_LV_SENSING_ADAPTATION_MODE^MDC¦¦FixedSensing^Fixed ¦¦¦¦¦¦F¦¦¦¦¦¦¦
147 | NM | 729987 | MDC_IDC_SET_LEADCHNL_LV_PACING_AMPLITUDE | MDC |  | 2.0 |  | V |  | F |  |  |  |  | OBX¦147¦NM¦729987^MDC_IDC_SET_LEADCHNL_LV_PACING_AMPLITUDE^MDC¦¦2.0¦V¦¦¦¦¦F¦¦¦¦¦¦¦
148 | NM | 730051 | MDC_IDC_SET_LEADCHNL_LV_PACING_PULSEWIDTH | MDC |  | 0.5 |  | ms |  | F |  |  |  |  | OBX¦148¦NM¦730051^MDC_IDC_SET_LEADCHNL_LV_PACING_PULSEWIDTH^MDC¦¦0.5¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
149 | CWE | 730115 | MDC_IDC_SET_LEADCHNL_LV_PACING_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦149¦CWE¦730115^MDC_IDC_SET_LEADCHNL_LV_PACING_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
150 | CWE | 730189 | MDC_IDC_SET_LEADCHNL_LV_PACING_ANODE_LOCATION_1 | MDC |  | LV | ^Left Ventricle |  |  | F |  |  |  |  | OBX¦150¦CWE¦730189^MDC_IDC_SET_LEADCHNL_LV_PACING_ANODE_LOCATION_1^MDC¦¦LV^Left Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
151 | CWE | 730253 | MDC_IDC_SET_LEADCHNL_LV_PACING_ANODE_ELECTRODE_1 | MDC |  | Tip | ^Tip |  |  | F |  |  |  |  | OBX¦151¦CWE¦730253^MDC_IDC_SET_LEADCHNL_LV_PACING_ANODE_ELECTRODE_1^MDC¦¦Tip^Tip¦¦¦¦¦¦F¦¦¦¦¦¦¦
152 | CWE | 730317 | MDC_IDC_SET_LEADCHNL_LV_PACING_CATHODE_LOCATION_1 | MDC |  | RV | ^Right Ventricle |  |  | F |  |  |  |  | OBX¦152¦CWE¦730317^MDC_IDC_SET_LEADCHNL_LV_PACING_CATHODE_LOCATION_1^MDC¦¦RV^Right Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
153 | CWE | 730381 | MDC_IDC_SET_LEADCHNL_LV_PACING_CATHODE_ELECTRODE_1 | MDC |  | Ring | ^Ring |  |  | F |  |  |  |  | OBX¦153¦CWE¦730381^MDC_IDC_SET_LEADCHNL_LV_PACING_CATHODE_ELECTRODE_1^MDC¦¦Ring^Ring¦¦¦¦¦¦F¦¦¦¦¦¦¦
154 | CWE | 730435 | MDC_IDC_SET_LEADCHNL_LV_PACING_CAPTURE_MODE | MDC |  | AdaptiveCapture | ^Adaptive |  |  | F |  |  |  |  | OBX¦154¦CWE¦730435^MDC_IDC_SET_LEADCHNL_LV_PACING_CAPTURE_MODE^MDC¦¦AdaptiveCapture^Adaptive¦¦¦¦¦¦F¦¦¦¦¦¦¦
155 | CWE | 730497 | MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_LOCATION_1 | MDC |  | RV | ^Right Ventricle |  |  | F |  |  |  |  | OBX¦155¦CWE¦730497^MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_LOCATION_1^MDC¦¦RV^Right Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
156 | CWE | 730561 | MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_ELECTRODE_1 | MDC |  | Coil | ^Coil |  |  | F |  |  |  |  | OBX¦156¦CWE¦730561^MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_ELECTRODE_1^MDC¦¦Coil^Coil¦¦¦¦¦¦F¦¦¦¦¦¦¦
157 | CWE | 730498 | MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_LOCATION_2 | MDC |  | RA | ^Right Atrium |  |  | F |  |  |  |  | OBX¦157¦CWE¦730498^MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_LOCATION_2^MDC¦¦RA^Right Atrium¦¦¦¦¦¦F¦¦¦¦¦¦¦
158 | CWE | 730562 | MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_ELECTRODE_2 | MDC |  | Coil | ^Coil |  |  | F |  |  |  |  | OBX¦158¦CWE¦730562^MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_ELECTRODE_2^MDC¦¦Coil^Coil¦¦¦¦¦¦F¦¦¦¦¦¦¦
159 | CWE | 730625 | MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_CATHODE_LOCATION_1 | MDC |  | Other | ^Other |  |  | F |  |  |  |  | OBX¦159¦CWE¦730625^MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_CATHODE_LOCATION_1^MDC¦¦Other^Other¦¦¦¦¦¦F¦¦¦¦¦¦¦
160 | CWE | 730689 | MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_CATHODE_ELECTRODE_1 | MDC |  | Can | ^Can |  |  | F |  |  |  |  | OBX¦160¦CWE¦730689^MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_CATHODE_ELECTRODE_1^MDC¦¦Can^Can¦¦¦¦¦¦F¦¦¦¦¦¦¦
161 | CWE | 730752 | MDC_IDC_SET_BRADY_MODE | MDC |  | DDDR | ^DDDR |  |  | F |  |  |  |  | OBX¦161¦CWE¦730752^MDC_IDC_SET_BRADY_MODE^MDC¦¦DDDR^DDDR¦¦¦¦¦¦F¦¦¦¦¦¦¦
162 | CWE | 730816 | MDC_IDC_SET_BRADY_VENDOR_MODE | MDC |  | DDDR | ^DDDR |  |  | F |  |  |  |  | OBX¦162¦CWE¦730816^MDC_IDC_SET_BRADY_VENDOR_MODE^MDC¦¦DDDR^DDDR¦¦¦¦¦¦F¦¦¦¦¦¦¦
163 | NM | 730880 | MDC_IDC_SET_BRADY_LOWRATE | MDC |  | 60 |  | {beats}/min |  | F |  |  |  |  | OBX¦163¦NM¦730880^MDC_IDC_SET_BRADY_LOWRATE^MDC¦¦60¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
164 | NM | 730944 | MDC_IDC_SET_BRADY_HYSTRATE | MDC |  | 55 |  | {beats}/min |  | F |  |  |  |  | OBX¦164¦NM¦730944^MDC_IDC_SET_BRADY_HYSTRATE^MDC¦¦55¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
165 | NM | 731008 | MDC_IDC_SET_BRADY_NIGHT_RATE | MDC |  | 50 |  | {beats}/min |  | F |  |  |  |  | OBX¦165¦NM¦731008^MDC_IDC_SET_BRADY_NIGHT_RATE^MDC¦¦50¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
166 | ST | 731072 | MDC_IDC_SET_BRADY_SENSOR_TYPE | MDC |  | Accelerometer |  |  |  | F |  |  |  |  | OBX¦166¦ST¦731072^MDC_IDC_SET_BRADY_SENSOR_TYPE^MDC¦¦Accelerometer¦¦¦¦¦¦F¦¦¦¦¦¦¦
167 | NM | 731136 | MDC_IDC_SET_BRADY_MAX_TRACKING_RATE | MDC |  | 130 |  | {beats}/min |  | F |  |  |  |  | OBX¦167¦NM¦731136^MDC_IDC_SET_BRADY_MAX_TRACKING_RATE^MDC¦¦130¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
168 | NM | 731200 | MDC_IDC_SET_BRADY_MAX_SENSOR_RATE | MDC |  | 120 |  | {beats}/min |  | F |  |  |  |  | OBX¦168¦NM¦731200^MDC_IDC_SET_BRADY_MAX_SENSOR_RATE^MDC¦¦120¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
169 | NM | 731266 | MDC_IDC_SET_BRADY_SAV_DELAY_LOW | MDC |  | 140 |  | ms |  | F |  |  |  |  | OBX¦169¦NM¦731266^MDC_IDC_SET_BRADY_SAV_DELAY_LOW^MDC¦¦140¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
170 | NM | 731265 | MDC_IDC_SET_BRADY_SAV_DELAY_HIGH | MDC |  | 180 |  | ms |  | F |  |  |  |  | OBX¦170¦NM¦731265^MDC_IDC_SET_BRADY_SAV_DELAY_HIGH^MDC¦¦180¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
171 | NM | 731330 | MDC_IDC_SET_BRADY_PAV_DELAY_LOW | MDC |  | 110 |  | ms |  | F |  |  |  |  | OBX¦171¦NM¦731330^MDC_IDC_SET_BRADY_PAV_DELAY_LOW^MDC¦¦110¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
172 | NM | 731329 | MDC_IDC_SET_BRADY_PAV_DELAY_HIGH | MDC |  | 150 |  | ms |  | F |  |  |  |  | OBX¦172¦NM¦731329^MDC_IDC_SET_BRADY_PAV_DELAY_HIGH^MDC¦¦150¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
173 | CWE | 731392 | MDC_IDC_SET_BRADY_AT_MODE_SWITCH_MODE | MDC |  | DDIR | ^DDIR |  |  | F |  |  |  |  | OBX¦173¦CWE¦731392^MDC_IDC_SET_BRADY_AT_MODE_SWITCH_MODE^MDC¦¦DDIR^DDIR¦¦¦¦¦¦F¦¦¦¦¦¦¦
174 | NM | 731456 | MDC_IDC_SET_BRADY_AT_MODE_SWITCH_RATE | MDC |  | 180 |  | {beats}/min |  | F |  |  |  |  | OBX¦174¦NM¦731456^MDC_IDC_SET_BRADY_AT_MODE_SWITCH_RATE^MDC¦¦180¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
175 | CWE | 731520 | MDC_IDC_SET_TACHYTHERAPY_VSTAT | MDC |  | On | ^On |  |  | F |  |  |  |  | OBX¦175¦CWE¦731520^MDC_IDC_SET_TACHYTHERAPY_VSTAT^MDC¦¦On^On¦¦¦¦¦¦F¦¦¦¦¦¦¦
176 | CWE | 731584 | MDC_IDC_SET_TACHYTHERAPY_ASTAT | MDC |  | Off | ^Off |  |  | F |  |  |  |  | OBX¦176¦CWE¦731584^MDC_IDC_SET_TACHYTHERAPY_ASTAT^MDC¦¦Off^Off¦¦¦¦¦¦F¦¦¦¦¦¦¦
177 | CWE | 731648 | MDC_IDC_SET_ZONE_TYPE | MDC | 1 | Zone_VF | ^VF |  |  | F |  |  |  |  | OBX¦177¦CWE¦731648^MDC_IDC_SET_ZONE_TYPE^MDC¦1¦Zone_VF^VF¦¦¦¦¦¦F¦¦¦¦¦¦¦
178 | CWE | 731712 | MDC_IDC_SET_ZONE_VENDOR_TYPE | MDC | 1 | VF | ^VF |  |  | F |  |  |  |  | OBX¦178¦CWE¦731712^MDC_IDC_SET_ZONE_VENDOR_TYPE^MDC¦1¦VF^VF¦¦¦¦¦¦F¦¦¦¦¦¦¦
179 | CWE | 731776 | MDC_IDC_SET_ZONE_STATUS | MDC | 1 | Active | ^Active |  |  | F |  |  |  |  | OBX¦179¦CWE¦731776^MDC_IDC_SET_ZONE_STATUS^MDC¦1¦Active^Active¦¦¦¦¦¦F¦¦¦¦¦¦¦
180 | NM | 731840 | MDC_IDC_SET_ZONE_DETECTION_INTERVAL | MDC | 1 | 195 |  | ms |  | F |  |  |  |  | OBX¦180¦NM¦731840^MDC_IDC_SET_ZONE_DETECTION_INTERVAL^MDC¦1¦195¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
181 | NM | 731904 | MDC_IDC_SET_ZONE_DETECTION_BEATS_NUMERATOR | MDC | 1 | 12 |  | {beats} |  | F |  |  |  |  | OBX¦181¦NM¦731904^MDC_IDC_SET_ZONE_DETECTION_BEATS_NUMERATOR^MDC¦1¦12¦{beats}¦¦¦¦¦F¦¦¦¦¦¦¦
182 | NM | 731968 | MDC_IDC_SET_ZONE_DETECTION_BEATS_DENOMINATOR | MDC | 1 | 18 |  | {beats} |  | F |  |  |  |  | OBX¦182¦NM¦731968^MDC_IDC_SET_ZONE_DETECTION_BEATS_DENOMINATOR^MDC¦1¦18¦{beats}¦¦¦¦¦F¦¦¦¦¦¦¦
183 | ST | 732032 | MDC_IDC_SET_ZONE_DETECTION_DETAILS | MDC | 1 |  |  |  |  | F |  |  |  |  | OBX¦183¦ST¦732032^MDC_IDC_SET_ZONE_DETECTION_DETAILS^MDC¦1¦¦¦¦¦¦¦F¦¦¦¦¦¦¦
184 | CWE | 732097 | MDC_IDC_SET_ZONE_TYPE_ATP_1 | MDC | 1 | Ramp | ^Ramp |  |  | F |  |  |  |  | OBX¦184¦CWE¦732097^MDC_IDC_SET_ZONE_TYPE_ATP_1^MDC¦1¦Ramp^Ramp¦¦¦¦¦¦F¦¦¦¦¦¦¦
185 | NM | 732161 | MDC_IDC_SET_ZONE_NUM_ATP_SEQS_1 | MDC | 1 | 1 |  | {seq} |  | F |  |  |  |  | OBX¦185¦NM¦732161^MDC_IDC_SET_ZONE_NUM_ATP_SEQS_1^MDC¦1¦1¦{seq}¦¦¦¦¦F¦¦¦¦¦¦¦
186 | NM | 732225 | MDC_IDC_SET_ZONE_SHOCK_ENERGY_1 | MDC | 1 | 30 |  | J |  | F |  |  |  |  | OBX¦186¦NM¦732225^MDC_IDC_SET_ZONE_SHOCK_ENERGY_1^MDC¦1¦30¦J¦¦¦¦¦F¦¦¦¦¦¦¦
187 | NM | 732289 | MDC_IDC_SET_ZONE_NUM_SHOCKS_1 | MDC | 1 | 5 |  | {shocks} |  | F |  |  |  |  | OBX¦187¦NM¦732289^MDC_IDC_SET_ZONE_NUM_SHOCKS_1^MDC¦1¦5¦{shocks}¦¦¦¦¦F¦¦¦¦¦¦¦
188 | CWE | 731648 | MDC_IDC_SET_ZONE_TYPE | MDC | 2 | Zone_VT | ^VT |  |  | F |  |  |  |  | OBX¦188¦CWE¦731648^MDC_IDC_SET_ZONE_TYPE^MDC¦2¦Zone_VT^VT¦¦¦¦¦¦F¦¦¦¦¦¦¦
189 | CWE | 731712 | MDC_IDC_SET_ZONE_VENDOR_TYPE | MDC | 2 | VT1 | ^VT1 |  |  | F |  |  |  |  | OBX¦189¦CWE¦731712^MDC_IDC_SET_ZONE_VENDOR_TYPE^MDC¦2¦VT1^VT1¦¦¦¦¦¦F¦¦¦¦¦¦¦
190 | CWE | 731776 | MDC_IDC_SET_ZONE_STATUS | MDC | 2 | Active | ^Active |  |  | F |  |  |  |  | OBX¦190¦CWE¦731776^MDC_IDC_SET_ZONE_STATUS^MDC¦2¦Active^Active¦¦¦¦¦¦F¦¦¦¦¦¦¦
191 | NM | 731840 | MDC_IDC_SET_ZONE_DETECTION_INTERVAL | MDC | 2 | 165 |  | ms |  | F |  |  |  |  | OBX¦191¦NM¦731840^MDC_IDC_SET_ZONE_DETECTION_INTERVAL^MDC¦2¦165¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
192 | NM | 731904 | MDC_IDC_SET_ZONE_DETECTION_BEATS_NUMERATOR | MDC | 2 | 9 |  | {beats} |  | F |  |  |  |  | OBX¦192¦NM¦731904^MDC_IDC_SET_ZONE_DETECTION_BEATS_NUMERATOR^MDC¦2¦9¦{beats}¦¦¦¦¦F¦¦¦¦¦¦¦
193 | NM | 731968 | MDC_IDC_SET_ZONE_DETECTION_BEATS_DENOMINATOR | MDC | 2 | 12 |  | {beats} |  | F |  |  |  |  | OBX¦193¦NM¦731968^MDC_IDC_SET_ZONE_DETECTION_BEATS_DENOMINATOR^MDC¦2¦12¦{beats}¦¦¦¦¦F¦¦¦¦¦¦¦
194 | ST | 732032 | MDC_IDC_SET_ZONE_DETECTION_DETAILS | MDC | 2 | Enhanced detection |  |  |  | F |  |  |  |  | OBX¦194¦ST¦732032^MDC_IDC_SET_ZONE_DETECTION_DETAILS^MDC¦2¦Enhanced detection¦¦¦¦¦¦F¦¦¦¦¦¦¦
195 | CWE | 732097 | MDC_IDC_SET_ZONE_TYPE_ATP_1 | MDC | 2 | Burst | ^Burst |  |  | F |  |  |  |  | OBX¦195¦CWE¦732097^MDC_IDC_SET_ZONE_TYPE_ATP_1^MDC¦2¦Burst^Burst¦¦¦¦¦¦F¦¦¦¦¦¦¦
196 | NM | 732161 | MDC_IDC_SET_ZONE_NUM_ATP_SEQS_1 | MDC | 2 | 5 |  | {seq} |  | F |  |  |  |  | OBX¦196¦NM¦732161^MDC_IDC_SET_ZONE_NUM_ATP_SEQS_1^MDC¦2¦5¦{seq}¦¦¦¦¦F¦¦¦¦¦¦¦
197 | NM | 732225 | MDC_IDC_SET_ZONE_SHOCK_ENERGY_1 | MDC | 2 | 20 |  | J |  | F |  |  |  |  | OBX¦197¦NM¦732225^MDC_IDC_SET_ZONE_SHOCK_ENERGY_1^MDC¦2¦20¦J¦¦¦¦¦F¦¦¦¦¦¦¦
198 | NM | 732289 | MDC_IDC_SET_ZONE_NUM_SHOCKS_1 | MDC | 2 | 1 |  | {shocks} |  | F |  |  |  |  | OBX¦198¦NM¦732289^MDC_IDC_SET_ZONE_NUM_SHOCKS_1^MDC¦2¦1¦{shocks}¦¦¦¦¦F¦¦¦¦¦¦¦
199 | NM | 732226 | MDC_IDC_SET_ZONE_SHOCK_ENERGY_2 | MDC | 2 | 30 |  | J |  | F |  |  |  |  | OBX¦199¦NM¦732226^MDC_IDC_SET_ZONE_SHOCK_ENERGY_2^MDC¦2¦30¦J¦¦¦¦¦F¦¦¦¦¦¦¦
200 | NM | 732290 | MDC_IDC_SET_ZONE_NUM_SHOCKS_2 | MDC | 2 | 1 |  | {shocks} |  | F |  |  |  |  | OBX¦200¦NM¦732290^MDC_IDC_SET_ZONE_NUM_SHOCKS_2^MDC¦2¦1¦{shocks}¦¦¦¦¦F¦¦¦¦¦¦¦
201 | NM | 732227 | MDC_IDC_SET_ZONE_SHOCK_ENERGY_3 | MDC | 2 | 30 |  | J |  | F |  |  |  |  | OBX¦201¦NM¦732227^MDC_IDC_SET_ZONE_SHOCK_ENERGY_3^MDC¦2¦30¦J¦¦¦¦¦F¦¦¦¦¦¦¦
202 | NM | 732291 | MDC_IDC_SET_ZONE_NUM_SHOCKS_3 | MDC | 2 | 5 |  | {shocks} |  | F |  |  |  |  | OBX¦202¦NM¦732291^MDC_IDC_SET_ZONE_NUM_SHOCKS_3^MDC¦2¦5¦{shocks}¦¦¦¦¦F¦¦¦¦¦¦¦
203 | CWE | 731648 | MDC_IDC_SET_ZONE_TYPE | MDC | 3 | Zone_VT | ^VT |  |  | F |  |  |  |  | OBX¦203¦CWE¦731648^MDC_IDC_SET_ZONE_TYPE^MDC¦3¦Zone_VT^VT¦¦¦¦¦¦F¦¦¦¦¦¦¦
204 | CWE | 731712 | MDC_IDC_SET_ZONE_VENDOR_TYPE | MDC | 3 | FastVT | ^FastVT |  |  | F |  |  |  |  | OBX¦204¦CWE¦731712^MDC_IDC_SET_ZONE_VENDOR_TYPE^MDC¦3¦FastVT^FastVT¦¦¦¦¦¦F¦¦¦¦¦¦¦
205 | CWE | 731776 | MDC_IDC_SET_ZONE_STATUS | MDC | 3 | Active | ^Active |  |  | F |  |  |  |  | OBX¦205¦CWE¦731776^MDC_IDC_SET_ZONE_STATUS^MDC¦3¦Active^Active¦¦¦¦¦¦F¦¦¦¦¦¦¦
206 | NM | 731840 | MDC_IDC_SET_ZONE_DETECTION_INTERVAL | MDC | 3 | 165 |  | ms |  | F |  |  |  |  | OBX¦206¦NM¦731840^MDC_IDC_SET_ZONE_DETECTION_INTERVAL^MDC¦3¦165¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
207 | NM | 731904 | MDC_IDC_SET_ZONE_DETECTION_BEATS_NUMERATOR | MDC | 3 | 9 |  | {beats} |  | F |  |  |  |  | OBX¦207¦NM¦731904^MDC_IDC_SET_ZONE_DETECTION_BEATS_NUMERATOR^MDC¦3¦9¦{beats}¦¦¦¦¦F¦¦¦¦¦¦¦
208 | NM | 731968 | MDC_IDC_SET_ZONE_DETECTION_BEATS_DENOMINATOR | MDC | 3 | 12 |  | {beats} |  | F |  |  |  |  | OBX¦208¦NM¦731968^MDC_IDC_SET_ZONE_DETECTION_BEATS_DENOMINATOR^MDC¦3¦12¦{beats}¦¦¦¦¦F¦¦¦¦¦¦¦
209 | ST | 732032 | MDC_IDC_SET_ZONE_DETECTION_DETAILS | MDC | 3 | Progressive Therapy |  |  |  | F |  |  |  |  | OBX¦209¦ST¦732032^MDC_IDC_SET_ZONE_DETECTION_DETAILS^MDC¦3¦Progressive Therapy¦¦¦¦¦¦F¦¦¦¦¦¦¦
210 | CWE | 732097 | MDC_IDC_SET_ZONE_TYPE_ATP_1 | MDC | 3 | RampScan | ^Ramp+Scan |  |  | F |  |  |  |  | OBX¦210¦CWE¦732097^MDC_IDC_SET_ZONE_TYPE_ATP_1^MDC¦3¦RampScan^Ramp+Scan¦¦¦¦¦¦F¦¦¦¦¦¦¦
211 | NM | 732161 | MDC_IDC_SET_ZONE_NUM_ATP_SEQS_1 | MDC | 3 | 5 |  | {seq} |  | F |  |  |  |  | OBX¦211¦NM¦732161^MDC_IDC_SET_ZONE_NUM_ATP_SEQS_1^MDC¦3¦5¦{seq}¦¦¦¦¦F¦¦¦¦¦¦¦
212 | NM | 732225 | MDC_IDC_SET_ZONE_SHOCK_ENERGY_1 | MDC | 3 | 20 |  | J |  | F |  |  |  |  | OBX¦212¦NM¦732225^MDC_IDC_SET_ZONE_SHOCK_ENERGY_1^MDC¦3¦20¦J¦¦¦¦¦F¦¦¦¦¦¦¦
213 | NM | 732289 | MDC_IDC_SET_ZONE_NUM_SHOCKS_1 | MDC | 3 | 1 |  | {shocks} |  | F |  |  |  |  | OBX¦213¦NM¦732289^MDC_IDC_SET_ZONE_NUM_SHOCKS_1^MDC¦3¦1¦{shocks}¦¦¦¦¦F¦¦¦¦¦¦¦
214 | NM | 732226 | MDC_IDC_SET_ZONE_SHOCK_ENERGY_2 | MDC | 3 | 30 |  | J |  | F |  |  |  |  | OBX¦214¦NM¦732226^MDC_IDC_SET_ZONE_SHOCK_ENERGY_2^MDC¦3¦30¦J¦¦¦¦¦F¦¦¦¦¦¦¦
215 | NM | 732290 | MDC_IDC_SET_ZONE_NUM_SHOCKS_2 | MDC | 3 | 5 |  | {shocks} |  | F |  |  |  |  | OBX¦215¦NM¦732290^MDC_IDC_SET_ZONE_NUM_SHOCKS_2^MDC¦3¦5¦{shocks}¦¦¦¦¦F¦¦¦¦¦¦¦
216 | DTM | 737617 | MDC_IDC_STAT_HEART_RATE_DTM_START | MDC |  | 20110925000000 |  |  |  | F |  |  |  |  | OBX¦216¦DTM¦737617^MDC_IDC_STAT_HEART_RATE_DTM_START^MDC¦¦20110925000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
217 | DTM | 737618 | MDC_IDC_STAT_HEART_RATE_DTM_END | MDC |  | 20111025100000 |  |  |  | F |  |  |  |  | OBX¦217¦DTM¦737618^MDC_IDC_STAT_HEART_RATE_DTM_END^MDC¦¦20111025100000¦¦¦¦¦¦F¦¦¦¦¦¦¦
218 | NM | 737635 | MDC_IDC_STAT_HEART_RATE_ATRIAL_MEAN | MDC |  | 72 |  | {beats}/min |  | F |  |  |  |  | OBX¦218¦NM¦737635^MDC_IDC_STAT_HEART_RATE_ATRIAL_MEAN^MDC¦¦72¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
219 | NM | 737633 | MDC_IDC_STAT_HEART_RATE_ATRIAL_MAX | MDC |  | 165 |  | {beats}/min |  | F |  |  |  |  | OBX¦219¦NM¦737633^MDC_IDC_STAT_HEART_RATE_ATRIAL_MAX^MDC¦¦165¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
220 | NM | 737634 | MDC_IDC_STAT_HEART_RATE_ATRIAL_MIN | MDC |  | 60 |  | {beats}/min |  | F |  |  |  |  | OBX¦220¦NM¦737634^MDC_IDC_STAT_HEART_RATE_ATRIAL_MIN^MDC¦¦60¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
221 | NM | 737651 | MDC_IDC_STAT_HEART_RATE_VENTRICULAR_MEAN | MDC |  | 72 |  | {beats}/min |  | F |  |  |  |  | OBX¦221¦NM¦737651^MDC_IDC_STAT_HEART_RATE_VENTRICULAR_MEAN^MDC¦¦72¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
222 | NM | 737649 | MDC_IDC_STAT_HEART_RATE_VENTRICULAR_MAX | MDC |  | 165 |  | {beats}/min |  | F |  |  |  |  | OBX¦222¦NM¦737649^MDC_IDC_STAT_HEART_RATE_VENTRICULAR_MAX^MDC¦¦165¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
223 | NM | 737650 | MDC_IDC_STAT_HEART_RATE_VENTRICULAR_MIN | MDC |  | 60 |  | {beats}/min |  | F |  |  |  |  | OBX¦223¦NM¦737650^MDC_IDC_STAT_HEART_RATE_VENTRICULAR_MIN^MDC¦¦60¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
224 | DTM | 737505 | MDC_IDC_STAT_BRADY_DTM_START | MDC |  | 20110925000000 |  |  |  | F |  |  |  |  | OBX¦224¦DTM¦737505^MDC_IDC_STAT_BRADY_DTM_START^MDC¦¦20110925000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
225 | DTM | 737506 | MDC_IDC_STAT_BRADY_DTM_END | MDC |  | 20111025100000 |  |  |  | F |  |  |  |  | OBX¦225¦DTM¦737506^MDC_IDC_STAT_BRADY_DTM_END^MDC¦¦20111025100000¦¦¦¦¦¦F¦¦¦¦¦¦¦
226 | NM | 737520 | MDC_IDC_STAT_BRADY_RA_PERCENT_PACED | MDC |  | 50 |  | % |  | F |  |  |  |  | OBX¦226¦NM¦737520^MDC_IDC_STAT_BRADY_RA_PERCENT_PACED^MDC¦¦50¦%¦¦¦¦¦F¦¦¦¦¦¦¦
227 | NM | 737536 | MDC_IDC_STAT_BRADY_RV_PERCENT_PACED | MDC |  | 30 |  | % |  | F |  |  |  |  | OBX¦227¦NM¦737536^MDC_IDC_STAT_BRADY_RV_PERCENT_PACED^MDC¦¦30¦%¦¦¦¦¦F¦¦¦¦¦¦¦
228 | NM | 737552 | MDC_IDC_STAT_BRADY_AP_VP_PERCENT | MDC |  | 10 |  | % |  | F |  |  |  |  | OBX¦228¦NM¦737552^MDC_IDC_STAT_BRADY_AP_VP_PERCENT^MDC¦¦10¦%¦¦¦¦¦F¦¦¦¦¦¦¦
229 | NM | 737568 | MDC_IDC_STAT_BRADY_AS_VP_PERCENT | MDC |  | 20 |  | % |  | F |  |  |  |  | OBX¦229¦NM¦737568^MDC_IDC_STAT_BRADY_AS_VP_PERCENT^MDC¦¦20¦%¦¦¦¦¦F¦¦¦¦¦¦¦
230 | NM | 737584 | MDC_IDC_STAT_BRADY_AP_VS_PERCENT | MDC |  | 40 |  | % |  | F |  |  |  |  | OBX¦230¦NM¦737584^MDC_IDC_STAT_BRADY_AP_VS_PERCENT^MDC¦¦40¦%¦¦¦¦¦F¦¦¦¦¦¦¦
231 | NM | 737600 | MDC_IDC_STAT_BRADY_AS_VS_PERCENT | MDC |  | 60 |  | % |  | F |  |  |  |  | OBX¦231¦NM¦737600^MDC_IDC_STAT_BRADY_AS_VS_PERCENT^MDC¦¦60¦%¦¦¦¦¦F¦¦¦¦¦¦¦
232 | DTM | 737665 | MDC_IDC_STAT_AT_DTM_START | MDC |  | 20110925000000 |  |  |  | F |  |  |  |  | OBX¦232¦DTM¦737665^MDC_IDC_STAT_AT_DTM_START^MDC¦¦20110925000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
233 | DTM | 737666 | MDC_IDC_STAT_AT_DTM_END | MDC |  | 20111025100000 |  |  |  | F |  |  |  |  | OBX¦233¦DTM¦737666^MDC_IDC_STAT_AT_DTM_END^MDC¦¦20111025100000¦¦¦¦¦¦F¦¦¦¦¦¦¦
234 | NM | 737696 | MDC_IDC_STAT_AT_BURDEN_PERCENT | MDC |  | 10 |  | % |  | F |  |  |  |  | OBX¦234¦NM¦737696^MDC_IDC_STAT_AT_BURDEN_PERCENT^MDC¦¦10¦%¦¦¦¦¦F¦¦¦¦¦¦¦
235 | NM | 737712 | MDC_IDC_STAT_AT_MODE_SW_PERCENT_TIME | MDC |  | 5 |  | % |  | F |  |  |  |  | OBX¦235¦NM¦737712^MDC_IDC_STAT_AT_MODE_SW_PERCENT_TIME^MDC¦¦5¦%¦¦¦¦¦F¦¦¦¦¦¦¦
236 | NM | 737728 | MDC_IDC_STAT_AT_MODE_SW_PERCENT_TIME_PER_DAY | MDC |  | 5 |  | % |  | F |  |  |  |  | OBX¦236¦NM¦737728^MDC_IDC_STAT_AT_MODE_SW_PERCENT_TIME_PER_DAY^MDC¦¦5¦%¦¦¦¦¦F¦¦¦¦¦¦¦
237 | NM | 737760 | MDC_IDC_STAT_AT_MODE_SW_COUNT_PER_DAY | MDC |  | 360 |  | {switches} |  | F |  |  |  |  | OBX¦237¦NM¦737760^MDC_IDC_STAT_AT_MODE_SW_COUNT_PER_DAY^MDC¦¦360¦{switches}¦¦¦¦¦F¦¦¦¦¦¦¦
238 | NM | 737744 | MDC_IDC_STAT_AT_MODE_SW_COUNT | MDC |  | 10800 |  | {switches} |  | F |  |  |  |  | OBX¦238¦NM¦737744^MDC_IDC_STAT_AT_MODE_SW_COUNT^MDC¦¦10800¦{switches}¦¦¦¦¦F¦¦¦¦¦¦¦
239 | NM | 737680 | MDC_IDC_STAT_AT_MODE_SW_MAX_DURATION | MDC |  | 60 |  | s |  | F |  |  |  |  | OBX¦239¦NM¦737680^MDC_IDC_STAT_AT_MODE_SW_MAX_DURATION^MDC¦¦60¦s¦¦¦¦¦F¦¦¦¦¦¦¦
240 | DTM | 737777 | MDC_IDC_STAT_CRT_DTM_START | MDC |  | 20110925000000 |  |  |  | F |  |  |  |  | OBX¦240¦DTM¦737777^MDC_IDC_STAT_CRT_DTM_START^MDC¦¦20110925000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
241 | DTM | 737778 | MDC_IDC_STAT_CRT_DTM_END | MDC |  | 20111025100000 |  |  |  | F |  |  |  |  | OBX¦241¦DTM¦737778^MDC_IDC_STAT_CRT_DTM_END^MDC¦¦20111025100000¦¦¦¦¦¦F¦¦¦¦¦¦¦
242 | NM | 737792 | MDC_IDC_STAT_CRT_LV_PERCENT_PACED | MDC |  | 95 |  | % |  | F |  |  |  |  | OBX¦242¦NM¦737792^MDC_IDC_STAT_CRT_LV_PERCENT_PACED^MDC¦¦95¦%¦¦¦¦¦F¦¦¦¦¦¦¦
243 | NM | 737808 | MDC_IDC_STAT_CRT_PERCENT_PACED | MDC |  | 80 |  | % |  | F |  |  |  |  | OBX¦243¦NM¦737808^MDC_IDC_STAT_CRT_PERCENT_PACED^MDC¦¦80¦%¦¦¦¦¦F¦¦¦¦¦¦¦
244 | NM | 737824 | MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_DELIVERED_RECENT | MDC |  | 1 |  | {shocks} |  | F |  |  |  |  | OBX¦244¦NM¦737824^MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_DELIVERED_RECENT^MDC¦¦1¦{shocks}¦¦¦¦¦F¦¦¦¦¦¦¦
245 | NM | 737840 | MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_DELIVERED_TOTAL | MDC |  | 4 |  | {shocks} |  | F |  |  |  |  | OBX¦245¦NM¦737840^MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_DELIVERED_TOTAL^MDC¦¦4¦{shocks}¦¦¦¦¦F¦¦¦¦¦¦¦
246 | NM | 737856 | MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_ABORTED_RECENT | MDC |  | 0 |  | {shocks} |  | F |  |  |  |  | OBX¦246¦NM¦737856^MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_ABORTED_RECENT^MDC¦¦0¦{shocks}¦¦¦¦¦F¦¦¦¦¦¦¦
247 | NM | 737872 | MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_ABORTED_TOTAL | MDC |  | 1 |  | {shocks} |  | F |  |  |  |  | OBX¦247¦NM¦737872^MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_ABORTED_TOTAL^MDC¦¦1¦{shocks}¦¦¦¦¦F¦¦¦¦¦¦¦
248 | NM | 737888 | MDC_IDC_STAT_TACHYTHERAPY_ATP_DELIVERED_RECENT | MDC |  | 1 |  | {seq} |  | F |  |  |  |  | OBX¦248¦NM¦737888^MDC_IDC_STAT_TACHYTHERAPY_ATP_DELIVERED_RECENT^MDC¦¦1¦{seq}¦¦¦¦¦F¦¦¦¦¦¦¦
249 | NM | 737904 | MDC_IDC_STAT_TACHYTHERAPY_ATP_DELIVERED_TOTAL | MDC |  | 3 |  | {seq} |  | F |  |  |  |  | OBX¦249¦NM¦737904^MDC_IDC_STAT_TACHYTHERAPY_ATP_DELIVERED_TOTAL^MDC¦¦3¦{seq}¦¦¦¦¦F¦¦¦¦¦¦¦
250 | DTM | 737921 | MDC_IDC_STAT_TACHYTHERAPY_TOTAL_DTM_START | MDC |  | 20110501000000 |  |  |  | F |  |  |  |  | OBX¦250¦DTM¦737921^MDC_IDC_STAT_TACHYTHERAPY_TOTAL_DTM_START^MDC¦¦20110501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
251 | DTM | 737922 | MDC_IDC_STAT_TACHYTHERAPY_TOTAL_DTM_END | MDC |  | 20111025100000 |  |  |  | F |  |  |  |  | OBX¦251¦DTM¦737922^MDC_IDC_STAT_TACHYTHERAPY_TOTAL_DTM_END^MDC¦¦20111025100000¦¦¦¦¦¦F¦¦¦¦¦¦¦
252 | DTM | 737937 | MDC_IDC_STAT_TACHYTHERAPY_RECENT_DTM_START | MDC |  | 20110925000000 |  |  |  | F |  |  |  |  | OBX¦252¦DTM¦737937^MDC_IDC_STAT_TACHYTHERAPY_RECENT_DTM_START^MDC¦¦20110925000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
253 | DTM | 737938 | MDC_IDC_STAT_TACHYTHERAPY_RECENT_DTM_END | MDC |  | 20111025100000 |  |  |  | F |  |  |  |  | OBX¦253¦DTM¦737938^MDC_IDC_STAT_TACHYTHERAPY_RECENT_DTM_END^MDC¦¦20111025100000¦¦¦¦¦¦F¦¦¦¦¦¦¦
254 | CWE | 737952 | MDC_IDC_STAT_EPISODE_TYPE | MDC | 1 | Epis_VF | ^VF |  |  | F |  |  |  |  | OBX¦254¦CWE¦737952^MDC_IDC_STAT_EPISODE_TYPE^MDC¦1¦Epis_VF^VF¦¦¦¦¦¦F¦¦¦¦¦¦¦
255 | CWE | 737968 | MDC_IDC_STAT_EPISODE_TYPE_INDUCED | MDC | 1 | NO | ^NO |  |  | F |  |  |  |  | OBX¦255¦CWE¦737968^MDC_IDC_STAT_EPISODE_TYPE_INDUCED^MDC¦1¦NO^NO¦¦¦¦¦¦F¦¦¦¦¦¦¦
256 | CWE | 737984 | MDC_IDC_STAT_EPISODE_VENDOR_TYPE | MDC | 1 | VF | ^VF |  |  | F |  |  |  |  | OBX¦256¦CWE¦737984^MDC_IDC_STAT_EPISODE_VENDOR_TYPE^MDC¦1¦VF^VF¦¦¦¦¦¦F¦¦¦¦¦¦¦
257 | NM | 738000 | MDC_IDC_STAT_EPISODE_RECENT_COUNT | MDC | 1 | 1 |  |  |  | F |  |  |  |  | OBX¦257¦NM¦738000^MDC_IDC_STAT_EPISODE_RECENT_COUNT^MDC¦1¦1¦¦¦¦¦¦F¦¦¦¦¦¦¦
258 | NM | 738032 | MDC_IDC_STAT_EPISODE_TOTAL_COUNT | MDC | 1 | 4 |  |  |  | F |  |  |  |  | OBX¦258¦NM¦738032^MDC_IDC_STAT_EPISODE_TOTAL_COUNT^MDC¦1¦4¦¦¦¦¦¦F¦¦¦¦¦¦¦
259 | CWE | 737952 | MDC_IDC_STAT_EPISODE_TYPE | MDC | 2 | Epis_VT | ^VT |  |  | F |  |  |  |  | OBX¦259¦CWE¦737952^MDC_IDC_STAT_EPISODE_TYPE^MDC¦2¦Epis_VT^VT¦¦¦¦¦¦F¦¦¦¦¦¦¦
260 | CWE | 737968 | MDC_IDC_STAT_EPISODE_TYPE_INDUCED | MDC | 2 | NO | ^NO |  |  | F |  |  |  |  | OBX¦260¦CWE¦737968^MDC_IDC_STAT_EPISODE_TYPE_INDUCED^MDC¦2¦NO^NO¦¦¦¦¦¦F¦¦¦¦¦¦¦
261 | CWE | 737984 | MDC_IDC_STAT_EPISODE_VENDOR_TYPE | MDC | 2 | VT1 | ^VT1 |  |  | F |  |  |  |  | OBX¦261¦CWE¦737984^MDC_IDC_STAT_EPISODE_VENDOR_TYPE^MDC¦2¦VT1^VT1¦¦¦¦¦¦F¦¦¦¦¦¦¦
262 | NM | 738000 | MDC_IDC_STAT_EPISODE_RECENT_COUNT | MDC | 2 | 1 |  |  |  | F |  |  |  |  | OBX¦262¦NM¦738000^MDC_IDC_STAT_EPISODE_RECENT_COUNT^MDC¦2¦1¦¦¦¦¦¦F¦¦¦¦¦¦¦
263 | NM | 738032 | MDC_IDC_STAT_EPISODE_TOTAL_COUNT | MDC | 2 | 0 |  |  |  | F |  |  |  |  | OBX¦263¦NM¦738032^MDC_IDC_STAT_EPISODE_TOTAL_COUNT^MDC¦2¦0¦¦¦¦¦¦F¦¦¦¦¦¦¦
264 | CWE | 737952 | MDC_IDC_STAT_EPISODE_TYPE | MDC | 3 | Epis_ATAF | ^AT/AF |  |  | F |  |  |  |  | OBX¦264¦CWE¦737952^MDC_IDC_STAT_EPISODE_TYPE^MDC¦3¦Epis_ATAF^AT/AF¦¦¦¦¦¦F¦¦¦¦¦¦¦
265 | CWE | 737968 | MDC_IDC_STAT_EPISODE_TYPE_INDUCED | MDC | 3 | NO | ^NO |  |  | F |  |  |  |  | OBX¦265¦CWE¦737968^MDC_IDC_STAT_EPISODE_TYPE_INDUCED^MDC¦3¦NO^NO¦¦¦¦¦¦F¦¦¦¦¦¦¦
266 | CWE | 737984 | MDC_IDC_STAT_EPISODE_VENDOR_TYPE | MDC | 3 | ATAF | ^ATAF |  |  | F |  |  |  |  | OBX¦266¦CWE¦737984^MDC_IDC_STAT_EPISODE_VENDOR_TYPE^MDC¦3¦ATAF^ATAF¦¦¦¦¦¦F¦¦¦¦¦¦¦
267 | NM | 738000 | MDC_IDC_STAT_EPISODE_RECENT_COUNT | MDC | 3 | 2 |  |  |  | F |  |  |  |  | OBX¦267¦NM¦738000^MDC_IDC_STAT_EPISODE_RECENT_COUNT^MDC¦3¦2¦¦¦¦¦¦F¦¦¦¦¦¦¦
268 | NM | 738032 | MDC_IDC_STAT_EPISODE_TOTAL_COUNT | MDC | 3 | 150 |  |  |  | F |  |  |  |  | OBX¦268¦NM¦738032^MDC_IDC_STAT_EPISODE_TOTAL_COUNT^MDC¦3¦150¦¦¦¦¦¦F¦¦¦¦¦¦¦
269 | NM | 738017 | MDC_IDC_STAT_EPISODE_RECENT_COUNT_DTM_START | MDC |  | 20110501000000 |  |  |  | F |  |  |  |  | OBX¦269¦NM¦738017^MDC_IDC_STAT_EPISODE_RECENT_COUNT_DTM_START^MDC¦¦20110501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
270 | NM | 738018 | MDC_IDC_STAT_EPISODE_RECENT_COUNT_DTM_END | MDC |  | 20111025100000 |  |  |  | F |  |  |  |  | OBX¦270¦NM¦738018^MDC_IDC_STAT_EPISODE_RECENT_COUNT_DTM_END^MDC¦¦20111025100000¦¦¦¦¦¦F¦¦¦¦¦¦¦
271 | NM | 738049 | MDC_IDC_STAT_EPISODE_TOTAL_COUNT_DTM_START | MDC |  | 20110925000000 |  |  |  | F |  |  |  |  | OBX¦271¦NM¦738049^MDC_IDC_STAT_EPISODE_TOTAL_COUNT_DTM_START^MDC¦¦20110925000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
272 | NM | 738050 | MDC_IDC_STAT_EPISODE_TOTAL_COUNT_DTM_END | MDC |  | 20111025100000 |  |  |  | F |  |  |  |  | OBX¦272¦NM¦738050^MDC_IDC_STAT_EPISODE_TOTAL_COUNT_DTM_END^MDC¦¦20111025100000¦¦¦¦¦¦F¦¦¦¦¦¦¦
273 | ST | 739536 | MDC_IDC_EPISODE_ID | MDC | 1 | 1000 |  |  |  | F |  |  |  |  | OBX¦273¦ST¦739536^MDC_IDC_EPISODE_ID^MDC¦1¦1000¦¦¦¦¦¦F¦¦¦¦¦¦¦
274 | DTM | 739552 | MDC_IDC_EPISODE_DTM | MDC | 1 | 20110927000000 |  |  |  | F |  |  |  |  | OBX¦274¦DTM¦739552^MDC_IDC_EPISODE_DTM^MDC¦1¦20110927000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
275 | CWE | 739568 | MDC_IDC_EPISODE_TYPE | MDC | 1 | Epis_VT | ^VT |  |  | F |  |  |  |  | OBX¦275¦CWE¦739568^MDC_IDC_EPISODE_TYPE^MDC¦1¦Epis_VT^VT¦¦¦¦¦¦F¦¦¦¦¦¦¦
276 | CWE | 739584 | MDC_IDC_EPISODE_TYPE_INDUCED | MDC | 1 | NO | ^NO |  |  | F |  |  |  |  | OBX¦276¦CWE¦739584^MDC_IDC_EPISODE_TYPE_INDUCED^MDC¦1¦NO^NO¦¦¦¦¦¦F¦¦¦¦¦¦¦
277 | CWE | 739600 | MDC_IDC_EPISODE_VENDOR_TYPE | MDC | 1 | VT1 | ^VT1 |  |  | F |  |  |  |  | OBX¦277¦CWE¦739600^MDC_IDC_EPISODE_VENDOR_TYPE^MDC¦1¦VT1^VT1¦¦¦¦¦¦F¦¦¦¦¦¦¦
278 | NM | 739648 | MDC_IDC_EPISODE_VENTRICULAR_INTERVAL_AT_DETECTION | MDC | 1 | 400 |  | ms |  | F |  |  |  |  | OBX¦278¦NM¦739648^MDC_IDC_EPISODE_VENTRICULAR_INTERVAL_AT_DETECTION^MDC¦1¦400¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
279 | NM | 739664 | MDC_IDC_EPISODE_VENTRICULAR_INTERVAL_AT_TERMINATION | MDC | 1 | 820 |  | ms |  | F |  |  |  |  | OBX¦279¦NM¦739664^MDC_IDC_EPISODE_VENTRICULAR_INTERVAL_AT_TERMINATION^MDC¦1¦820¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
280 | ST | 739680 | MDC_IDC_EPISODE_DETECTION_THERAPY_DETAILS | MDC | 1 | 2 ATP, 5x 30J |  |  |  | F |  |  |  |  | OBX¦280¦ST¦739680^MDC_IDC_EPISODE_DETECTION_THERAPY_DETAILS^MDC¦1¦2 ATP, 5x 30J¦¦¦¦¦¦F¦¦¦¦¦¦¦
281 | CWE | 739696 | MDC_IDC_EPISODE_THERAPY_RESULT | MDC | 1 | Successful | ^Successful |  |  | F |  |  |  |  | OBX¦281¦CWE¦739696^MDC_IDC_EPISODE_THERAPY_RESULT^MDC¦1¦Successful^Successful¦¦¦¦¦¦F¦¦¦¦¦¦¦
282 | NM | 739712 | MDC_IDC_EPISODE_DURATION | MDC | 1 | 27 |  | s |  | F |  |  |  |  | OBX¦282¦NM¦739712^MDC_IDC_EPISODE_DURATION^MDC¦1¦27¦s¦¦¦¦¦F¦¦¦¦¦¦¦
283 | ST | 739536 | MDC_IDC_EPISODE_ID | MDC | 2 | 1001 |  |  |  | F |  |  |  |  | OBX¦283¦ST¦739536^MDC_IDC_EPISODE_ID^MDC¦2¦1001¦¦¦¦¦¦F¦¦¦¦¦¦¦
284 | DTM | 739552 | MDC_IDC_EPISODE_DTM | MDC | 2 | 20110927000000 |  |  |  | F |  |  |  |  | OBX¦284¦DTM¦739552^MDC_IDC_EPISODE_DTM^MDC¦2¦20110927000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
285 | CWE | 739568 | MDC_IDC_EPISODE_TYPE | MDC | 2 | Epis_ATAF | ^AT/AF |  |  | F |  |  |  |  | OBX¦285¦CWE¦739568^MDC_IDC_EPISODE_TYPE^MDC¦2¦Epis_ATAF^AT/AF¦¦¦¦¦¦F¦¦¦¦¦¦¦
286 | CWE | 739584 | MDC_IDC_EPISODE_TYPE_INDUCED | MDC | 2 | NO | ^NO |  |  | F |  |  |  |  | OBX¦286¦CWE¦739584^MDC_IDC_EPISODE_TYPE_INDUCED^MDC¦2¦NO^NO¦¦¦¦¦¦F¦¦¦¦¦¦¦
287 | CWE | 739600 | MDC_IDC_EPISODE_VENDOR_TYPE | MDC | 2 | ATR | ^ATR |  |  | F |  |  |  |  | OBX¦287¦CWE¦739600^MDC_IDC_EPISODE_VENDOR_TYPE^MDC¦2¦ATR^ATR¦¦¦¦¦¦F¦¦¦¦¦¦¦
288 | NM | 739616 | MDC_IDC_EPISODE_ATRIAL_INTERVAL_AT_DETECTION | MDC | 2 | 200 |  | ms |  | F |  |  |  |  | OBX¦288¦NM¦739616^MDC_IDC_EPISODE_ATRIAL_INTERVAL_AT_DETECTION^MDC¦2¦200¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
289 | NM | 739632 | MDC_IDC_EPISODE_ATRIAL_INTERVAL_AT_TERMINATION | MDC | 2 | 60 |  | ms |  | F |  |  |  |  | OBX¦289¦NM¦739632^MDC_IDC_EPISODE_ATRIAL_INTERVAL_AT_TERMINATION^MDC¦2¦60¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
290 | ST | 739680 | MDC_IDC_EPISODE_DETECTION_THERAPY_DETAILS | MDC | 2 | 10 ATP |  |  |  | F |  |  |  |  | OBX¦290¦ST¦739680^MDC_IDC_EPISODE_DETECTION_THERAPY_DETAILS^MDC¦2¦10 ATP¦¦¦¦¦¦F¦¦¦¦¦¦¦
291 | CWE | 739696 | MDC_IDC_EPISODE_THERAPY_RESULT | MDC | 2 | Successful | ^Successful |  |  | F |  |  |  |  | OBX¦291¦CWE¦739696^MDC_IDC_EPISODE_THERAPY_RESULT^MDC¦2¦Successful^Successful¦¦¦¦¦¦F¦¦¦¦¦¦¦
292 | NM | 739712 | MDC_IDC_EPISODE_DURATION | MDC | 2 | 10800 |  | s |  | F |  |  |  |  | OBX¦292¦NM¦739712^MDC_IDC_EPISODE_DURATION^MDC¦2¦10800¦s¦¦¦¦¦F¦¦¦¦¦¦¦
IMPLANT MESSAGE
OBR¦2¦¦123456¦IMPLANT¦¦¦20090501031234¦20090501031234¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦F
NTE¦1¦P¦Initial implant observation (a comment can go here concerning the OBR/interrogation section).
1 | CWE | 720897 | MDC_IDC_DEV_TYPE | MDC |  | CRT_D | ^Cardiac Resynchronization Therapy - Defibrillator |  |  | F |  |  |  |  | OBX¦1¦CWE¦720897^MDC_IDC_DEV_TYPE^MDC¦¦CRT_D^Cardiac Resynchronization Therapy - Defibrillator¦¦¦¦¦¦F¦¦¦¦¦¦¦
2 | ST | 720898 | MDC_IDC_DEV_MODEL | MDC |  | Device Model Name |  |  |  | F |  |  |  |  | OBX¦2¦ST¦720898^MDC_IDC_DEV_MODEL^MDC¦¦Device Model Name¦¦¦¦¦¦F¦¦¦¦¦¦¦
3 | ST | 720899 | MDC_IDC_DEV_SERIAL | MDC |  | 5867463524 |  |  |  | F |  |  |  |  | OBX¦3¦ST¦720899^MDC_IDC_DEV_SERIAL^MDC¦¦5867463524¦¦¦¦¦¦F¦¦¦¦¦¦¦
4 | CWE | 720900 | MDC_IDC_DEV_MFG | MDC |  | STJ | ^St.Jude Medical |  |  | F |  |  |  |  | OBX¦4¦CWE¦720900^MDC_IDC_DEV_MFG^MDC¦¦STJ^St.Jude Medical¦¦¦¦¦¦F¦¦¦¦¦¦¦
5 | DTM | 720901 | MDC_IDC_DEV_IMPLANT_DT | MDC |  | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦5¦DTM¦720901^MDC_IDC_DEV_IMPLANT_DT^MDC¦¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
6 | ST | 720902 | MDC_IDC_DEV_IMPLANTER | MDC |  | Dr. Miller |  |  |  | F |  |  |  |  | OBX¦6¦ST¦720902^MDC_IDC_DEV_IMPLANTER^MDC¦¦Dr. Miller¦¦¦¦¦¦F¦¦¦¦¦¦¦
7 | ST | 720903 | MDC_IDC_DEV_IMPLANTER_CONTACT_INFO | MDC |  | Phone: +1 12 345 6789 |  |  |  | F |  |  |  |  | OBX¦7¦ST¦720903^MDC_IDC_DEV_IMPLANTER_CONTACT_INFO^MDC¦¦Phone: +1 12 345 6789¦¦¦¦¦¦F¦¦¦¦¦¦¦
8 | ST | 720904 | MDC_IDC_DEV_IMPLANTING_FACILITY | MDC |  | Main Heart Center New York |  |  |  | F |  |  |  |  | OBX¦8¦ST¦720904^MDC_IDC_DEV_IMPLANTING_FACILITY^MDC¦¦Main Heart Center New York¦¦¦¦¦¦F¦¦¦¦¦¦¦
9 | ST | 720961 | MDC_IDC_LEAD_MODEL | MDC | 1 | Supersense |  |  |  | F |  |  |  |  | OBX¦9¦ST¦720961^MDC_IDC_LEAD_MODEL^MDC¦1¦Supersense¦¦¦¦¦¦F¦¦¦¦¦¦¦
10 | ST | 720962 | MDC_IDC_LEAD_SERIAL | MDC | 1 | 1234567812 |  |  |  | F |  |  |  |  | OBX¦10¦ST¦720962^MDC_IDC_LEAD_SERIAL^MDC¦1¦1234567812¦¦¦¦¦¦F¦¦¦¦¦¦¦
11 | CWE | 720963 | MDC_IDC_LEAD_MFG | MDC | 1 | BIO | ^Biotronik |  |  | F |  |  |  |  | OBX¦11¦CWE¦720963^MDC_IDC_LEAD_MFG^MDC¦1¦BIO^Biotronik¦¦¦¦¦¦F¦¦¦¦¦¦¦
12 | DTM | 720964 | MDC_IDC_LEAD_IMPLANT_DT | MDC | 1 | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦12¦DTM¦720964^MDC_IDC_LEAD_IMPLANT_DT^MDC¦1¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
13 | CWE | 720965 | MDC_IDC_LEAD_POLARITY_TYPE | MDC | 1 | UNI | ^Unipolar Lead |  |  | F |  |  |  |  | OBX¦13¦CWE¦720965^MDC_IDC_LEAD_POLARITY_TYPE^MDC¦1¦UNI^Unipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
14 | CWE | 720966 | MDC_IDC_LEAD_LOCATION | MDC | 1 | RA | ^Right Atrium |  |  | F |  |  |  |  | OBX¦14¦CWE¦720966^MDC_IDC_LEAD_LOCATION^MDC¦1¦RA^Right Atrium¦¦¦¦¦¦F¦¦¦¦¦¦¦
15 | CWE | 720967 | MDC_IDC_LEAD_LOCATION_DETAIL_1 | MDC | 1 | Appendage | ^Appendage |  |  | F |  |  |  |  | OBX¦15¦CWE¦720967^MDC_IDC_LEAD_LOCATION_DETAIL_1^MDC¦1¦Appendage^Appendage¦¦¦¦¦¦F¦¦¦¦¦¦¦
16 | CWE | 720970 | MDC_IDC_LEAD_CONNECTION_STATUS | MDC | 1 | Connected | ^Connected |  |  | F |  |  |  |  | OBX¦16¦CWE¦720970^MDC_IDC_LEAD_CONNECTION_STATUS^MDC¦1¦Connected^Connected¦¦¦¦¦¦F¦¦¦¦¦¦¦
17 | ST | 720971 | MDC_IDC_LEAD_SPECIAL_FUNCTION | MDC | 1 | Pressure Sensor |  |  |  | F |  |  |  |  | OBX¦17¦ST¦720971^MDC_IDC_LEAD_SPECIAL_FUNCTION^MDC¦1¦Pressure Sensor¦¦¦¦¦¦F¦¦¦¦¦¦¦
18 | ST | 720961 | MDC_IDC_LEAD_MODEL | MDC | 2 | Supersense |  |  |  | F |  |  |  |  | OBX¦18¦ST¦720961^MDC_IDC_LEAD_MODEL^MDC¦2¦Supersense¦¦¦¦¦¦F¦¦¦¦¦¦¦
19 | ST | 720962 | MDC_IDC_LEAD_SERIAL | MDC | 2 | 1234567813 |  |  |  | F |  |  |  |  | OBX¦19¦ST¦720962^MDC_IDC_LEAD_SERIAL^MDC¦2¦1234567813¦¦¦¦¦¦F¦¦¦¦¦¦¦
20 | CWE | 720963 | MDC_IDC_LEAD_MFG | MDC | 2 | BSX | ^Boston Scientific |  |  | F |  |  |  |  | OBX¦20¦CWE¦720963^MDC_IDC_LEAD_MFG^MDC¦2¦BSX^Boston Scientific¦¦¦¦¦¦F¦¦¦¦¦¦¦
21 | DTM | 720964 | MDC_IDC_LEAD_IMPLANT_DT | MDC | 2 | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦21¦DTM¦720964^MDC_IDC_LEAD_IMPLANT_DT^MDC¦2¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
22 | CWE | 720965 | MDC_IDC_LEAD_POLARITY_TYPE | MDC | 2 | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦22¦CWE¦720965^MDC_IDC_LEAD_POLARITY_TYPE^MDC¦2¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
23 | CWE | 720966 | MDC_IDC_LEAD_LOCATION | MDC | 2 | RV | ^Right Ventricle |  |  | F |  |  |  |  | OBX¦23¦CWE¦720966^MDC_IDC_LEAD_LOCATION^MDC¦2¦RV^Right Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
24 | CWE | 720967 | MDC_IDC_LEAD_LOCATION_DETAIL_1 | MDC | 2 | Apex | ^Apex |  |  | F |  |  |  |  | OBX¦24¦CWE¦720967^MDC_IDC_LEAD_LOCATION_DETAIL_1^MDC¦2¦Apex^Apex¦¦¦¦¦¦F¦¦¦¦¦¦¦
25 | CWE | 720970 | MDC_IDC_LEAD_CONNECTION_STATUS | MDC | 2 | Connected | ^Connected |  |  | F |  |  |  |  | OBX¦25¦CWE¦720970^MDC_IDC_LEAD_CONNECTION_STATUS^MDC¦2¦Connected^Connected¦¦¦¦¦¦F¦¦¦¦¦¦¦
26 | ST | 720971 | MDC_IDC_LEAD_SPECIAL_FUNCTION | MDC | 2 | N/A |  |  |  | F |  |  |  |  | OBX¦26¦ST¦720971^MDC_IDC_LEAD_SPECIAL_FUNCTION^MDC¦2¦N/A¦¦¦¦¦¦F¦¦¦¦¦¦¦
27 | ST | 720961 | MDC_IDC_LEAD_MODEL | MDC | 3 | Supersense |  |  |  | F |  |  |  |  | OBX¦27¦ST¦720961^MDC_IDC_LEAD_MODEL^MDC¦3¦Supersense¦¦¦¦¦¦F¦¦¦¦¦¦¦
28 | ST | 720962 | MDC_IDC_LEAD_SERIAL | MDC | 3 | 1234567814 |  |  |  | F |  |  |  |  | OBX¦28¦ST¦720962^MDC_IDC_LEAD_SERIAL^MDC¦3¦1234567814¦¦¦¦¦¦F¦¦¦¦¦¦¦
29 | CWE | 720963 | MDC_IDC_LEAD_MFG | MDC | 3 | MDT | ^Medtronic |  |  | F |  |  |  |  | OBX¦29¦CWE¦720963^MDC_IDC_LEAD_MFG^MDC¦3¦MDT^Medtronic¦¦¦¦¦¦F¦¦¦¦¦¦¦
30 | DTM | 720964 | MDC_IDC_LEAD_IMPLANT_DT | MDC | 3 | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦30¦DTM¦720964^MDC_IDC_LEAD_IMPLANT_DT^MDC¦3¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
31 | CWE | 720965 | MDC_IDC_LEAD_POLARITY_TYPE | MDC | 3 | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦31¦CWE¦720965^MDC_IDC_LEAD_POLARITY_TYPE^MDC¦3¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
32 | CWE | 720966 | MDC_IDC_LEAD_LOCATION | MDC | 3 | LV | ^Left Ventricle |  |  | F |  |  |  |  | OBX¦32¦CWE¦720966^MDC_IDC_LEAD_LOCATION^MDC¦3¦LV^Left Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
33 | CWE | 720967 | MDC_IDC_LEAD_LOCATION_DETAIL_1 | MDC | 3 | FreeWall | ^Free Wall |  |  | F |  |  |  |  | OBX¦33¦CWE¦720967^MDC_IDC_LEAD_LOCATION_DETAIL_1^MDC¦3¦FreeWall^Free Wall¦¦¦¦¦¦F¦¦¦¦¦¦¦
34 | CWE | 720970 | MDC_IDC_LEAD_CONNECTION_STATUS | MDC | 3 | Connected | ^Connected |  |  | F |  |  |  |  | OBX¦34¦CWE¦720970^MDC_IDC_LEAD_CONNECTION_STATUS^MDC¦3¦Connected^Connected¦¦¦¦¦¦F¦¦¦¦¦¦¦
35 | ST | 720971 | MDC_IDC_LEAD_SPECIAL_FUNCTION | MDC | 3 | N/A |  |  |  | F |  |  |  |  | OBX¦35¦ST¦720971^MDC_IDC_LEAD_SPECIAL_FUNCTION^MDC¦3¦N/A¦¦¦¦¦¦F¦¦¦¦¦¦¦
36 | DTM | 721025 | MDC_IDC_SESS_DTM | MDC |  | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦36¦DTM¦721025^MDC_IDC_SESS_DTM^MDC¦¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
37 | CWE | 721026 | MDC_IDC_SESS_TYPE | MDC |  | InClinic | ^In Clinic |  |  | F |  |  |  |  | OBX¦37¦CWE¦721026^MDC_IDC_SESS_TYPE^MDC¦¦InClinic^In Clinic¦¦¦¦¦¦F¦¦¦¦¦¦¦
38 | CWE | 721027 | MDC_IDC_SESS_REPROGRAMMED | MDC |  | YES | ^YES |  |  | F |  |  |  |  | OBX¦38¦CWE¦721027^MDC_IDC_SESS_REPROGRAMMED^MDC¦¦YES^YES¦¦¦¦¦¦F¦¦¦¦¦¦¦
39 | DTM | 721028 | MDC_IDC_SESS_DTM_PREVIOUS | MDC |  | 20090925100000 |  |  |  | F |  |  |  |  | OBX¦39¦DTM¦721028^MDC_IDC_SESS_DTM_PREVIOUS^MDC¦¦20090925100000¦¦¦¦¦¦F¦¦¦¦¦¦¦
40 | CWE | 721029 | MDC_IDC_SESS_TYPE_PREVIOUS | MDC |  | InClinic | ^In Clinic |  |  | F |  |  |  |  | OBX¦40¦CWE¦721029^MDC_IDC_SESS_TYPE_PREVIOUS^MDC¦¦InClinic^In Clinic¦¦¦¦¦¦F¦¦¦¦¦¦¦
41 | CWE | 721030 | MDC_IDC_SESS_REPROGRAMMED_PREVIOUS | MDC |  | YES | ^YES |  |  | F |  |  |  |  | OBX¦41¦CWE¦721030^MDC_IDC_SESS_REPROGRAMMED_PREVIOUS^MDC¦¦YES^YES¦¦¦¦¦¦F¦¦¦¦¦¦¦
42 | ST | 721031 | MDC_IDC_SESS_CLINICIAN_NAME | MDC |  | Dr. Anderson |  |  |  | F |  |  |  |  | OBX¦42¦ST¦721031^MDC_IDC_SESS_CLINICIAN_NAME^MDC¦¦Dr. Anderson¦¦¦¦¦¦F¦¦¦¦¦¦¦
43 | ST | 721032 | MDC_IDC_SESS_CLINICIAN_CONTACT_INFORMATION | MDC |  | Phone: +1 12 345 6789 |  |  |  | F |  |  |  |  | OBX¦43¦ST¦721032^MDC_IDC_SESS_CLINICIAN_CONTACT_INFORMATION^MDC¦¦Phone: +1 12 345 6789¦¦¦¦¦¦F¦¦¦¦¦¦¦
44 | ST | 721033 | MDC_IDC_SESS_CLINIC_NAME | MDC |  | Main Heart Center New Jersey |  |  |  | F |  |  |  |  | OBX¦44¦ST¦721033^MDC_IDC_SESS_CLINIC_NAME^MDC¦¦Main Heart Center New Jersey¦¦¦¦¦¦F¦¦¦¦¦¦¦
45 | DTM | 721216 | MDC_IDC_MSMT_BATTERY_DTM | MDC | 1 | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦45¦DTM¦721216^MDC_IDC_MSMT_BATTERY_DTM^MDC¦1¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
46 | CWE | 721280 | MDC_IDC_MSMT_BATTERY_STATUS | MDC | 1 | BOS | ^Beginning of Service |  |  | F |  |  |  |  | OBX¦46¦CWE¦721280^MDC_IDC_MSMT_BATTERY_STATUS^MDC¦1¦BOS^Beginning of Service¦¦¦¦¦¦F¦¦¦¦¦¦¦
47 | NM | 721344 | MDC_IDC_MSMT_BATTERY_VOLTAGE | MDC | 1 | 6.4 |  | V |  | F |  |  |  |  | OBX¦47¦NM¦721344^MDC_IDC_MSMT_BATTERY_VOLTAGE^MDC¦1¦6.4¦V¦¦¦¦¦F¦¦¦¦¦¦¦
48 | NM | 721408 | MDC_IDC_MSMT_BATTERY_IMPEDANCE | MDC | 1 | 2500 |  | Ohm |  | F |  |  |  |  | OBX¦48¦NM¦721408^MDC_IDC_MSMT_BATTERY_IMPEDANCE^MDC¦1¦2500¦Ohm¦¦¦¦¦F¦¦¦¦¦¦¦
49 | NM | 721472 | MDC_IDC_MSMT_BATTERY_REMAINING_LONGEVITY | MDC | 1 | 52 |  | mo |  | F |  |  |  |  | OBX¦49¦NM¦721472^MDC_IDC_MSMT_BATTERY_REMAINING_LONGEVITY^MDC¦1¦52¦mo¦¦¦¦¦F¦¦¦¦¦¦¦
50 | NM | 721536 | MDC_IDC_MSMT_BATTERY_REMAINING_PERCENTAGE | MDC | 1 | 99 |  | % |  | F |  |  |  |  | OBX¦50¦NM¦721536^MDC_IDC_MSMT_BATTERY_REMAINING_PERCENTAGE^MDC¦1¦99¦%¦¦¦¦¦F¦¦¦¦¦¦¦
51 | ST | 721600 | MDC_IDC_MSMT_BATTERY_RRT_TRIGGER | MDC | 1 | Battery voltage < 5.7 V / Cap. Charge time > 12 s |  |  |  | F |  |  |  |  | OBX¦51¦ST¦721600^MDC_IDC_MSMT_BATTERY_RRT_TRIGGER^MDC¦1¦Battery voltage < 5.7 V / Cap. Charge time > 12 s¦¦¦¦¦¦F¦¦¦¦¦¦¦
52 | DTM | 721664 | MDC_IDC_MSMT_CAP_CHARGE_DTM | MDC | 1 | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦52¦DTM¦721664^MDC_IDC_MSMT_CAP_CHARGE_DTM^MDC¦1¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
53 | NM | 721728 | MDC_IDC_MSMT_CAP_CHARGE_TIME | MDC | 1 | 8.4 |  | s |  | F |  |  |  |  | OBX¦53¦NM¦721728^MDC_IDC_MSMT_CAP_CHARGE_TIME^MDC¦1¦8.4¦s¦¦¦¦¦F¦¦¦¦¦¦¦
54 | NM | 721792 | MDC_IDC_MSMT_CAP_CHARGE_ENERGY | MDC | 1 | 36 |  | J |  | F |  |  |  |  | OBX¦54¦NM¦721792^MDC_IDC_MSMT_CAP_CHARGE_ENERGY^MDC¦1¦36¦J¦¦¦¦¦F¦¦¦¦¦¦¦
55 | CWE | 721856 | MDC_IDC_MSMT_CAP_CHARGE_TYPE | MDC | 1 | Reformation | ^Reformation |  |  | F |  |  |  |  | OBX¦55¦CWE¦721856^MDC_IDC_MSMT_CAP_CHARGE_TYPE^MDC¦1¦Reformation^Reformation¦¦¦¦¦¦F¦¦¦¦¦¦¦
56 | DTM | 721920 | MDC_IDC_MSMT_LEADCHNL_RA_DTM | MDC |  | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦56¦DTM¦721920^MDC_IDC_MSMT_LEADCHNL_RA_DTM^MDC¦¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
57 | DTM | 721921 | MDC_IDC_MSMT_LEADCHNL_RA_DTM_START | MDC |  | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦57¦DTM¦721921^MDC_IDC_MSMT_LEADCHNL_RA_DTM_START^MDC¦¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
58 | DTM | 721922 | MDC_IDC_MSMT_LEADCHNL_RA_DTM_END | MDC |  | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦58¦DTM¦721922^MDC_IDC_MSMT_LEADCHNL_RA_DTM_END^MDC¦¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
59 | CWE | 721984 | MDC_IDC_MSMT_LEADCHNL_RA_LEAD_CHANNEL_STATUS | MDC |  | NULL | ^Null |  |  | F |  |  |  |  | OBX¦59¦CWE¦721984^MDC_IDC_MSMT_LEADCHNL_RA_LEAD_CHANNEL_STATUS^MDC¦¦NULL^Null¦¦¦¦¦¦F¦¦¦¦¦¦¦
60 | NM | 722048 | MDC_IDC_MSMT_LEADCHNL_RA_SENSING_INTR_AMPL | MDC |  | 2.5 |  | mV |  | F |  |  |  |  | OBX¦60¦NM¦722048^MDC_IDC_MSMT_LEADCHNL_RA_SENSING_INTR_AMPL^MDC¦¦2.5¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
61 | NM | 722049 | MDC_IDC_MSMT_LEADCHNL_RA_SENSING_INTR_AMPL_MAX | MDC |  | 2.6 |  | mV |  | F |  |  |  |  | OBX¦61¦NM¦722049^MDC_IDC_MSMT_LEADCHNL_RA_SENSING_INTR_AMPL_MAX^MDC¦¦2.6¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
62 | NM | 722050 | MDC_IDC_MSMT_LEADCHNL_RA_SENSING_INTR_AMPL_MIN | MDC |  | 2.2 |  | mV |  | F |  |  |  |  | OBX¦62¦NM¦722050^MDC_IDC_MSMT_LEADCHNL_RA_SENSING_INTR_AMPL_MIN^MDC¦¦2.2¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
63 | NM | 722051 | MDC_IDC_MSMT_LEADCHNL_RA_SENSING_INTR_AMPL_MEAN | MDC |  | 2.5 |  | mV |  | F |  |  |  |  | OBX¦63¦NM¦722051^MDC_IDC_MSMT_LEADCHNL_RA_SENSING_INTR_AMPL_MEAN^MDC¦¦2.5¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
64 | CWE | 722112 | MDC_IDC_MSMT_LEADCHNL_RA_SENSING_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦64¦CWE¦722112^MDC_IDC_MSMT_LEADCHNL_RA_SENSING_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
65 | NM | 722176 | MDC_IDC_MSMT_LEADCHNL_RA_PACING_THRESHOLD_AMPLITUDE | MDC |  | 0.7 |  | V |  | F |  |  |  |  | OBX¦65¦NM¦722176^MDC_IDC_MSMT_LEADCHNL_RA_PACING_THRESHOLD_AMPLITUDE^MDC¦¦0.7¦V¦¦¦¦¦F¦¦¦¦¦¦¦
66 | NM | 722240 | MDC_IDC_MSMT_LEADCHNL_RA_PACING_THRESHOLD_PULSEWIDTH | MDC |  | 0.5 |  | ms |  | F |  |  |  |  | OBX¦66¦NM¦722240^MDC_IDC_MSMT_LEADCHNL_RA_PACING_THRESHOLD_PULSEWIDTH^MDC¦¦0.5¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
67 | CWE | 722304 | MDC_IDC_MSMT_LEADCHNL_RA_PACING_THRESHOLD_MEASUREMENT_METHOD | MDC |  | ProgrammerManual | ^Programmer Manual |  |  | F |  |  |  |  | OBX¦67¦CWE¦722304^MDC_IDC_MSMT_LEADCHNL_RA_PACING_THRESHOLD_MEASUREMENT_METHOD^MDC¦¦ProgrammerManual^Programmer Manual¦¦¦¦¦¦F¦¦¦¦¦¦¦
68 | CWE | 722368 | MDC_IDC_MSMT_LEADCHNL_RA_PACING_THRESHOLD_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦68¦CWE¦722368^MDC_IDC_MSMT_LEADCHNL_RA_PACING_THRESHOLD_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
69 | NM | 722432 | MDC_IDC_MSMT_LEADCHNL_RA_IMPEDANCE_VALUE | MDC |  | 500 |  | Ohm |  | F |  |  |  |  | OBX¦69¦NM¦722432^MDC_IDC_MSMT_LEADCHNL_RA_IMPEDANCE_VALUE^MDC¦¦500¦Ohm¦¦¦¦¦F¦¦¦¦¦¦¦
70 | CWE | 722496 | MDC_IDC_MSMT_LEADCHNL_RA_IMPEDANCE_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦70¦CWE¦722496^MDC_IDC_MSMT_LEADCHNL_RA_IMPEDANCE_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
71 | DTM | 721924 | MDC_IDC_MSMT_LEADCHNL_RV_DTM | MDC |  | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦71¦DTM¦721924^MDC_IDC_MSMT_LEADCHNL_RV_DTM^MDC¦¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
72 | DTM | 721925 | MDC_IDC_MSMT_LEADCHNL_RV_DTM_START | MDC |  | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦72¦DTM¦721925^MDC_IDC_MSMT_LEADCHNL_RV_DTM_START^MDC¦¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
73 | DTM | 721926 | MDC_IDC_MSMT_LEADCHNL_RV_DTM_END | MDC |  | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦73¦DTM¦721926^MDC_IDC_MSMT_LEADCHNL_RV_DTM_END^MDC¦¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
74 | CWE | 721985 | MDC_IDC_MSMT_LEADCHNL_RV_LEAD_CHANNEL_STATUS | MDC |  | CheckLead | ^Check Lead |  |  | F |  |  |  |  | OBX¦74¦CWE¦721985^MDC_IDC_MSMT_LEADCHNL_RV_LEAD_CHANNEL_STATUS^MDC¦¦CheckLead^Check Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
75 | NM | 722052 | MDC_IDC_MSMT_LEADCHNL_RV_SENSING_INTR_AMPL | MDC |  | 4.7 |  | mV |  | F |  |  |  |  | OBX¦75¦NM¦722052^MDC_IDC_MSMT_LEADCHNL_RV_SENSING_INTR_AMPL^MDC¦¦4.7¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
76 | NM | 722053 | MDC_IDC_MSMT_LEADCHNL_RV_SENSING_INTR_AMPL_MAX | MDC |  | 4.9 |  | mV |  | F |  |  |  |  | OBX¦76¦NM¦722053^MDC_IDC_MSMT_LEADCHNL_RV_SENSING_INTR_AMPL_MAX^MDC¦¦4.9¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
77 | NM | 722054 | MDC_IDC_MSMT_LEADCHNL_RV_SENSING_INTR_AMPL_MIN | MDC |  | 4.0 |  | mV |  | F |  |  |  |  | OBX¦77¦NM¦722054^MDC_IDC_MSMT_LEADCHNL_RV_SENSING_INTR_AMPL_MIN^MDC¦¦4.0¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
78 | NM | 722055 | MDC_IDC_MSMT_LEADCHNL_RV_SENSING_INTR_AMPL_MEAN | MDC |  | 4.7 |  | mV |  | F |  |  |  |  | OBX¦78¦NM¦722055^MDC_IDC_MSMT_LEADCHNL_RV_SENSING_INTR_AMPL_MEAN^MDC¦¦4.7¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
79 | CWE | 722113 | MDC_IDC_MSMT_LEADCHNL_RV_SENSING_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦79¦CWE¦722113^MDC_IDC_MSMT_LEADCHNL_RV_SENSING_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
80 | NM | 722177 | MDC_IDC_MSMT_LEADCHNL_RV_PACING_THRESHOLD_AMPLITUDE | MDC |  | 0.6 |  | V |  | F |  |  |  |  | OBX¦80¦NM¦722177^MDC_IDC_MSMT_LEADCHNL_RV_PACING_THRESHOLD_AMPLITUDE^MDC¦¦0.6¦V¦¦¦¦¦F¦¦¦¦¦¦¦
81 | NM | 722241 | MDC_IDC_MSMT_LEADCHNL_RV_PACING_THRESHOLD_PULSEWIDTH | MDC |  | 0.5 |  | ms |  | F |  |  |  |  | OBX¦81¦NM¦722241^MDC_IDC_MSMT_LEADCHNL_RV_PACING_THRESHOLD_PULSEWIDTH^MDC¦¦0.5¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
82 | CWE | 722305 | MDC_IDC_MSMT_LEADCHNL_RV_PACING_THRESHOLD_MEASUREMENT_METHOD | MDC |  | DeviceAutomatic | ^Device Automatic |  |  | F |  |  |  |  | OBX¦82¦CWE¦722305^MDC_IDC_MSMT_LEADCHNL_RV_PACING_THRESHOLD_MEASUREMENT_METHOD^MDC¦¦DeviceAutomatic^Device Automatic¦¦¦¦¦¦F¦¦¦¦¦¦¦
83 | CWE | 722369 | MDC_IDC_MSMT_LEADCHNL_RV_PACING_THRESHOLD_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦83¦CWE¦722369^MDC_IDC_MSMT_LEADCHNL_RV_PACING_THRESHOLD_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
84 | NM | 722433 | MDC_IDC_MSMT_LEADCHNL_RV_IMPEDANCE_VALUE | MDC |  | 3000 |  | Ohm |  | F |  |  |  |  | OBX¦84¦NM¦722433^MDC_IDC_MSMT_LEADCHNL_RV_IMPEDANCE_VALUE^MDC¦¦3000¦Ohm¦¦¦¦¦F¦¦¦¦¦¦¦
85 | CWE | 722497 | MDC_IDC_MSMT_LEADCHNL_RV_IMPEDANCE_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦85¦CWE¦722497^MDC_IDC_MSMT_LEADCHNL_RV_IMPEDANCE_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
86 | DTM | 721932 | MDC_IDC_MSMT_LEADCHNL_LV_DTM | MDC |  | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦86¦DTM¦721932^MDC_IDC_MSMT_LEADCHNL_LV_DTM^MDC¦¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
87 | DTM | 721933 | MDC_IDC_MSMT_LEADCHNL_LV_DTM_START | MDC |  | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦87¦DTM¦721933^MDC_IDC_MSMT_LEADCHNL_LV_DTM_START^MDC¦¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
88 | DTM | 721934 | MDC_IDC_MSMT_LEADCHNL_LV_DTM_END | MDC |  | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦88¦DTM¦721934^MDC_IDC_MSMT_LEADCHNL_LV_DTM_END^MDC¦¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
89 | CWE | 721987 | MDC_IDC_MSMT_LEADCHNL_LV_LEAD_CHANNEL_STATUS | MDC |  | NULL | ^Null |  |  | F |  |  |  |  | OBX¦89¦CWE¦721987^MDC_IDC_MSMT_LEADCHNL_LV_LEAD_CHANNEL_STATUS^MDC¦¦NULL^Null¦¦¦¦¦¦F¦¦¦¦¦¦¦
90 | NM | 722060 | MDC_IDC_MSMT_LEADCHNL_LV_SENSING_INTR_AMPL | MDC |  | 3.5 |  | mV |  | F |  |  |  |  | OBX¦90¦NM¦722060^MDC_IDC_MSMT_LEADCHNL_LV_SENSING_INTR_AMPL^MDC¦¦3.5¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
91 | NM | 722061 | MDC_IDC_MSMT_LEADCHNL_LV_SENSING_INTR_AMPL_MAX | MDC |  | 3.8 |  | mV |  | F |  |  |  |  | OBX¦91¦NM¦722061^MDC_IDC_MSMT_LEADCHNL_LV_SENSING_INTR_AMPL_MAX^MDC¦¦3.8¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
92 | NM | 722062 | MDC_IDC_MSMT_LEADCHNL_LV_SENSING_INTR_AMPL_MIN | MDC |  | 2.2 |  | mV |  | F |  |  |  |  | OBX¦92¦NM¦722062^MDC_IDC_MSMT_LEADCHNL_LV_SENSING_INTR_AMPL_MIN^MDC¦¦2.2¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
93 | NM | 722063 | MDC_IDC_MSMT_LEADCHNL_LV_SENSING_INTR_AMPL_MEAN | MDC |  | 3.5 |  | mV |  | F |  |  |  |  | OBX¦93¦NM¦722063^MDC_IDC_MSMT_LEADCHNL_LV_SENSING_INTR_AMPL_MEAN^MDC¦¦3.5¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
94 | CWE | 722115 | MDC_IDC_MSMT_LEADCHNL_LV_SENSING_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦94¦CWE¦722115^MDC_IDC_MSMT_LEADCHNL_LV_SENSING_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
95 | NM | 722179 | MDC_IDC_MSMT_LEADCHNL_LV_PACING_THRESHOLD_AMPLITUDE | MDC |  | 0.6 |  | V |  | F |  |  |  |  | OBX¦95¦NM¦722179^MDC_IDC_MSMT_LEADCHNL_LV_PACING_THRESHOLD_AMPLITUDE^MDC¦¦0.6¦V¦¦¦¦¦F¦¦¦¦¦¦¦
96 | NM | 722243 | MDC_IDC_MSMT_LEADCHNL_LV_PACING_THRESHOLD_PULSEWIDTH | MDC |  | 0.5 |  | ms |  | F |  |  |  |  | OBX¦96¦NM¦722243^MDC_IDC_MSMT_LEADCHNL_LV_PACING_THRESHOLD_PULSEWIDTH^MDC¦¦0.5¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
97 | CWE | 722307 | MDC_IDC_MSMT_LEADCHNL_LV_PACING_THRESHOLD_MEASUREMENT_METHOD | MDC |  | ProgrammerAutomatic | ^Programmer Automatic |  |  | F |  |  |  |  | OBX¦97¦CWE¦722307^MDC_IDC_MSMT_LEADCHNL_LV_PACING_THRESHOLD_MEASUREMENT_METHOD^MDC¦¦ProgrammerAutomatic^Programmer Automatic¦¦¦¦¦¦F¦¦¦¦¦¦¦
98 | CWE | 722371 | MDC_IDC_MSMT_LEADCHNL_LV_PACING_THRESHOLD_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦98¦CWE¦722371^MDC_IDC_MSMT_LEADCHNL_LV_PACING_THRESHOLD_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
99 | NM | 722435 | MDC_IDC_MSMT_LEADCHNL_LV_IMPEDANCE_VALUE | MDC |  | 500 |  | Ohm |  | F |  |  |  |  | OBX¦99¦NM¦722435^MDC_IDC_MSMT_LEADCHNL_LV_IMPEDANCE_VALUE^MDC¦¦500¦Ohm¦¦¦¦¦F¦¦¦¦¦¦¦
100 | CWE | 722499 | MDC_IDC_MSMT_LEADCHNL_LV_IMPEDANCE_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦100¦CWE¦722499^MDC_IDC_MSMT_LEADCHNL_LV_IMPEDANCE_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
101 | DTM | 722560 | MDC_IDC_MSMT_LEADHVCHNL_DTM | MDC | 1 | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦101¦DTM¦722560^MDC_IDC_MSMT_LEADHVCHNL_DTM^MDC¦1¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
102 | DTM | 722561 | MDC_IDC_MSMT_LEADHVCHNL_DTM_START | MDC | 1 | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦102¦DTM¦722561^MDC_IDC_MSMT_LEADHVCHNL_DTM_START^MDC¦1¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
103 | DTM | 722562 | MDC_IDC_MSMT_LEADHVCHNL_DTM_END | MDC | 1 | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦103¦DTM¦722562^MDC_IDC_MSMT_LEADHVCHNL_DTM_END^MDC¦1¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
104 | NM | 722624 | MDC_IDC_MSMT_LEADHVCHNL_IMPEDANCE | MDC | 1 | 330 |  | Ohm |  | F |  |  |  |  | OBX¦104¦NM¦722624^MDC_IDC_MSMT_LEADHVCHNL_IMPEDANCE^MDC¦1¦330¦Ohm¦¦¦¦¦F¦¦¦¦¦¦¦
105 | CWE | 722688 | MDC_IDC_MSMT_LEADHVCHNL_MEASUREMENT_TYPE | MDC | 1 | LowVoltage | ^Low Voltage Pulse |  |  | F |  |  |  |  | OBX¦105¦CWE¦722688^MDC_IDC_MSMT_LEADHVCHNL_MEASUREMENT_TYPE^MDC¦1¦LowVoltage^Low Voltage Pulse¦¦¦¦¦¦F¦¦¦¦¦¦¦
106 | CWE | 722752 | MDC_IDC_MSMT_LEADHVCHNL_STATUS | MDC | 1 | CheckLead | ^Check Lead |  |  | F |  |  |  |  | OBX¦106¦CWE¦722752^MDC_IDC_MSMT_LEADHVCHNL_STATUS^MDC¦1¦CheckLead^Check Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
107 | NM | 729344 | MDC_IDC_SET_CRT_LVRV_DELAY | MDC |  | -20 |  | ms |  | F |  |  |  |  | OBX¦107¦NM¦729344^MDC_IDC_SET_CRT_LVRV_DELAY^MDC¦¦-20¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
108 | CWE | 729408 | MDC_IDC_SET_CRT_PACED_CHAMBERS | MDC |  | BiV | ^BiV |  |  | F |  |  |  |  | OBX¦108¦CWE¦729408^MDC_IDC_SET_CRT_PACED_CHAMBERS^MDC¦¦BiV^BiV¦¦¦¦¦¦F¦¦¦¦¦¦¦
109 | ST | 729472 | MDC_IDC_SET_MAGNET_RESP | MDC |  | Therapies inactive, VVIR |  |  |  | F |  |  |  |  | OBX¦109¦ST¦729472^MDC_IDC_SET_MAGNET_RESP^MDC¦¦Therapies inactive, VVIR¦¦¦¦¦¦F¦¦¦¦¦¦¦
110 | NM | 729536 | MDC_IDC_SET_LEADCHNL_RA_SENSING_SENSITIVITY | MDC |  | 0.8 |  | mV |  | F |  |  |  |  | OBX¦110¦NM¦729536^MDC_IDC_SET_LEADCHNL_RA_SENSING_SENSITIVITY^MDC¦¦0.8¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
111 | CWE | 729600 | MDC_IDC_SET_LEADCHNL_RA_SENSING_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦111¦CWE¦729600^MDC_IDC_SET_LEADCHNL_RA_SENSING_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
112 | CWE | 729665 | MDC_IDC_SET_LEADCHNL_RA_SENSING_ANODE_LOCATION_1 | MDC |  | RA | ^Right Atrium |  |  | F |  |  |  |  | OBX¦112¦CWE¦729665^MDC_IDC_SET_LEADCHNL_RA_SENSING_ANODE_LOCATION_1^MDC¦¦RA^Right Atrium¦¦¦¦¦¦F¦¦¦¦¦¦¦
113 | CWE | 729729 | MDC_IDC_SET_LEADCHNL_RA_SENSING_ANODE_ELECTRODE_1 | MDC |  | Tip | ^Tip |  |  | F |  |  |  |  | OBX¦113¦CWE¦729729^MDC_IDC_SET_LEADCHNL_RA_SENSING_ANODE_ELECTRODE_1^MDC¦¦Tip^Tip¦¦¦¦¦¦F¦¦¦¦¦¦¦
114 | CWE | 729793 | MDC_IDC_SET_LEADCHNL_RA_SENSING_CATHODE_LOCATION_1 | MDC |  | Other | ^Other |  |  | F |  |  |  |  | OBX¦114¦CWE¦729793^MDC_IDC_SET_LEADCHNL_RA_SENSING_CATHODE_LOCATION_1^MDC¦¦Other^Other¦¦¦¦¦¦F¦¦¦¦¦¦¦
115 | CWE | 729857 | MDC_IDC_SET_LEADCHNL_RA_SENSING_CATHODE_ELECTRODE_1 | MDC |  | Can | ^Can |  |  | F |  |  |  |  | OBX¦115¦CWE¦729857^MDC_IDC_SET_LEADCHNL_RA_SENSING_CATHODE_ELECTRODE_1^MDC¦¦Can^Can¦¦¦¦¦¦F¦¦¦¦¦¦¦
116 | CWE | 729920 | MDC_IDC_SET_LEADCHNL_RA_SENSING_ADAPTATION_MODE | MDC |  | AdaptiveSensing | ^Adaptive |  |  | F |  |  |  |  | OBX¦116¦CWE¦729920^MDC_IDC_SET_LEADCHNL_RA_SENSING_ADAPTATION_MODE^MDC¦¦AdaptiveSensing^Adaptive¦¦¦¦¦¦F¦¦¦¦¦¦¦
117 | NM | 729984 | MDC_IDC_SET_LEADCHNL_RA_PACING_AMPLITUDE | MDC |  | 1.8 |  | V |  | F |  |  |  |  | OBX¦117¦NM¦729984^MDC_IDC_SET_LEADCHNL_RA_PACING_AMPLITUDE^MDC¦¦1.8¦V¦¦¦¦¦F¦¦¦¦¦¦¦
118 | NM | 730048 | MDC_IDC_SET_LEADCHNL_RA_PACING_PULSEWIDTH | MDC |  | 0.5 |  | ms |  | F |  |  |  |  | OBX¦118¦NM¦730048^MDC_IDC_SET_LEADCHNL_RA_PACING_PULSEWIDTH^MDC¦¦0.5¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
119 | CWE | 730112 | MDC_IDC_SET_LEADCHNL_RA_PACING_POLARITY | MDC |  | UNI | ^Unipolar Lead |  |  | F |  |  |  |  | OBX¦119¦CWE¦730112^MDC_IDC_SET_LEADCHNL_RA_PACING_POLARITY^MDC¦¦UNI^Unipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
120 | CWE | 730177 | MDC_IDC_SET_LEADCHNL_RA_PACING_ANODE_LOCATION_1 | MDC |  | RA | ^Right Atrium |  |  | F |  |  |  |  | OBX¦120¦CWE¦730177^MDC_IDC_SET_LEADCHNL_RA_PACING_ANODE_LOCATION_1^MDC¦¦RA^Right Atrium¦¦¦¦¦¦F¦¦¦¦¦¦¦
121 | CWE | 730241 | MDC_IDC_SET_LEADCHNL_RA_PACING_ANODE_ELECTRODE_1 | MDC |  | Ring | ^Ring |  |  | F |  |  |  |  | OBX¦121¦CWE¦730241^MDC_IDC_SET_LEADCHNL_RA_PACING_ANODE_ELECTRODE_1^MDC¦¦Ring^Ring¦¦¦¦¦¦F¦¦¦¦¦¦¦
122 | CWE | 730305 | MDC_IDC_SET_LEADCHNL_RA_PACING_CATHODE_LOCATION_1 | MDC |  | Other | ^Other |  |  | F |  |  |  |  | OBX¦122¦CWE¦730305^MDC_IDC_SET_LEADCHNL_RA_PACING_CATHODE_LOCATION_1^MDC¦¦Other^Other¦¦¦¦¦¦F¦¦¦¦¦¦¦
123 | CWE | 730369 | MDC_IDC_SET_LEADCHNL_RA_PACING_CATHODE_ELECTRODE_1 | MDC |  | Can | ^Can |  |  | F |  |  |  |  | OBX¦123¦CWE¦730369^MDC_IDC_SET_LEADCHNL_RA_PACING_CATHODE_ELECTRODE_1^MDC¦¦Can^Can¦¦¦¦¦¦F¦¦¦¦¦¦¦
124 | CWE | 730432 | MDC_IDC_SET_LEADCHNL_RA_PACING_CAPTURE_MODE | MDC |  | FixedPacing | ^Fixed Pacing |  |  | F |  |  |  |  | OBX¦124¦CWE¦730432^MDC_IDC_SET_LEADCHNL_RA_PACING_CAPTURE_MODE^MDC¦¦FixedPacing^Fixed Pacing¦¦¦¦¦¦F¦¦¦¦¦¦¦
125 | NM | 729537 | MDC_IDC_SET_LEADCHNL_RV_SENSING_SENSITIVITY | MDC |  | 1.3 |  | mV |  | F |  |  |  |  | OBX¦125¦NM¦729537^MDC_IDC_SET_LEADCHNL_RV_SENSING_SENSITIVITY^MDC¦¦1.3¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
126 | CWE | 729601 | MDC_IDC_SET_LEADCHNL_RV_SENSING_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦126¦CWE¦729601^MDC_IDC_SET_LEADCHNL_RV_SENSING_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
127 | CWE | 729669 | MDC_IDC_SET_LEADCHNL_RV_SENSING_ANODE_LOCATION_1 | MDC |  | RV | ^Right Ventricle |  |  | F |  |  |  |  | OBX¦127¦CWE¦729669^MDC_IDC_SET_LEADCHNL_RV_SENSING_ANODE_LOCATION_1^MDC¦¦RV^Right Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
128 | CWE | 729733 | MDC_IDC_SET_LEADCHNL_RV_SENSING_ANODE_ELECTRODE_1 | MDC |  | Tip | ^Tip |  |  | F |  |  |  |  | OBX¦128¦CWE¦729733^MDC_IDC_SET_LEADCHNL_RV_SENSING_ANODE_ELECTRODE_1^MDC¦¦Tip^Tip¦¦¦¦¦¦F¦¦¦¦¦¦¦
129 | CWE | 729797 | MDC_IDC_SET_LEADCHNL_RV_SENSING_CATHODE_LOCATION_1 | MDC |  | RV | ^Right Ventricle |  |  | F |  |  |  |  | OBX¦129¦CWE¦729797^MDC_IDC_SET_LEADCHNL_RV_SENSING_CATHODE_LOCATION_1^MDC¦¦RV^Right Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
130 | CWE | 729861 | MDC_IDC_SET_LEADCHNL_RV_SENSING_CATHODE_ELECTRODE_1 | MDC |  | Ring | ^Ring |  |  | F |  |  |  |  | OBX¦130¦CWE¦729861^MDC_IDC_SET_LEADCHNL_RV_SENSING_CATHODE_ELECTRODE_1^MDC¦¦Ring^Ring¦¦¦¦¦¦F¦¦¦¦¦¦¦
131 | CWE | 729921 | MDC_IDC_SET_LEADCHNL_RV_SENSING_ADAPTATION_MODE | MDC |  | AdaptiveSensing | ^Adaptive |  |  | F |  |  |  |  | OBX¦131¦CWE¦729921^MDC_IDC_SET_LEADCHNL_RV_SENSING_ADAPTATION_MODE^MDC¦¦AdaptiveSensing^Adaptive¦¦¦¦¦¦F¦¦¦¦¦¦¦
132 | NM | 729985 | MDC_IDC_SET_LEADCHNL_RV_PACING_AMPLITUDE | MDC |  | 2.0 |  | V |  | F |  |  |  |  | OBX¦132¦NM¦729985^MDC_IDC_SET_LEADCHNL_RV_PACING_AMPLITUDE^MDC¦¦2.0¦V¦¦¦¦¦F¦¦¦¦¦¦¦
133 | NM | 730049 | MDC_IDC_SET_LEADCHNL_RV_PACING_PULSEWIDTH | MDC |  | 0.5 |  | ms |  | F |  |  |  |  | OBX¦133¦NM¦730049^MDC_IDC_SET_LEADCHNL_RV_PACING_PULSEWIDTH^MDC¦¦0.5¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
134 | CWE | 730113 | MDC_IDC_SET_LEADCHNL_RV_PACING_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦134¦CWE¦730113^MDC_IDC_SET_LEADCHNL_RV_PACING_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
135 | CWE | 730181 | MDC_IDC_SET_LEADCHNL_RV_PACING_ANODE_LOCATION_1 | MDC |  | RV | ^Right Ventricle |  |  | F |  |  |  |  | OBX¦135¦CWE¦730181^MDC_IDC_SET_LEADCHNL_RV_PACING_ANODE_LOCATION_1^MDC¦¦RV^Right Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
136 | CWE | 730245 | MDC_IDC_SET_LEADCHNL_RV_PACING_ANODE_ELECTRODE_1 | MDC |  | Tip | ^Tip |  |  | F |  |  |  |  | OBX¦136¦CWE¦730245^MDC_IDC_SET_LEADCHNL_RV_PACING_ANODE_ELECTRODE_1^MDC¦¦Tip^Tip¦¦¦¦¦¦F¦¦¦¦¦¦¦
137 | CWE | 730309 | MDC_IDC_SET_LEADCHNL_RV_PACING_CATHODE_LOCATION_1 | MDC |  | RV | ^Right Ventricle |  |  | F |  |  |  |  | OBX¦137¦CWE¦730309^MDC_IDC_SET_LEADCHNL_RV_PACING_CATHODE_LOCATION_1^MDC¦¦RV^Right Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
138 | CWE | 730373 | MDC_IDC_SET_LEADCHNL_RV_PACING_CATHODE_ELECTRODE_1 | MDC |  | Ring | ^Ring |  |  | F |  |  |  |  | OBX¦138¦CWE¦730373^MDC_IDC_SET_LEADCHNL_RV_PACING_CATHODE_ELECTRODE_1^MDC¦¦Ring^Ring¦¦¦¦¦¦F¦¦¦¦¦¦¦
139 | CWE | 730433 | MDC_IDC_SET_LEADCHNL_RV_PACING_CAPTURE_MODE | MDC |  | AdaptiveCapture | ^Adaptive |  |  | F |  |  |  |  | OBX¦139¦CWE¦730433^MDC_IDC_SET_LEADCHNL_RV_PACING_CAPTURE_MODE^MDC¦¦AdaptiveCapture^Adaptive¦¦¦¦¦¦F¦¦¦¦¦¦¦
140 | NM | 729539 | MDC_IDC_SET_LEADCHNL_LV_SENSING_SENSITIVITY | MDC |  | 1.0 |  | mV |  | F |  |  |  |  | OBX¦140¦NM¦729539^MDC_IDC_SET_LEADCHNL_LV_SENSING_SENSITIVITY^MDC¦¦1.0¦mV¦¦¦¦¦F¦¦¦¦¦¦¦
141 | CWE | 729603 | MDC_IDC_SET_LEADCHNL_LV_SENSING_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦141¦CWE¦729603^MDC_IDC_SET_LEADCHNL_LV_SENSING_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
142 | CWE | 729677 | MDC_IDC_SET_LEADCHNL_LV_SENSING_ANODE_LOCATION_1 | MDC |  | LV | ^Left Ventricle |  |  | F |  |  |  |  | OBX¦142¦CWE¦729677^MDC_IDC_SET_LEADCHNL_LV_SENSING_ANODE_LOCATION_1^MDC¦¦LV^Left Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
143 | CWE | 729741 | MDC_IDC_SET_LEADCHNL_LV_SENSING_ANODE_ELECTRODE_1 | MDC |  | Tip | ^Tip |  |  | F |  |  |  |  | OBX¦143¦CWE¦729741^MDC_IDC_SET_LEADCHNL_LV_SENSING_ANODE_ELECTRODE_1^MDC¦¦Tip^Tip¦¦¦¦¦¦F¦¦¦¦¦¦¦
144 | CWE | 729805 | MDC_IDC_SET_LEADCHNL_LV_SENSING_CATHODE_LOCATION_1 | MDC |  | LV | ^Left Ventricle |  |  | F |  |  |  |  | OBX¦144¦CWE¦729805^MDC_IDC_SET_LEADCHNL_LV_SENSING_CATHODE_LOCATION_1^MDC¦¦LV^Left Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
145 | CWE | 729869 | MDC_IDC_SET_LEADCHNL_LV_SENSING_CATHODE_ELECTRODE_1 | MDC |  | Ring | ^Ring |  |  | F |  |  |  |  | OBX¦145¦CWE¦729869^MDC_IDC_SET_LEADCHNL_LV_SENSING_CATHODE_ELECTRODE_1^MDC¦¦Ring^Ring¦¦¦¦¦¦F¦¦¦¦¦¦¦
146 | CWE | 729923 | MDC_IDC_SET_LEADCHNL_LV_SENSING_ADAPTATION_MODE | MDC |  | FixedSensing | ^Fixed |  |  | F |  |  |  |  | OBX¦146¦CWE¦729923^MDC_IDC_SET_LEADCHNL_LV_SENSING_ADAPTATION_MODE^MDC¦¦FixedSensing^Fixed ¦¦¦¦¦¦F¦¦¦¦¦¦¦
147 | NM | 729987 | MDC_IDC_SET_LEADCHNL_LV_PACING_AMPLITUDE | MDC |  | 2.0 |  | V |  | F |  |  |  |  | OBX¦147¦NM¦729987^MDC_IDC_SET_LEADCHNL_LV_PACING_AMPLITUDE^MDC¦¦2.0¦V¦¦¦¦¦F¦¦¦¦¦¦¦
148 | NM | 730051 | MDC_IDC_SET_LEADCHNL_LV_PACING_PULSEWIDTH | MDC |  | 0.5 |  | ms |  | F |  |  |  |  | OBX¦148¦NM¦730051^MDC_IDC_SET_LEADCHNL_LV_PACING_PULSEWIDTH^MDC¦¦0.5¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
149 | CWE | 730115 | MDC_IDC_SET_LEADCHNL_LV_PACING_POLARITY | MDC |  | BI | ^Bipolar Lead |  |  | F |  |  |  |  | OBX¦149¦CWE¦730115^MDC_IDC_SET_LEADCHNL_LV_PACING_POLARITY^MDC¦¦BI^Bipolar Lead¦¦¦¦¦¦F¦¦¦¦¦¦¦
150 | CWE | 730189 | MDC_IDC_SET_LEADCHNL_LV_PACING_ANODE_LOCATION_1 | MDC |  | LV | ^Left Ventricle |  |  | F |  |  |  |  | OBX¦150¦CWE¦730189^MDC_IDC_SET_LEADCHNL_LV_PACING_ANODE_LOCATION_1^MDC¦¦LV^Left Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
151 | CWE | 730253 | MDC_IDC_SET_LEADCHNL_LV_PACING_ANODE_ELECTRODE_1 | MDC |  | Tip | ^Tip |  |  | F |  |  |  |  | OBX¦151¦CWE¦730253^MDC_IDC_SET_LEADCHNL_LV_PACING_ANODE_ELECTRODE_1^MDC¦¦Tip^Tip¦¦¦¦¦¦F¦¦¦¦¦¦¦
152 | CWE | 730317 | MDC_IDC_SET_LEADCHNL_LV_PACING_CATHODE_LOCATION_1 | MDC |  | RV | ^Right Ventricle |  |  | F |  |  |  |  | OBX¦152¦CWE¦730317^MDC_IDC_SET_LEADCHNL_LV_PACING_CATHODE_LOCATION_1^MDC¦¦RV^Right Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
153 | CWE | 730381 | MDC_IDC_SET_LEADCHNL_LV_PACING_CATHODE_ELECTRODE_1 | MDC |  | Ring | ^Ring |  |  | F |  |  |  |  | OBX¦153¦CWE¦730381^MDC_IDC_SET_LEADCHNL_LV_PACING_CATHODE_ELECTRODE_1^MDC¦¦Ring^Ring¦¦¦¦¦¦F¦¦¦¦¦¦¦
154 | CWE | 730435 | MDC_IDC_SET_LEADCHNL_LV_PACING_CAPTURE_MODE | MDC |  | AdaptiveCapture | ^Adaptive |  |  | F |  |  |  |  | OBX¦154¦CWE¦730435^MDC_IDC_SET_LEADCHNL_LV_PACING_CAPTURE_MODE^MDC¦¦AdaptiveCapture^Adaptive¦¦¦¦¦¦F¦¦¦¦¦¦¦
155 | CWE | 730497 | MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_LOCATION_1 | MDC |  | RV | ^Right Ventricle |  |  | F |  |  |  |  | OBX¦155¦CWE¦730497^MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_LOCATION_1^MDC¦¦RV^Right Ventricle¦¦¦¦¦¦F¦¦¦¦¦¦¦
156 | CWE | 730561 | MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_ELECTRODE_1 | MDC |  | Coil | ^Coil |  |  | F |  |  |  |  | OBX¦156¦CWE¦730561^MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_ELECTRODE_1^MDC¦¦Coil^Coil¦¦¦¦¦¦F¦¦¦¦¦¦¦
157 | CWE | 730498 | MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_LOCATION_2 | MDC |  | RA | ^Right Atrium |  |  | F |  |  |  |  | OBX¦157¦CWE¦730498^MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_LOCATION_2^MDC¦¦RA^Right Atrium¦¦¦¦¦¦F¦¦¦¦¦¦¦
158 | CWE | 730562 | MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_ELECTRODE_2 | MDC |  | Coil | ^Coil |  |  | F |  |  |  |  | OBX¦158¦CWE¦730562^MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_ANODE_ELECTRODE_2^MDC¦¦Coil^Coil¦¦¦¦¦¦F¦¦¦¦¦¦¦
159 | CWE | 730625 | MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_CATHODE_LOCATION_1 | MDC |  | Other | ^Other |  |  | F |  |  |  |  | OBX¦159¦CWE¦730625^MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_CATHODE_LOCATION_1^MDC¦¦Other^Other¦¦¦¦¦¦F¦¦¦¦¦¦¦
160 | CWE | 730689 | MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_CATHODE_ELECTRODE_1 | MDC |  | Can | ^Can |  |  | F |  |  |  |  | OBX¦160¦CWE¦730689^MDC_IDC_SET_LEADHVCHNL_SHOCK_VECTOR_CATHODE_ELECTRODE_1^MDC¦¦Can^Can¦¦¦¦¦¦F¦¦¦¦¦¦¦
161 | CWE | 730752 | MDC_IDC_SET_BRADY_MODE | MDC |  | DDDR | ^DDDR |  |  | F |  |  |  |  | OBX¦161¦CWE¦730752^MDC_IDC_SET_BRADY_MODE^MDC¦¦DDDR^DDDR¦¦¦¦¦¦F¦¦¦¦¦¦¦
162 | CWE | 730816 | MDC_IDC_SET_BRADY_VENDOR_MODE | MDC |  | DDDR | ^DDDR |  |  | F |  |  |  |  | OBX¦162¦CWE¦730816^MDC_IDC_SET_BRADY_VENDOR_MODE^MDC¦¦DDDR^DDDR¦¦¦¦¦¦F¦¦¦¦¦¦¦
163 | NM | 730880 | MDC_IDC_SET_BRADY_LOWRATE | MDC |  | 60 |  | {beats}/min |  | F |  |  |  |  | OBX¦163¦NM¦730880^MDC_IDC_SET_BRADY_LOWRATE^MDC¦¦60¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
164 | NM | 730944 | MDC_IDC_SET_BRADY_HYSTRATE | MDC |  | 55 |  | {beats}/min |  | F |  |  |  |  | OBX¦164¦NM¦730944^MDC_IDC_SET_BRADY_HYSTRATE^MDC¦¦55¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
165 | NM | 731008 | MDC_IDC_SET_BRADY_NIGHT_RATE | MDC |  | 50 |  | {beats}/min |  | F |  |  |  |  | OBX¦165¦NM¦731008^MDC_IDC_SET_BRADY_NIGHT_RATE^MDC¦¦50¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
166 | ST | 731072 | MDC_IDC_SET_BRADY_SENSOR_TYPE | MDC |  | Accelerometer |  |  |  | F |  |  |  |  | OBX¦166¦ST¦731072^MDC_IDC_SET_BRADY_SENSOR_TYPE^MDC¦¦Accelerometer¦¦¦¦¦¦F¦¦¦¦¦¦¦
167 | NM | 731136 | MDC_IDC_SET_BRADY_MAX_TRACKING_RATE | MDC |  | 130 |  | {beats}/min |  | F |  |  |  |  | OBX¦167¦NM¦731136^MDC_IDC_SET_BRADY_MAX_TRACKING_RATE^MDC¦¦130¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
168 | NM | 731200 | MDC_IDC_SET_BRADY_MAX_SENSOR_RATE | MDC |  | 120 |  | {beats}/min |  | F |  |  |  |  | OBX¦168¦NM¦731200^MDC_IDC_SET_BRADY_MAX_SENSOR_RATE^MDC¦¦120¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
169 | NM | 731266 | MDC_IDC_SET_BRADY_SAV_DELAY_LOW | MDC |  | 140 |  | ms |  | F |  |  |  |  | OBX¦169¦NM¦731266^MDC_IDC_SET_BRADY_SAV_DELAY_LOW^MDC¦¦140¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
170 | NM | 731265 | MDC_IDC_SET_BRADY_SAV_DELAY_HIGH | MDC |  | 180 |  | ms |  | F |  |  |  |  | OBX¦170¦NM¦731265^MDC_IDC_SET_BRADY_SAV_DELAY_HIGH^MDC¦¦180¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
171 | NM | 731330 | MDC_IDC_SET_BRADY_PAV_DELAY_LOW | MDC |  | 110 |  | ms |  | F |  |  |  |  | OBX¦171¦NM¦731330^MDC_IDC_SET_BRADY_PAV_DELAY_LOW^MDC¦¦110¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
172 | NM | 731329 | MDC_IDC_SET_BRADY_PAV_DELAY_HIGH | MDC |  | 150 |  | ms |  | F |  |  |  |  | OBX¦172¦NM¦731329^MDC_IDC_SET_BRADY_PAV_DELAY_HIGH^MDC¦¦150¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
173 | CWE | 731392 | MDC_IDC_SET_BRADY_AT_MODE_SWITCH_MODE | MDC |  | DDIR | ^DDIR |  |  | F |  |  |  |  | OBX¦173¦CWE¦731392^MDC_IDC_SET_BRADY_AT_MODE_SWITCH_MODE^MDC¦¦DDIR^DDIR¦¦¦¦¦¦F¦¦¦¦¦¦¦
174 | NM | 731456 | MDC_IDC_SET_BRADY_AT_MODE_SWITCH_RATE | MDC |  | 180 |  | {beats}/min |  | F |  |  |  |  | OBX¦174¦NM¦731456^MDC_IDC_SET_BRADY_AT_MODE_SWITCH_RATE^MDC¦¦180¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
175 | CWE | 731520 | MDC_IDC_SET_TACHYTHERAPY_VSTAT | MDC |  | On | ^On |  |  | F |  |  |  |  | OBX¦175¦CWE¦731520^MDC_IDC_SET_TACHYTHERAPY_VSTAT^MDC¦¦On^On¦¦¦¦¦¦F¦¦¦¦¦¦¦
176 | CWE | 731584 | MDC_IDC_SET_TACHYTHERAPY_ASTAT | MDC |  | Off | ^Off |  |  | F |  |  |  |  | OBX¦176¦CWE¦731584^MDC_IDC_SET_TACHYTHERAPY_ASTAT^MDC¦¦Off^Off¦¦¦¦¦¦F¦¦¦¦¦¦¦
177 | CWE | 731648 | MDC_IDC_SET_ZONE_TYPE | MDC | 1 | Zone_VF | ^VF |  |  | F |  |  |  |  | OBX¦177¦CWE¦731648^MDC_IDC_SET_ZONE_TYPE^MDC¦1¦Zone_VF^VF¦¦¦¦¦¦F¦¦¦¦¦¦¦
178 | CWE | 731712 | MDC_IDC_SET_ZONE_VENDOR_TYPE | MDC | 1 | VF | ^VF |  |  | F |  |  |  |  | OBX¦178¦CWE¦731712^MDC_IDC_SET_ZONE_VENDOR_TYPE^MDC¦1¦VF^VF¦¦¦¦¦¦F¦¦¦¦¦¦¦
179 | CWE | 731776 | MDC_IDC_SET_ZONE_STATUS | MDC | 1 | Active | ^Active |  |  | F |  |  |  |  | OBX¦179¦CWE¦731776^MDC_IDC_SET_ZONE_STATUS^MDC¦1¦Active^Active¦¦¦¦¦¦F¦¦¦¦¦¦¦
180 | NM | 731840 | MDC_IDC_SET_ZONE_DETECTION_INTERVAL | MDC | 1 | 195 |  | ms |  | F |  |  |  |  | OBX¦180¦NM¦731840^MDC_IDC_SET_ZONE_DETECTION_INTERVAL^MDC¦1¦195¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
181 | NM | 731904 | MDC_IDC_SET_ZONE_DETECTION_BEATS_NUMERATOR | MDC | 1 | 12 |  | {beats} |  | F |  |  |  |  | OBX¦181¦NM¦731904^MDC_IDC_SET_ZONE_DETECTION_BEATS_NUMERATOR^MDC¦1¦12¦{beats}¦¦¦¦¦F¦¦¦¦¦¦¦
182 | NM | 731968 | MDC_IDC_SET_ZONE_DETECTION_BEATS_DENOMINATOR | MDC | 1 | 18 |  | {beats} |  | F |  |  |  |  | OBX¦182¦NM¦731968^MDC_IDC_SET_ZONE_DETECTION_BEATS_DENOMINATOR^MDC¦1¦18¦{beats}¦¦¦¦¦F¦¦¦¦¦¦¦
183 | ST | 732032 | MDC_IDC_SET_ZONE_DETECTION_DETAILS | MDC | 1 |  |  |  |  | F |  |  |  |  | OBX¦183¦ST¦732032^MDC_IDC_SET_ZONE_DETECTION_DETAILS^MDC¦1¦¦¦¦¦¦¦F¦¦¦¦¦¦¦
184 | CWE | 732097 | MDC_IDC_SET_ZONE_TYPE_ATP_1 | MDC | 1 | Ramp | ^Ramp |  |  | F |  |  |  |  | OBX¦184¦CWE¦732097^MDC_IDC_SET_ZONE_TYPE_ATP_1^MDC¦1¦Ramp^Ramp¦¦¦¦¦¦F¦¦¦¦¦¦¦
185 | NM | 732161 | MDC_IDC_SET_ZONE_NUM_ATP_SEQS_1 | MDC | 1 | 1 |  | {seq} |  | F |  |  |  |  | OBX¦185¦NM¦732161^MDC_IDC_SET_ZONE_NUM_ATP_SEQS_1^MDC¦1¦1¦{seq}¦¦¦¦¦F¦¦¦¦¦¦¦
186 | NM | 732225 | MDC_IDC_SET_ZONE_SHOCK_ENERGY_1 | MDC | 1 | 30 |  | J |  | F |  |  |  |  | OBX¦186¦NM¦732225^MDC_IDC_SET_ZONE_SHOCK_ENERGY_1^MDC¦1¦30¦J¦¦¦¦¦F¦¦¦¦¦¦¦
187 | NM | 732289 | MDC_IDC_SET_ZONE_NUM_SHOCKS_1 | MDC | 1 | 5 |  | {shocks} |  | F |  |  |  |  | OBX¦187¦NM¦732289^MDC_IDC_SET_ZONE_NUM_SHOCKS_1^MDC¦1¦5¦{shocks}¦¦¦¦¦F¦¦¦¦¦¦¦
188 | CWE | 731648 | MDC_IDC_SET_ZONE_TYPE | MDC | 2 | Zone_VT | ^VT |  |  | F |  |  |  |  | OBX¦188¦CWE¦731648^MDC_IDC_SET_ZONE_TYPE^MDC¦2¦Zone_VT^VT¦¦¦¦¦¦F¦¦¦¦¦¦¦
189 | CWE | 731712 | MDC_IDC_SET_ZONE_VENDOR_TYPE | MDC | 2 | VT1 | ^VT1 |  |  | F |  |  |  |  | OBX¦189¦CWE¦731712^MDC_IDC_SET_ZONE_VENDOR_TYPE^MDC¦2¦VT1^VT1¦¦¦¦¦¦F¦¦¦¦¦¦¦
190 | CWE | 731776 | MDC_IDC_SET_ZONE_STATUS | MDC | 2 | Active | ^Active |  |  | F |  |  |  |  | OBX¦190¦CWE¦731776^MDC_IDC_SET_ZONE_STATUS^MDC¦2¦Active^Active¦¦¦¦¦¦F¦¦¦¦¦¦¦
191 | NM | 731840 | MDC_IDC_SET_ZONE_DETECTION_INTERVAL | MDC | 2 | 165 |  | ms |  | F |  |  |  |  | OBX¦191¦NM¦731840^MDC_IDC_SET_ZONE_DETECTION_INTERVAL^MDC¦2¦165¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
192 | NM | 731904 | MDC_IDC_SET_ZONE_DETECTION_BEATS_NUMERATOR | MDC | 2 | 9 |  | {beats} |  | F |  |  |  |  | OBX¦192¦NM¦731904^MDC_IDC_SET_ZONE_DETECTION_BEATS_NUMERATOR^MDC¦2¦9¦{beats}¦¦¦¦¦F¦¦¦¦¦¦¦
193 | NM | 731968 | MDC_IDC_SET_ZONE_DETECTION_BEATS_DENOMINATOR | MDC | 2 | 12 |  | {beats} |  | F |  |  |  |  | OBX¦193¦NM¦731968^MDC_IDC_SET_ZONE_DETECTION_BEATS_DENOMINATOR^MDC¦2¦12¦{beats}¦¦¦¦¦F¦¦¦¦¦¦¦
194 | ST | 732032 | MDC_IDC_SET_ZONE_DETECTION_DETAILS | MDC | 2 | Enhanced detection |  |  |  | F |  |  |  |  | OBX¦194¦ST¦732032^MDC_IDC_SET_ZONE_DETECTION_DETAILS^MDC¦2¦Enhanced detection¦¦¦¦¦¦F¦¦¦¦¦¦¦
195 | CWE | 732097 | MDC_IDC_SET_ZONE_TYPE_ATP_1 | MDC | 2 | Burst | ^Burst |  |  | F |  |  |  |  | OBX¦195¦CWE¦732097^MDC_IDC_SET_ZONE_TYPE_ATP_1^MDC¦2¦Burst^Burst¦¦¦¦¦¦F¦¦¦¦¦¦¦
196 | NM | 732161 | MDC_IDC_SET_ZONE_NUM_ATP_SEQS_1 | MDC | 2 | 5 |  | {seq} |  | F |  |  |  |  | OBX¦196¦NM¦732161^MDC_IDC_SET_ZONE_NUM_ATP_SEQS_1^MDC¦2¦5¦{seq}¦¦¦¦¦F¦¦¦¦¦¦¦
197 | NM | 732225 | MDC_IDC_SET_ZONE_SHOCK_ENERGY_1 | MDC | 2 | 20 |  | J |  | F |  |  |  |  | OBX¦197¦NM¦732225^MDC_IDC_SET_ZONE_SHOCK_ENERGY_1^MDC¦2¦20¦J¦¦¦¦¦F¦¦¦¦¦¦¦
198 | NM | 732289 | MDC_IDC_SET_ZONE_NUM_SHOCKS_1 | MDC | 2 | 1 |  | {shocks} |  | F |  |  |  |  | OBX¦198¦NM¦732289^MDC_IDC_SET_ZONE_NUM_SHOCKS_1^MDC¦2¦1¦{shocks}¦¦¦¦¦F¦¦¦¦¦¦¦
199 | NM | 732226 | MDC_IDC_SET_ZONE_SHOCK_ENERGY_2 | MDC | 2 | 30 |  | J |  | F |  |  |  |  | OBX¦199¦NM¦732226^MDC_IDC_SET_ZONE_SHOCK_ENERGY_2^MDC¦2¦30¦J¦¦¦¦¦F¦¦¦¦¦¦¦
200 | NM | 732290 | MDC_IDC_SET_ZONE_NUM_SHOCKS_2 | MDC | 2 | 1 |  | {shocks} |  | F |  |  |  |  | OBX¦200¦NM¦732290^MDC_IDC_SET_ZONE_NUM_SHOCKS_2^MDC¦2¦1¦{shocks}¦¦¦¦¦F¦¦¦¦¦¦¦
201 | NM | 732227 | MDC_IDC_SET_ZONE_SHOCK_ENERGY_3 | MDC | 2 | 30 |  | J |  | F |  |  |  |  | OBX¦201¦NM¦732227^MDC_IDC_SET_ZONE_SHOCK_ENERGY_3^MDC¦2¦30¦J¦¦¦¦¦F¦¦¦¦¦¦¦
202 | NM | 732291 | MDC_IDC_SET_ZONE_NUM_SHOCKS_3 | MDC | 2 | 5 |  | {shocks} |  | F |  |  |  |  | OBX¦202¦NM¦732291^MDC_IDC_SET_ZONE_NUM_SHOCKS_3^MDC¦2¦5¦{shocks}¦¦¦¦¦F¦¦¦¦¦¦¦
203 | CWE | 731648 | MDC_IDC_SET_ZONE_TYPE | MDC | 3 | Zone_VT | ^VT |  |  | F |  |  |  |  | OBX¦203¦CWE¦731648^MDC_IDC_SET_ZONE_TYPE^MDC¦3¦Zone_VT^VT¦¦¦¦¦¦F¦¦¦¦¦¦¦
204 | CWE | 731712 | MDC_IDC_SET_ZONE_VENDOR_TYPE | MDC | 3 | FastVT | ^FastVT |  |  | F |  |  |  |  | OBX¦204¦CWE¦731712^MDC_IDC_SET_ZONE_VENDOR_TYPE^MDC¦3¦FastVT^FastVT¦¦¦¦¦¦F¦¦¦¦¦¦¦
205 | CWE | 731776 | MDC_IDC_SET_ZONE_STATUS | MDC | 3 | Active | ^Active |  |  | F |  |  |  |  | OBX¦205¦CWE¦731776^MDC_IDC_SET_ZONE_STATUS^MDC¦3¦Active^Active¦¦¦¦¦¦F¦¦¦¦¦¦¦
206 | NM | 731840 | MDC_IDC_SET_ZONE_DETECTION_INTERVAL | MDC | 3 | 165 |  | ms |  | F |  |  |  |  | OBX¦206¦NM¦731840^MDC_IDC_SET_ZONE_DETECTION_INTERVAL^MDC¦3¦165¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
207 | NM | 731904 | MDC_IDC_SET_ZONE_DETECTION_BEATS_NUMERATOR | MDC | 3 | 9 |  | {beats} |  | F |  |  |  |  | OBX¦207¦NM¦731904^MDC_IDC_SET_ZONE_DETECTION_BEATS_NUMERATOR^MDC¦3¦9¦{beats}¦¦¦¦¦F¦¦¦¦¦¦¦
208 | NM | 731968 | MDC_IDC_SET_ZONE_DETECTION_BEATS_DENOMINATOR | MDC | 3 | 12 |  | {beats} |  | F |  |  |  |  | OBX¦208¦NM¦731968^MDC_IDC_SET_ZONE_DETECTION_BEATS_DENOMINATOR^MDC¦3¦12¦{beats}¦¦¦¦¦F¦¦¦¦¦¦¦
209 | ST | 732032 | MDC_IDC_SET_ZONE_DETECTION_DETAILS | MDC | 3 | Progressive Therapy |  |  |  | F |  |  |  |  | OBX¦209¦ST¦732032^MDC_IDC_SET_ZONE_DETECTION_DETAILS^MDC¦3¦Progressive Therapy¦¦¦¦¦¦F¦¦¦¦¦¦¦
210 | CWE | 732097 | MDC_IDC_SET_ZONE_TYPE_ATP_1 | MDC | 3 | RampScan | ^Ramp+Scan |  |  | F |  |  |  |  | OBX¦210¦CWE¦732097^MDC_IDC_SET_ZONE_TYPE_ATP_1^MDC¦3¦RampScan^Ramp+Scan¦¦¦¦¦¦F¦¦¦¦¦¦¦
211 | NM | 732161 | MDC_IDC_SET_ZONE_NUM_ATP_SEQS_1 | MDC | 3 | 5 |  | {seq} |  | F |  |  |  |  | OBX¦211¦NM¦732161^MDC_IDC_SET_ZONE_NUM_ATP_SEQS_1^MDC¦3¦5¦{seq}¦¦¦¦¦F¦¦¦¦¦¦¦
212 | NM | 732225 | MDC_IDC_SET_ZONE_SHOCK_ENERGY_1 | MDC | 3 | 20 |  | J |  | F |  |  |  |  | OBX¦212¦NM¦732225^MDC_IDC_SET_ZONE_SHOCK_ENERGY_1^MDC¦3¦20¦J¦¦¦¦¦F¦¦¦¦¦¦¦
213 | NM | 732289 | MDC_IDC_SET_ZONE_NUM_SHOCKS_1 | MDC | 3 | 1 |  | {shocks} |  | F |  |  |  |  | OBX¦213¦NM¦732289^MDC_IDC_SET_ZONE_NUM_SHOCKS_1^MDC¦3¦1¦{shocks}¦¦¦¦¦F¦¦¦¦¦¦¦
214 | NM | 732226 | MDC_IDC_SET_ZONE_SHOCK_ENERGY_2 | MDC | 3 | 30 |  | J |  | F |  |  |  |  | OBX¦214¦NM¦732226^MDC_IDC_SET_ZONE_SHOCK_ENERGY_2^MDC¦3¦30¦J¦¦¦¦¦F¦¦¦¦¦¦¦
215 | NM | 732290 | MDC_IDC_SET_ZONE_NUM_SHOCKS_2 | MDC | 3 | 5 |  | {shocks} |  | F |  |  |  |  | OBX¦215¦NM¦732290^MDC_IDC_SET_ZONE_NUM_SHOCKS_2^MDC¦3¦5¦{shocks}¦¦¦¦¦F¦¦¦¦¦¦¦
216 | DTM | 737617 | MDC_IDC_STAT_HEART_RATE_DTM_START | MDC |  | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦216¦DTM¦737617^MDC_IDC_STAT_HEART_RATE_DTM_START^MDC¦¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
217 | DTM | 737618 | MDC_IDC_STAT_HEART_RATE_DTM_END | MDC |  | 20090501030000 |  |  |  | F |  |  |  |  | OBX¦217¦DTM¦737618^MDC_IDC_STAT_HEART_RATE_DTM_END^MDC¦¦20090501030000¦¦¦¦¦¦F¦¦¦¦¦¦¦
218 | NM | 737635 | MDC_IDC_STAT_HEART_RATE_ATRIAL_MEAN | MDC |  | 72 |  | {beats}/min |  | F |  |  |  |  | OBX¦218¦NM¦737635^MDC_IDC_STAT_HEART_RATE_ATRIAL_MEAN^MDC¦¦72¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
219 | NM | 737633 | MDC_IDC_STAT_HEART_RATE_ATRIAL_MAX | MDC |  | 165 |  | {beats}/min |  | F |  |  |  |  | OBX¦219¦NM¦737633^MDC_IDC_STAT_HEART_RATE_ATRIAL_MAX^MDC¦¦165¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
220 | NM | 737634 | MDC_IDC_STAT_HEART_RATE_ATRIAL_MIN | MDC |  | 60 |  | {beats}/min |  | F |  |  |  |  | OBX¦220¦NM¦737634^MDC_IDC_STAT_HEART_RATE_ATRIAL_MIN^MDC¦¦60¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
221 | NM | 737651 | MDC_IDC_STAT_HEART_RATE_VENTRICULAR_MEAN | MDC |  | 72 |  | {beats}/min |  | F |  |  |  |  | OBX¦221¦NM¦737651^MDC_IDC_STAT_HEART_RATE_VENTRICULAR_MEAN^MDC¦¦72¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
222 | NM | 737649 | MDC_IDC_STAT_HEART_RATE_VENTRICULAR_MAX | MDC |  | 165 |  | {beats}/min |  | F |  |  |  |  | OBX¦222¦NM¦737649^MDC_IDC_STAT_HEART_RATE_VENTRICULAR_MAX^MDC¦¦165¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
223 | NM | 737650 | MDC_IDC_STAT_HEART_RATE_VENTRICULAR_MIN | MDC |  | 60 |  | {beats}/min |  | F |  |  |  |  | OBX¦223¦NM¦737650^MDC_IDC_STAT_HEART_RATE_VENTRICULAR_MIN^MDC¦¦60¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
224 | DTM | 737505 | MDC_IDC_STAT_BRADY_DTM_START | MDC |  | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦224¦DTM¦737505^MDC_IDC_STAT_BRADY_DTM_START^MDC¦¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
225 | DTM | 737506 | MDC_IDC_STAT_BRADY_DTM_END | MDC |  | 20090501030000 |  |  |  | F |  |  |  |  | OBX¦225¦DTM¦737506^MDC_IDC_STAT_BRADY_DTM_END^MDC¦¦20090501030000¦¦¦¦¦¦F¦¦¦¦¦¦¦
226 | NM | 737520 | MDC_IDC_STAT_BRADY_RA_PERCENT_PACED | MDC |  | 50 |  | % |  | F |  |  |  |  | OBX¦226¦NM¦737520^MDC_IDC_STAT_BRADY_RA_PERCENT_PACED^MDC¦¦50¦%¦¦¦¦¦F¦¦¦¦¦¦¦
227 | NM | 737536 | MDC_IDC_STAT_BRADY_RV_PERCENT_PACED | MDC |  | 30 |  | % |  | F |  |  |  |  | OBX¦227¦NM¦737536^MDC_IDC_STAT_BRADY_RV_PERCENT_PACED^MDC¦¦30¦%¦¦¦¦¦F¦¦¦¦¦¦¦
228 | NM | 737552 | MDC_IDC_STAT_BRADY_AP_VP_PERCENT | MDC |  | 10 |  | % |  | F |  |  |  |  | OBX¦228¦NM¦737552^MDC_IDC_STAT_BRADY_AP_VP_PERCENT^MDC¦¦10¦%¦¦¦¦¦F¦¦¦¦¦¦¦
229 | NM | 737568 | MDC_IDC_STAT_BRADY_AS_VP_PERCENT | MDC |  | 20 |  | % |  | F |  |  |  |  | OBX¦229¦NM¦737568^MDC_IDC_STAT_BRADY_AS_VP_PERCENT^MDC¦¦20¦%¦¦¦¦¦F¦¦¦¦¦¦¦
230 | NM | 737584 | MDC_IDC_STAT_BRADY_AP_VS_PERCENT | MDC |  | 40 |  | % |  | F |  |  |  |  | OBX¦230¦NM¦737584^MDC_IDC_STAT_BRADY_AP_VS_PERCENT^MDC¦¦40¦%¦¦¦¦¦F¦¦¦¦¦¦¦
231 | NM | 737600 | MDC_IDC_STAT_BRADY_AS_VS_PERCENT | MDC |  | 60 |  | % |  | F |  |  |  |  | OBX¦231¦NM¦737600^MDC_IDC_STAT_BRADY_AS_VS_PERCENT^MDC¦¦60¦%¦¦¦¦¦F¦¦¦¦¦¦¦
232 | DTM | 737665 | MDC_IDC_STAT_AT_DTM_START | MDC |  | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦232¦DTM¦737665^MDC_IDC_STAT_AT_DTM_START^MDC¦¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
233 | DTM | 737666 | MDC_IDC_STAT_AT_DTM_END | MDC |  | 20090501030000 |  |  |  | F |  |  |  |  | OBX¦233¦DTM¦737666^MDC_IDC_STAT_AT_DTM_END^MDC¦¦20090501030000¦¦¦¦¦¦F¦¦¦¦¦¦¦
234 | NM | 737696 | MDC_IDC_STAT_AT_BURDEN_PERCENT | MDC |  | 10 |  | % |  | F |  |  |  |  | OBX¦234¦NM¦737696^MDC_IDC_STAT_AT_BURDEN_PERCENT^MDC¦¦10¦%¦¦¦¦¦F¦¦¦¦¦¦¦
235 | NM | 737712 | MDC_IDC_STAT_AT_MODE_SW_PERCENT_TIME | MDC |  | 5 |  | % |  | F |  |  |  |  | OBX¦235¦NM¦737712^MDC_IDC_STAT_AT_MODE_SW_PERCENT_TIME^MDC¦¦5¦%¦¦¦¦¦F¦¦¦¦¦¦¦
236 | NM | 737728 | MDC_IDC_STAT_AT_MODE_SW_PERCENT_TIME_PER_DAY | MDC |  | 5 |  | % |  | F |  |  |  |  | OBX¦236¦NM¦737728^MDC_IDC_STAT_AT_MODE_SW_PERCENT_TIME_PER_DAY^MDC¦¦5¦%¦¦¦¦¦F¦¦¦¦¦¦¦
237 | NM | 737760 | MDC_IDC_STAT_AT_MODE_SW_COUNT_PER_DAY | MDC |  | 360 |  | {switches} |  | F |  |  |  |  | OBX¦237¦NM¦737760^MDC_IDC_STAT_AT_MODE_SW_COUNT_PER_DAY^MDC¦¦360¦{switches}¦¦¦¦¦F¦¦¦¦¦¦¦
238 | NM | 737744 | MDC_IDC_STAT_AT_MODE_SW_COUNT | MDC |  | 10800 |  | {switches} |  | F |  |  |  |  | OBX¦238¦NM¦737744^MDC_IDC_STAT_AT_MODE_SW_COUNT^MDC¦¦10800¦{switches}¦¦¦¦¦F¦¦¦¦¦¦¦
239 | NM | 737680 | MDC_IDC_STAT_AT_MODE_SW_MAX_DURATION | MDC |  | 60 |  | s |  | F |  |  |  |  | OBX¦239¦NM¦737680^MDC_IDC_STAT_AT_MODE_SW_MAX_DURATION^MDC¦¦60¦s¦¦¦¦¦F¦¦¦¦¦¦¦
240 | DTM | 737777 | MDC_IDC_STAT_CRT_DTM_START | MDC |  | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦240¦DTM¦737777^MDC_IDC_STAT_CRT_DTM_START^MDC¦¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
241 | DTM | 737778 | MDC_IDC_STAT_CRT_DTM_END | MDC |  | 20090501030000 |  |  |  | F |  |  |  |  | OBX¦241¦DTM¦737778^MDC_IDC_STAT_CRT_DTM_END^MDC¦¦20090501030000¦¦¦¦¦¦F¦¦¦¦¦¦¦
242 | NM | 737792 | MDC_IDC_STAT_CRT_LV_PERCENT_PACED | MDC |  | 95 |  | % |  | F |  |  |  |  | OBX¦242¦NM¦737792^MDC_IDC_STAT_CRT_LV_PERCENT_PACED^MDC¦¦95¦%¦¦¦¦¦F¦¦¦¦¦¦¦
243 | NM | 737808 | MDC_IDC_STAT_CRT_PERCENT_PACED | MDC |  | 80 |  | % |  | F |  |  |  |  | OBX¦243¦NM¦737808^MDC_IDC_STAT_CRT_PERCENT_PACED^MDC¦¦80¦%¦¦¦¦¦F¦¦¦¦¦¦¦
244 | NM | 737824 | MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_DELIVERED_RECENT | MDC |  | 1 |  | {shocks} |  | F |  |  |  |  | OBX¦244¦NM¦737824^MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_DELIVERED_RECENT^MDC¦¦1¦{shocks}¦¦¦¦¦F¦¦¦¦¦¦¦
245 | NM | 737840 | MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_DELIVERED_TOTAL | MDC |  | 4 |  | {shocks} |  | F |  |  |  |  | OBX¦245¦NM¦737840^MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_DELIVERED_TOTAL^MDC¦¦4¦{shocks}¦¦¦¦¦F¦¦¦¦¦¦¦
246 | NM | 737856 | MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_ABORTED_RECENT | MDC |  | 0 |  | {shocks} |  | F |  |  |  |  | OBX¦246¦NM¦737856^MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_ABORTED_RECENT^MDC¦¦0¦{shocks}¦¦¦¦¦F¦¦¦¦¦¦¦
247 | NM | 737872 | MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_ABORTED_TOTAL | MDC |  | 1 |  | {shocks} |  | F |  |  |  |  | OBX¦247¦NM¦737872^MDC_IDC_STAT_TACHYTHERAPY_SHOCKS_ABORTED_TOTAL^MDC¦¦1¦{shocks}¦¦¦¦¦F¦¦¦¦¦¦¦
248 | NM | 737888 | MDC_IDC_STAT_TACHYTHERAPY_ATP_DELIVERED_RECENT | MDC |  | 1 |  | {seq} |  | F |  |  |  |  | OBX¦248¦NM¦737888^MDC_IDC_STAT_TACHYTHERAPY_ATP_DELIVERED_RECENT^MDC¦¦1¦{seq}¦¦¦¦¦F¦¦¦¦¦¦¦
249 | NM | 737904 | MDC_IDC_STAT_TACHYTHERAPY_ATP_DELIVERED_TOTAL | MDC |  | 3 |  | {seq} |  | F |  |  |  |  | OBX¦249¦NM¦737904^MDC_IDC_STAT_TACHYTHERAPY_ATP_DELIVERED_TOTAL^MDC¦¦3¦{seq}¦¦¦¦¦F¦¦¦¦¦¦¦
250 | DTM | 737921 | MDC_IDC_STAT_TACHYTHERAPY_TOTAL_DTM_START | MDC |  | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦250¦DTM¦737921^MDC_IDC_STAT_TACHYTHERAPY_TOTAL_DTM_START^MDC¦¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
251 | DTM | 737922 | MDC_IDC_STAT_TACHYTHERAPY_TOTAL_DTM_END | MDC |  | 20090501030000 |  |  |  | F |  |  |  |  | OBX¦251¦DTM¦737922^MDC_IDC_STAT_TACHYTHERAPY_TOTAL_DTM_END^MDC¦¦20090501030000¦¦¦¦¦¦F¦¦¦¦¦¦¦
252 | DTM | 737937 | MDC_IDC_STAT_TACHYTHERAPY_RECENT_DTM_START | MDC |  | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦252¦DTM¦737937^MDC_IDC_STAT_TACHYTHERAPY_RECENT_DTM_START^MDC¦¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
253 | DTM | 737938 | MDC_IDC_STAT_TACHYTHERAPY_RECENT_DTM_END | MDC |  | 20090501030000 |  |  |  | F |  |  |  |  | OBX¦253¦DTM¦737938^MDC_IDC_STAT_TACHYTHERAPY_RECENT_DTM_END^MDC¦¦20090501030000¦¦¦¦¦¦F¦¦¦¦¦¦¦
254 | CWE | 737952 | MDC_IDC_STAT_EPISODE_TYPE | MDC | 1 | Epis_VF | ^VF |  |  | F |  |  |  |  | OBX¦254¦CWE¦737952^MDC_IDC_STAT_EPISODE_TYPE^MDC¦1¦Epis_VF^VF¦¦¦¦¦¦F¦¦¦¦¦¦¦
255 | CWE | 737968 | MDC_IDC_STAT_EPISODE_TYPE_INDUCED | MDC | 1 | NO | ^NO |  |  | F |  |  |  |  | OBX¦255¦CWE¦737968^MDC_IDC_STAT_EPISODE_TYPE_INDUCED^MDC¦1¦NO^NO¦¦¦¦¦¦F¦¦¦¦¦¦¦
256 | CWE | 737984 | MDC_IDC_STAT_EPISODE_VENDOR_TYPE | MDC | 1 | VF | ^VF |  |  | F |  |  |  |  | OBX¦256¦CWE¦737984^MDC_IDC_STAT_EPISODE_VENDOR_TYPE^MDC¦1¦VF^VF¦¦¦¦¦¦F¦¦¦¦¦¦¦
257 | NM | 738000 | MDC_IDC_STAT_EPISODE_RECENT_COUNT | MDC | 1 | 0 |  |  |  | F |  |  |  |  | OBX¦257¦NM¦738000^MDC_IDC_STAT_EPISODE_RECENT_COUNT^MDC¦1¦0¦¦¦¦¦¦F¦¦¦¦¦¦¦
258 | NM | 738032 | MDC_IDC_STAT_EPISODE_TOTAL_COUNT | MDC | 1 | 0 |  |  |  | F |  |  |  |  | OBX¦258¦NM¦738032^MDC_IDC_STAT_EPISODE_TOTAL_COUNT^MDC¦1¦0¦¦¦¦¦¦F¦¦¦¦¦¦¦
259 | CWE | 737952 | MDC_IDC_STAT_EPISODE_TYPE | MDC | 2 | Epis_VT | ^VT |  |  | F |  |  |  |  | OBX¦259¦CWE¦737952^MDC_IDC_STAT_EPISODE_TYPE^MDC¦2¦Epis_VT^VT¦¦¦¦¦¦F¦¦¦¦¦¦¦
260 | CWE | 737968 | MDC_IDC_STAT_EPISODE_TYPE_INDUCED | MDC | 2 | NO | ^NO |  |  | F |  |  |  |  | OBX¦260¦CWE¦737968^MDC_IDC_STAT_EPISODE_TYPE_INDUCED^MDC¦2¦NO^NO¦¦¦¦¦¦F¦¦¦¦¦¦¦
261 | CWE | 737984 | MDC_IDC_STAT_EPISODE_VENDOR_TYPE | MDC | 2 | VT1 | ^VT1 |  |  | F |  |  |  |  | OBX¦261¦CWE¦737984^MDC_IDC_STAT_EPISODE_VENDOR_TYPE^MDC¦2¦VT1^VT1¦¦¦¦¦¦F¦¦¦¦¦¦¦
262 | NM | 738000 | MDC_IDC_STAT_EPISODE_RECENT_COUNT | MDC | 2 | 1 |  |  |  | F |  |  |  |  | OBX¦262¦NM¦738000^MDC_IDC_STAT_EPISODE_RECENT_COUNT^MDC¦2¦1¦¦¦¦¦¦F¦¦¦¦¦¦¦
263 | NM | 738032 | MDC_IDC_STAT_EPISODE_TOTAL_COUNT | MDC | 2 | 1 |  |  |  | F |  |  |  |  | OBX¦263¦NM¦738032^MDC_IDC_STAT_EPISODE_TOTAL_COUNT^MDC¦2¦1¦¦¦¦¦¦F¦¦¦¦¦¦¦
264 | CWE | 737952 | MDC_IDC_STAT_EPISODE_TYPE | MDC | 3 | Epis_ATAF | ^AT/AF |  |  | F |  |  |  |  | OBX¦264¦CWE¦737952^MDC_IDC_STAT_EPISODE_TYPE^MDC¦3¦Epis_ATAF^AT/AF¦¦¦¦¦¦F¦¦¦¦¦¦¦
265 | CWE | 737968 | MDC_IDC_STAT_EPISODE_TYPE_INDUCED | MDC | 3 | NO | ^NO |  |  | F |  |  |  |  | OBX¦265¦CWE¦737968^MDC_IDC_STAT_EPISODE_TYPE_INDUCED^MDC¦3¦NO^NO¦¦¦¦¦¦F¦¦¦¦¦¦¦
266 | CWE | 737984 | MDC_IDC_STAT_EPISODE_VENDOR_TYPE | MDC | 3 | ATAF | ^ATAF |  |  | F |  |  |  |  | OBX¦266¦CWE¦737984^MDC_IDC_STAT_EPISODE_VENDOR_TYPE^MDC¦3¦ATAF^ATAF¦¦¦¦¦¦F¦¦¦¦¦¦¦
267 | NM | 738000 | MDC_IDC_STAT_EPISODE_RECENT_COUNT | MDC | 3 | 0 |  |  |  | F |  |  |  |  | OBX¦267¦NM¦738000^MDC_IDC_STAT_EPISODE_RECENT_COUNT^MDC¦3¦0¦¦¦¦¦¦F¦¦¦¦¦¦¦
268 | NM | 738032 | MDC_IDC_STAT_EPISODE_TOTAL_COUNT | MDC | 3 | 0 |  |  |  | F |  |  |  |  | OBX¦268¦NM¦738032^MDC_IDC_STAT_EPISODE_TOTAL_COUNT^MDC¦3¦0¦¦¦¦¦¦F¦¦¦¦¦¦¦
269 | NM | 738017 | MDC_IDC_STAT_EPISODE_RECENT_COUNT_DTM_START | MDC |  | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦269¦NM¦738017^MDC_IDC_STAT_EPISODE_RECENT_COUNT_DTM_START^MDC¦¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
270 | NM | 738018 | MDC_IDC_STAT_EPISODE_RECENT_COUNT_DTM_END | MDC |  | 20090501030000 |  |  |  | F |  |  |  |  | OBX¦270¦NM¦738018^MDC_IDC_STAT_EPISODE_RECENT_COUNT_DTM_END^MDC¦¦20090501030000¦¦¦¦¦¦F¦¦¦¦¦¦¦
271 | NM | 738049 | MDC_IDC_STAT_EPISODE_TOTAL_COUNT_DTM_START | MDC |  | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦271¦NM¦738049^MDC_IDC_STAT_EPISODE_TOTAL_COUNT_DTM_START^MDC¦¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
272 | NM | 738050 | MDC_IDC_STAT_EPISODE_TOTAL_COUNT_DTM_END | MDC |  | 20090501030000 |  |  |  | F |  |  |  |  | OBX¦272¦NM¦738050^MDC_IDC_STAT_EPISODE_TOTAL_COUNT_DTM_END^MDC¦¦20090501030000¦¦¦¦¦¦F¦¦¦¦¦¦¦
273 | ST | 739536 | MDC_IDC_EPISODE_ID | MDC | 1 | 1000 |  |  |  | F |  |  |  |  | OBX¦273¦ST¦739536^MDC_IDC_EPISODE_ID^MDC¦1¦1000¦¦¦¦¦¦F¦¦¦¦¦¦¦
274 | DTM | 739552 | MDC_IDC_EPISODE_DTM | MDC | 1 | 20090501000000 |  |  |  | F |  |  |  |  | OBX¦274¦DTM¦739552^MDC_IDC_EPISODE_DTM^MDC¦1¦20090501000000¦¦¦¦¦¦F¦¦¦¦¦¦¦
275 | CWE | 739568 | MDC_IDC_EPISODE_TYPE | MDC | 1 | Epis_VT | ^VT |  |  | F |  |  |  |  | OBX¦275¦CWE¦739568^MDC_IDC_EPISODE_TYPE^MDC¦1¦Epis_VT^VT¦¦¦¦¦¦F¦¦¦¦¦¦¦
276 | CWE | 739584 | MDC_IDC_EPISODE_TYPE_INDUCED | MDC | 1 | NO | ^NO |  |  | F |  |  |  |  | OBX¦276¦CWE¦739584^MDC_IDC_EPISODE_TYPE_INDUCED^MDC¦1¦NO^NO¦¦¦¦¦¦F¦¦¦¦¦¦¦
277 | CWE | 739600 | MDC_IDC_EPISODE_VENDOR_TYPE | MDC | 1 | VT1 | ^VT1 |  |  | F |  |  |  |  | OBX¦277¦CWE¦739600^MDC_IDC_EPISODE_VENDOR_TYPE^MDC¦1¦VT1^VT1¦¦¦¦¦¦F¦¦¦¦¦¦¦
278 | NM | 739648 | MDC_IDC_EPISODE_VENTRICULAR_INTERVAL_AT_DETECTION | MDC | 1 | 400 |  | ms |  | F |  |  |  |  | OBX¦278¦NM¦739648^MDC_IDC_EPISODE_VENTRICULAR_INTERVAL_AT_DETECTION^MDC¦1¦400¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
279 | NM | 739664 | MDC_IDC_EPISODE_VENTRICULAR_INTERVAL_AT_TERMINATION | MDC | 1 | 820 |  | ms |  | F |  |  |  |  | OBX¦279¦NM¦739664^MDC_IDC_EPISODE_VENTRICULAR_INTERVAL_AT_TERMINATION^MDC¦1¦820¦ms¦¦¦¦¦F¦¦¦¦¦¦¦
280 | ST | 739680 | MDC_IDC_EPISODE_DETECTION_THERAPY_DETAILS | MDC | 1 | 2 ATP, 5x 30J |  |  |  | F |  |  |  |  | OBX¦280¦ST¦739680^MDC_IDC_EPISODE_DETECTION_THERAPY_DETAILS^MDC¦1¦2 ATP, 5x 30J¦¦¦¦¦¦F¦¦¦¦¦¦¦
281 | CWE | 739696 | MDC_IDC_EPISODE_THERAPY_RESULT | MDC | 1 | Successful | ^Successful |  |  | F |  |  |  |  | OBX¦281¦CWE¦739696^MDC_IDC_EPISODE_THERAPY_RESULT^MDC¦1¦Successful^Successful¦¦¦¦¦¦F¦¦¦¦¦¦¦
282 | NM | 739712 | MDC_IDC_EPISODE_DURATION | MDC | 1 | 27 |  | s |  | F |  |  |  |  | OBX¦282¦NM¦739712^MDC_IDC_EPISODE_DURATION^MDC¦1¦27¦s¦¦¦¦¦F¦¦¦¦¦¦¦
MINIMUM MESSAGE
MSH¦^~\&¦APPNAME¦VENDOR¦CLINIC APP¦CLINIC ID¦20090422152341¦¦ORU^R01^ORU_R01¦12345¦P¦2.6
PID¦¦¦MODEL:XXX/SERIAL:YYY^^^BSC^U~123-12-1234^^^BSC^SS¦¦DOE^JOHN¦¦20070422153118¦M¦¦¦^^^^12345-1234
PV1¦1¦R¦¦¦¦¦MWELBY^Welby^Marcus^A^III^DR¦¦¦¦¦¦¦¦¦¦¦¦12345
OBR¦1¦¦123456¦REMOTE FOLLOW-UP¦¦¦20111025100000¦20111025100000¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦F
NTE¦1¦P¦Remote follow-up observation (a comment can go here concerning the OBR/interrogation section).
1 | CWE | 720897 | MDC_IDC_DEV_TYPE | MDC |  | CRT_D | ^Cardiac Resynchronization Therapy - Defibrillator |  |  | F |  |  |  |  | OBX¦1¦CWE¦720897^MDC_IDC_DEV_TYPE^MDC¦¦CRT_D^Cardiac Resynchronization Therapy - Defibrillator¦¦¦¦¦¦F¦¦¦¦¦¦¦
2 | ST | 720898 | MDC_IDC_DEV_MODEL | MDC |  | Device Model Name |  |  |  | F |  |  |  |  | OBX¦2¦ST¦720898^MDC_IDC_DEV_MODEL^MDC¦¦Device Model Name¦¦¦¦¦¦F¦¦¦¦¦¦¦
3 | ST | 720899 | MDC_IDC_DEV_SERIAL | MDC |  | 5867463524 |  |  |  | F |  |  |  |  | OBX¦3¦ST¦720899^MDC_IDC_DEV_SERIAL^MDC¦¦5867463524¦¦¦¦¦¦F¦¦¦¦¦¦¦
4 | CWE | 720900 | MDC_IDC_DEV_MFG | MDC |  | BSX | ^Boston Scientific |  |  | F |  |  |  |  | OBX¦4¦CWE¦720900^MDC_IDC_DEV_MFG^MDC¦¦BSX^Boston Scientific¦¦¦¦¦¦F¦¦¦¦¦¦¦
5 | DTM | 721025 | MDC_IDC_SESS_DTM | MDC |  | 20111025100000 |  |  |  | F |  |  |  |  | OBX¦5¦DTM¦721025^MDC_IDC_SESS_DTM^MDC¦¦20111025100000¦¦¦¦¦¦F¦¦¦¦¦¦¦
6 | CWE | 721026 | MDC_IDC_SESS_TYPE | MDC |  | Remote | ^Remote |  |  | F |  |  |  |  | OBX¦6¦CWE¦721026^MDC_IDC_SESS_TYPE^MDC¦¦Remote^Remote ¦¦¦¦¦¦F¦¦¦¦¦¦¦
7 | CWE | 721280 | MDC_IDC_MSMT_BATTERY_STATUS | MDC | 1 | MOS | ^Middle of Service |  |  | F |  |  |  |  | OBX¦7¦CWE¦721280^MDC_IDC_MSMT_BATTERY_STATUS^MDC¦1¦MOS^Middle of Service¦¦¦¦¦¦F¦¦¦¦¦¦¦
8 | NM | 721344 | MDC_IDC_MSMT_BATTERY_VOLTAGE | MDC | 1 | 6.3 |  | V |  | F |  |  |  |  | OBX¦8¦NM¦721344^MDC_IDC_MSMT_BATTERY_VOLTAGE^MDC¦1¦6.3¦V¦¦¦¦¦F¦¦¦¦¦¦¦
9 | CWE | 730752 | MDC_IDC_SET_BRADY_MODE | MDC |  | DDDR | ^DDDR |  |  | F |  |  |  |  | OBX¦9¦CWE¦730752^MDC_IDC_SET_BRADY_MODE^MDC¦¦DDDR^DDDR¦¦¦¦¦¦F¦¦¦¦¦¦¦
10 | NM | 730880 | MDC_IDC_SET_BRADY_LOWRATE | MDC |  | 60 |  | {beats}/min |  | F |  |  |  |  | OBX¦10¦NM¦730880^MDC_IDC_SET_BRADY_LOWRATE^MDC¦¦60¦{beats}/min¦¦¦¦¦F¦¦¦¦¦¦¦
11 | NM | 737520 | MDC_IDC_STAT_BRADY_RA_PERCENT_PACED | MDC |  | 50 |  | % |  | F |  |  |  |  | OBX¦11¦NM¦737520^MDC_IDC_STAT_BRADY_RA_PERCENT_PACED^MDC¦¦50¦%¦¦¦¦¦F¦¦¦¦¦¦¦
12 | NM | 737536 | MDC_IDC_STAT_BRADY_RV_PERCENT_PACED | MDC |  | 30 |  | % |  | F |  |  |  |  | OBX¦12¦NM¦737536^MDC_IDC_STAT_BRADY_RV_PERCENT_PACED^MDC¦¦30¦%¦¦¦¦¦F¦¦¦¦¦¦¦

## General Conformance
Index | Feature | Reference | Status | Support | Comment
GEN-1 | Implementation Description | -- | Identification of the device/application. Description of the functionality.
GEN-2 | Standard Document Revision | Standard document. | Identification of the supported revisions to IEEE 11073-10103. | Set of supported IEEE 11073-10103 revisions.
GEN-3 | Conformance Adherence - Level 1 - | -- | Base conformance declaration that device/application meets the IEEE 11073-10103 mandatory conformance requirements.
GEN-4 | Conformance Adherence - Level 2 - |  | In addition to GEN-3 declaration that device/application meets the IEEE 11073-10103 informative constraints.

## Manadatory Conformance Reqts
Index | Feature | Reference | Status | Support | Comment
REQ-1 | Identification | -- | -Mandatory- Does the implementation have strict adherence to the IEEE 11073-10103 contraints for unique reference identifier. | Yes/No (No implies NON-conformant)
REQ-2 | Coding | -- | -Mandatory- Does the implementation have strict adherence to the IEEE 11073-10103 contraints for term codes. | Yes/No (No implies NON-conformant)
REQ-3 | Attributes | -- | -Mandatory- Does the implementation have strict adherence to the IEEE 11073-10103 contraints for term attributes data type, data format, and unit of measure. | Yes/No (No implies NON-conformant)
REQ-4 | Enumerations | -- | -Mandatory- Does the implementation have strict adherence to the IEEE 11073-10103 contraints for general term enumeration. | Yes/No (No implies NON-conformant)
REQ-5 | Semantics | -- | -Mandatory- Does the implementation have strict adherence to the IEEE 11073-10103 contraints for term semantics (definition). | Yes/No (No implies NON-conformant)

## Informative Constraints
Index | Feature | Reference | Status | Support | Comment
OPT-1 | Vendor Enumeration | -- | -Optional- Does the implementation have strict adherence to the IEEE 11073-10103 informative contraints for vendor enumerations. | Yes/No (No implies NON-Level II conformance)
OPT-2 | Co-Contraints | -- | -Optional- Does the implementation have strict adherence to the IEEE 11073-10103 informative co-contraints for primary key, optionality, cardinality and co-constraint rules. | Yes/No (No implies NON-Level II conformance)